Gas Delivery and Virtual Process Chamber Concept for Gas Assisted Material Processing in Focused Ion Beam System

Valery Ray

Particle Beam Systems & Technology

290 Broadway Street Suite 298, Methuen, MA, 01844 USA;

E-mail: vray@partbeamsystech.com

ABSTRACT


Gas Assisted Etching (GAE) and material deposition using gaseous precursors have become a critical part of Integrated Circuit (IC) modification and other applications of Focused Ion Beam (FIB) systems. Widely used methods of gas injection in FIB systems are either needle type or shroud type gas delivery nozzles. Each of these methods of gas delivery has significant drawbacks. Release of the gas from the needle type nozzles detrimental for the vacuum level in main chamber of the system, as the gas quickly dissipates from the process area. Insertion of shroud concentrators into the primary ion beam path the secondary electron path leads to significant decline in S/N ratio of the FIB image. Proposed “CUPOLA” geometry of the gas delivery nozzle would allow to concentrate process gas in the area of interest, as it is currently done by shroud concentrators, and, according to secondary electron trajectory simulations, would significantly improve signal level of secondary electron image, as compared to shroud concentrators.

INTRODUCTION


Optimized gas-assisted processes are critical for effective utilization of the Focused Ion Beam (FIB) systems in industrial applications1, including semiconductor circuit modification, or circuit edit, and FIB machining of the silicon and diamond-based tips for Atomic Force Microscopy (AFM). Shrinkage of the IC feature dimensions to below 100 nm drives a growing interest in gas assisted focused electron beam processing and nanoscale depositon2, where optimal gas delivery would also be critical.


Reported in literature optimal precursor pressure levels for gas-assisted FIB etching processes are in the millitorr range. For example, measurements by Ochiai et al. indicated3, that the optimal Cl2 pressure for FIB assisted etching of GaAs is in 20 mTorr range and the optimal XeF2 pressure for FIB assisted etching of Si is close to 1 mTorr. Design of the modern FIB systems however, requires maintaining much lower pressure level in the main chamber of the system, typically  ~ 2 x 10-5 Torr4 or below, in order to achieve reliable operation. These conflicting pressure requirements create difficulty to operate at the optimal level of precursor gas pressure. 

Group of Ochiai et al. used separately pumped subchamber3, filled by the precursor gas and coupled to the main chamber of the FIB system only by the small orifice for passing primary ion beam and secondary electron signal, to study precursor pressure dependence of gas assisted FIB milling. Folch et al. also utilized precursor-filled subchamber5 to construct “EBD cell” for the study of electron beam induced deposition in SEM. While the subchamber approach to gas delivery is a convenient experimental setup, it is not used extensively in industrial FIB systems at this time. 


Capillary needles are the most commonly used method of delivery the gas close to the area of interaction between the ion or electron beam and the sample6. Nevertheless gas is released from the needle close to the surface of the sample it quickly dissipates from the work area, creating significant gas load in the main chamber of the FIB system. 

Shrouded-type gas concentrators7, such as “Beehives”8 and similar, are used today in commercial FIB systems to create high localized pressures of process gas for bulk material removal applications. High accuracy milling work, including the via milling, is also possible with the shroud concentrators and can be done with reduced level of the base pressure in a main vacuum chamber of the FIB system9. Unfortunately, introduction of the shroud concentrator into the path of the primary ion beam and the path of the secondary electrons significantly reducing the level of a secondary electron signal available for imaging. Poor imaging capability is a limiting factor for use of the “beehive” nozzles and other shroud concentrators in precision FIB applications.

Previous experimental work demonstrated10 that contrast level of the secondary electron signal, collected through the “beehive” nozzle, can be increased by up to factor of eight by biasing the nozzle positively, relatively to the potential of the system chamber. The S/N level of the image however remained too low for the high quality imaging of the small features. Improved gas delivery method is needed to optimal precursor pressure regime for the gas assisted charged particle beam processes with acceptable quality of through-nozzle imaging of the area of interest.

VIRTUAL CHAMBER CONCEPT DESCRIPTION

There is a long-settled agreement in literature11, 12 that the angle of departure of the gas molecule after the interaction with the solid sidewall of the chamber in vacuum is following a cosine distribution and that the most probable departure direction for the molecule is orthogonal to the wall plane. 

Interesting consequence of this departure angle distribution is that highly polished cylindrical or spherical surfaces could potentially have increase local pressure at the center of the sphere or along the axis of the cylinder. Nesterov et al. calculated13, that local pressure at the center of the spherical vacuum vessel can be by factor of seven larger then the pressure at the vessel walls.

It becomes logical therefore to make an end of the gas delivery nozzle semi-spherical in shape (Fig. 1) and position the process area at the center of semi-sphere.  In this configuration, area of a relatively high precursor gas pressure is confined within the virtual processing chamber, created between the semispherical end of the gas delivery nozzle and the surface of the sample. The delivery tube, supplying precursor gas to the 

[image: image1.png]



[image: image2.wmf]0.0

10.0

20.0

30.0

40.0

50.0

60.0

70.0

80.0

90.0

100.0

0

10

20

30

40

50

60

70

80

90

100

110

120

130

140

150

160

170

180

190

200

Shroud bias, Volt

Signal distribution, %%

Anode signal

Shroud signal


virtual processing chamber, is directed to the location of interaction between the primary particle beam and the sample at the center of the semi-sphere, to take additional advantage of the directionality of the molecular gas flow exiting the supply tube14. The length of the straight portion of the delivery line, directed toward the center of the virtual processing chamber, should be much longer then the diameter of the tube. 

“CUPOLA” gas delivery nozzle (Fig. 2) is designed as a practically realizable implementation of the virtual process chamber concept. It is intended for high-throughput operation of industrial FIB system in semiconductor circuit modification and AFM tip manufacturing applications. 

The nozzle should be aligned within the FIB apparatus in such a way that, while the nozzle is extended, the primary ion beam would pass through the center of the 500 (m aperture at the top of the nozzle. Circular bottom opening of the nozzle should be parallel to the sample plane and internal surface of the “CUPOLA” may need to be highly polished, to ensure that the internal surface of the nozzle is semi-spherical on the microscale.

[image: image3.wmf]0.0

10.0

20.0

30.0

40.0

50.0

60.0

70.0

80.0

90.0

100.0

0

10

20

30

40

50

60

70

80

90

100

110

120

130

140

150

160

170

180

190

200

Cupola bias, Volt

Signal distribution, %%

Anode signal

Cupola signal


[image: image4.wmf]0.0

10.0

20.0

30.0

40.0

50.0

60.0

70.0

80.0

90.0

100.0

0

10

20

30

40

50

60

70

80

90

100

110

120

130

140

150

160

170

180

190

200

Shroud bias, Volt

Signal distribution, %%

Anode signal

Shroud signal


ELECTRON OPTICS SIMULATIONS

[image: image5.png]


In order to make a comparative evaluation of the imaging capabilities of proposed “CUPOLA” nozzle design in respect to the imaging capabilities of the existing shroud gas concentrators (Fig. 3) a rotational-symmetric models of both nozzle configuration were 

[image: image6.wmf]0.0

10.0

20.0

30.0

40.0

50.0

60.0

70.0

80.0

90.0

100.0

0

10

20

30

40

50

60

70

80

90

100

110

120

130

140

150

160

170

180

190

200

Cupola bias, Volt

Signal distribution, %%

Anode signal

Cupola signal


compared in LORENTZ-2E particle trajectory simulation software from Integrated Engineering Software Company, Manitoba, Canada. 

Geometry of the “CUPOLA” nozzle was modeled after the Fig. 2. For the shroud, or “Beehive” nozzle, 2 mm width of the opening directed toward the sample was assumed to for the purposes of comparison with the “CUPOLA” design and the remaining dimensions of the shroud concentrator were calculated based on the proportions of the Fig. 3 representation of the shroud. Resulting overall height of the shroud was calculated as 9.18 mm and width of the opening at the top of the shroud calculated as 0.59 mm. 

Both nozzle models were placed on 250 (m distance from the surface of the sample. The surface of the sample, for the purposes of simulations, was assumed to be conductive and grounded. Simulated electron emission was launched from 5 (m above the surface of the sample on the axis of symmetry of each of the nozzle models with initial electron energies of 1, 2, 3, 4, 5, 7, 9, 11, 13, and 15 electron volts. Electrons with each level of energy were launched in 11 beams, equally distributed within ( 45-degree angle from the axis of symmetry of the models. Potential of both nozzles was varied from zero to two hundred volts, to evaluate possibility of the enhanced extraction of the electrons through the top aperture of the nozzle. Collection electrode, or anode, was placed on the distance of 40 mm from the surface of the sample and set to +400 Volt potential.

SIMULATION RESULTS AND DISCUSSION

[image: image7.emf][image: image8.emf]Results of the of the electron trajectory simulations were evaluated in terms of the ratio between the number of the electrons that successfully pass through the top of the nozzle and collected by the anode, to the number of the electrons that are captured on the internal surface of the nozzle. The results of evaluation are represented by the graphs on Fig. 4 for the “CUPOLA” and Fig. 5 the shroud, or “Beehive”, nozzles respectively. 

[image: image9.emf][image: image10.emf]
Even at the zero volts nozzle potential, corresponding to the grounded nozzle configuration, 28% of the simulated secondary electrons exit the “CUPOLA”. With the “Beehive” nozzle configuration only 9% of the simulated secondary electrons are exiting the grounded nozzle and collected by the anode. 

Increase of the “Beehive” nozzle potential results in a slow increase of the fraction of secondary electrons passing through the top of the nozzle, reaching the 30% of the simulated secondary electrons collected by the anode at + 200 Volt potential applied to the “Beehive” nozzle. As much as 70% of the electrons are still captured by the nozzle (Fig. 6). 

Increase of the “CUPOLA” nozzle potential results in a much stronger increase of the fraction of secondary electrons passing through the nozzle, reaching 94.5% of simulated secondary electrons collected by the anode at + 200 Volt potential applied to the “CUPOLA” nozzle. Slightly over 5% of the secondary electrons remain captured by the nozzle (Fig. 7).


It is evident from the secondary electron simulation results, that a fraction of the secondary electrons, generated by the interaction between the primary ion beam and the material of the sample during the gas assisted etching process, is collected by the gas delivery nozzle itself. This observation sets basis for the future experimental investigation of possibility to detect material-dependent transitions of the secondary electron emission during the FIB milling process by monitoring the nozzle-absorbed current. Such detection could have applications in via milling endpoint application, similarly to the sample absorbed current, or “stage current” endpoint techniques15, 16.

Current simulations did not investigate effects of interactions between the secondary electrons and the high-pressure gas within the “CUPOLA” nozzle. It is evident from the nozzle design however, that the path of the secondary electrons in the high gas pressure region within the “CUPOLA” nozzle, which is about 1 mm distance, is significantly shorter then the path within the “Beehive” or other shroud nozzle, which is close to 9 mm distance. Signal degradation due to high gas pressure region within the nozzle should therefore be reduced in the “CUPOLA”, comparatively to shroud design.

CONCLUSIONS

Proposed virtual chamber concept and the “CUPOLA” nozzle design is expected to allow carrying gas assisted processes in FIB and other focused particle beam systems under the increased process gas pressure, previously available only with shroud concentrators, but with significantly improved secondary electron image signal level, as compared to “Beehive” nozzle and other shroud-type gas concentrators. 

Electron trajectory simulations suggest that considerable fraction of secondary electron signal is absorbed by the gas delivery nozzle, creating a “nozzle absorbed current”. Detection of transitions in the nozzle current could have practical applications in FIB via milling endpoint detection application, similarly to the sample absorbed current technique.

ACKNOWLEDGEMENTS

Author would like to thank Mr. Nicholas Antoniou, Dr. Neil Bassom, Mr. Gavin Hartigan, and Dr. Kun Liu, all of FEI Company, for there’s contributions to the abstract of this manuscript.

Special thanks are due to Craig Armstrong of Integrated Engineering Software Company, Winnipeg, Canada, for providing evaluation copy of the LORENTZ-2E particle trajectory simulation software.

REFERENCES

1. J. Orloff et al., “High Resolution Focused Ion Beams”, Kluwer Academic / Plenum Publishers, 2003

2. H. Marchman et al, “The Impact of Feature Packing Density on FIB Editing of Advanced Technology ICs”, Proceedings ISTFA 2003, pp. 348 – 354

3. Y. Ochiai et al., “Pressure and Irradiation Angle Dependence of Maskless Ion Beam Assisted Etching of GaAs and Si”, J. Vac. Sci. Technol. B 3(1), 1985, pp. 67 – 70

4. N. Bassom and T. Mai, “Modeling and Optimizing XeF2-enhanced FIB Milling of Silicon”, Proceedings ISTFA 1999, pp. 255 – 261

5. A. Folch et al., “High Vacuum versus “Environmental” Electron Beam Deposition”, J. Vac. Sci. Technol. B 14(4) 1996, pp. 2609 – 2614   

6. J. Orloff, “Charged Particle Optics”, CRC Press, 1997, pp. 441 – 453

7. Casella et al, “Gas Delivery Systems for Particle Beam Processing”, US Patent #5,851,413

8. Libby et al, “Focused Particle Beam Systems and Methods Using a Tilt Column”, US Patent #6,497,194

9. V. Ray et al., “Gas Delivery and Virtual Process Chamber Concept”, Abstracts, 48th International Conference EIPBN 2004, pp. 145 – 146

10. V. Ray et al, “Improvements of Secondary Electron Imaging and Endpoint Detection in Focused Ion Beam Circuit Modification”, Proceedings ISTFA 2003, pp 338-342

11. H. A. Lorentz, “Lectures on Theoretical Physics” Vol. 1 Chapter 3, Macmillan, London 1927

12. P. Clausing, “The Flow of Highly Rarefield Gases Through Tubes of Arbitrary Length”, J. Vac. Sci. Technol. B 8(5) 1971, pp. 636 – 646

13. Nesterov et al., “Influence of the Vacuum Chamber Shape of the Non-Uniformity of Gas Distribution”, Vacuum 53 1999, pp. 193 – 196

14. J. F. O’Hanlon, “A Users Guide to Vacuum Technology”, John Wiley & Sons, 1989, pp. 44 – 45

15. V. Ray and N. Antoniou, "Stage Current Monitoring and Endpoint Detection in Focused Ion Beam Circuit Modification", 7th European FIB User Group meeting (EFUG 2003), http://www.imec.be/efug/EFUG2003_Ray1.pdf 

16. V. Ray et al., “Small Via High Aspect Ratio Circuit Edit: Challenges, Techniques and Developments”, Proceedings, ISTFA 2003, pp. 357 - 358

� EMBED MSPhotoEd.3  ���





Fig. 5





Shroud, of “Beehive” simulation results.





� EMBED Excel.Sheet.8  ���





Fig. 1  Virtual process chamber concept, comprising of semi-spherical attachment at the end of gas delivery nozzle





Fig. 2  “CUPOLA” nozzle designed as a practical implementation of the virtual process chamber concept.





� EMBED Excel.Sheet.8  ���





Fig. 3  





Shroud, or “Beehive”, nozzle presented on Fig. 2 in US Patent 5,851,413





Fig. 4





“CUPOLA” simulation results.





Fig. 6





Modeled trajectories of the secondary electrons, extracted through the positively biased “Beehive” gas delivery nozzle under +200 Volt nozzle bias potential. Most of electrons are absorbed by the internal surface of the shroud.








Fig. 7





Modeled trajectories of the secondary electrons, extracted through the positively biased “CUPOLA” gas delivery nozzle under + 200 Volt nozzle bias potential. Most of electrons are extracted through the aperture at the top of the “CUPOLA” nozzle.











© Valery Ray, PBS&T
Page 6 of 9
June, 2004

_1148647182.bin

_1148649768.xls
Chart1

		0		0

		10		10

		20		20

		30		30

		40		40

		50		50

		60		60

		70		70

		80		80

		90		90

		100		100

		110		110

		120		120

		130		130

		140		140

		150		150

		160		160

		170		170

		180		180

		190		190

		200		200



Anode signal

Shroud signal

Shroud bias, Volt

Signal distribution, %%

9.0909090909

90.9090909091

11.8181818182

88.1818181818

14.5454545455

85.4545454545

16.3636363636

83.6363636364

17.2727272727

82.7272727273

20

80

20.9090909091

79.0909090909

21.8181818182

78.1818181818

21.8181818182

78.1818181818

22.7272727273

77.2727272727

23.6363636364

76.3636363636

25

75

25.4545454545

74.5454545455

25.4545454545

74.5454545455

25.9090909091

74.0909090909

26.3636363636

73.6363636364

27.2727272727

72.7272727273

28.1818181818

71.8181818182

28.1818181818

71.8181818182

28.1818181818

71.8181818182

30

70



Summary

		

		Bias Potential		Anode		Shroud

				%		%

		0		9.1		90.9

		10		11.8		88.2

		20		14.5		85.5

		30		16.4		83.6

		40		17.3		82.7

		50		20.0		80.0

		60		20.9		79.1

		70		21.8		78.2

		80		21.8		78.2

		90		22.7		77.3

		100		23.6		76.4

		110		25.0		75.0

		120		25.5		74.5

		130		25.5		74.5

		140		25.9		74.1

		150		26.4		73.6

		160		27.3		72.7

		170		28.2		71.8

		180		28.2		71.8

		190		28.2		71.8

		200		30.0		70.0





Summary

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0

		0		0



Anode signal

Shroud signal

Shroud bias, Volt

Signal distribution, %%

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0

0



0 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		----------------------		-------		---------------------		----------------

		1		195		0.000163697		40000		0		0				1

		2		125		730.429		7313.3		0		0				0

		3		80		912.536		4568.36		0		0				0

		4		60		995.16		3322.95		0		0				0

		5		47		1000		2505.83		0		0				0

		6		39		1000		2005.87		0		0				0

		7		34		1000		1672.55		0		0				0

		8		31		1000		1434.47		0		0				0

		9		28		1000		1255.9		0		0				0

		10		27		1000		1117.02		0		0				0

		11		25		1000		1005.91		0		0				0

		12		25		-1000		1005.91		0		0				0

		13		27		-1000		1117.02		0		0				0

		14		28		-1000		1255.9		0		0				0

		15		31		-1000		1434.47		0		0				0

		16		34		-1000		1672.55		0		0				0

		17		39		-1000		2005.87		0		0				0

		18		47		-1000		2505.83		0		0				0

		19		60		-995.16		3322.95		0		0				0

		20		80		-912.536		4568.36		0		0				0

		21		125		-730.429		7313.3		0		0				0

		22		195		0.000163697		40000		0		0				1

		23		151		-4.47E-04		40000		0		0				1

		24		89		730.488		7312.4		0		0				0

		25		57		912.528		4568.48		0		0				0

		26		43		995.154		3323.05		0		0				0

		27		34		1000		2505.91		0		0				0

		28		28		1000		2005.93		0		0				0

		29		25		1000		1672.61		0		0				0

		30		22		1000		1434.52		0		0				0

		31		21		1000		1255.95		0		0				0

		32		19		1000		1117.06		0		0				0

		33		18		1000		1005.95		0		0				0

		34		18		-1000		1005.95		0		0				0

		35		19		-1000		1117.06		0		0				0

		36		21		-1000		1255.95		0		0				0

		37		22		-1000		1434.52		0		0				0

		38		25		-1000		1672.61		0		0				0

		39		28		-1000		2005.93		0		0				0

		40		34		-1000		2505.91		0		0				0

		41		43		-995.154		3323.05		0		0				0

		42		57		-912.528		4568.48		0		0				0

		43		89		-730.489		7312.39		0		0				0

		44		151		-4.47E-04		40000		0		0				1

		45		131		6.29E-04		40000		0		0				1

		46		73		730.508		7312.09		0		0				0

		47		47		912.526		4568.52		0		0				0

		48		35		995.151		3323.08		0		0				0

		49		28		1000		2505.94		0		0				0

		50		23		1000		2005.95		0		0				0

		51		20		1000		1672.63		0		0				0

		52		18		1000		1434.54		0		0				0

		53		17		1000		1255.97		0		0				0

		54		16		1000		1117.08		0		0				0

		55		15		1000		1005.97		0		0				0

		56		15		-1000		1005.97		0		0				0

		57		16		-1000		1117.08		0		0				0

		58		17		-1000		1255.97		0		0				0

		59		18		-1000		1434.54		0		0				0

		60		20		-1000		1672.63		0		0				0

		61		23		-1000		2005.95		0		0				0

		62		28		-1000		2505.94		0		0				0

		63		35		-995.151		3323.08		0		0				0

		64		47		-912.526		4568.52		0		0				0

		65		73		-730.509		7312.09		0		0				0

		66		131		6.29E-04		40000		0		0				1

		67		119		3.92E-05		40000		0		0				1

		68		63		730.519		7311.94		0		0				0

		69		41		912.524		4568.54		0		0				0

		70		31		995.15		3323.1		0		0				0

		71		24		1000		2505.95		0		0				0

		72		20		1000		2005.96		0		0				0

		73		18		1000		1672.64		0		0				0

		74		16		1000		1434.54		0		0				0

		75		15		1000		1255.97		0		0				0

		76		14		1000		1117.09		0		0				0

		77		13		1000		1005.98		0		0				0

		78		13		-1000		1005.98		0		0				0

		79		14		-1000		1117.09		0		0				0

		80		15		-1000		1255.97		0		0				0

		81		16		-1000		1434.54		0		0				0

		82		18		-1000		1672.64		0		0				0

		83		20		-1000		2005.96		0		0				0

		84		24		-1000		2505.95		0		0				0

		85		31		-995.15		3323.1		0		0				0

		86		41		-912.524		4568.54		0		0				0

		87		63		-730.519		7311.94		0		0				0

		88		119		3.92E-05		40000		0		0				1

		89		111		0.000223259		40000		0		0				1

		90		57		730.525		7311.84		0		0				0

		91		37		912.523		4568.55		0		0				0

		92		28		995.15		3323.11		0		0				0

		93		22		1000		2505.96		0		0				0

		94		18		1000		2005.97		0		0				0

		95		16		1000		1672.64		0		0				0

		96		15		1000		1434.55		0		0				0

		97		14		1000		1255.98		0		0				0

		98		13		1000		1117.09		0		0				0

		99		12		1000		1005.98		0		0				0

		100		12		-1000		1005.98		0		0				0

		101		13		-1000		1117.09		0		0				0

		102		14		-1000		1255.98		0		0				0

		103		15		-1000		1434.55		0		0				0

		104		16		-1000		1672.64		0		0				0

		105		18		-1000		2005.97		0		0				0

		106		22		-1000		2505.96		0		0				0

		107		28		-995.15		3323.11		0		0				0

		108		37		-912.524		4568.55		0		0				0

		109		57		-730.525		7311.84		0		0				0

		110		111		0.000223259		40000		0		0				1

		111		100		-4.22E-05		40000		0		0				1

		112		48		730.532		7311.74		0		0				0

		113		31		912.523		4568.56		0		0				0

		114		24		995.149		3323.12		0		0				0

		115		19		1000		2505.97		0		0				0

		116		16		1000		2005.98		0		0				0

		117		14		1000		1672.65		0		0				0

		118		13		1000		1434.55		0		0				0

		119		12		1000		1255.98		0		0				0

		120		11		1000		1117.1		0		0				0

		121		11		1000		1005.99		0		0				0

		122		11		-1000		1005.99		0		0				0

		123		11		-1000		1117.1		0		0				0

		124		12		-1000		1255.98		0		0				0

		125		13		-1000		1434.55		0		0				0

		126		14		-1000		1672.65		0		0				0

		127		16		-1000		2005.98		0		0				0

		128		19		-1000		2505.97		0		0				0

		129		24		-995.149		3323.12		0		0				0

		130		31		-912.523		4568.56		0		0				0

		131		48		-730.532		7311.74		0		0				0

		132		100		-4.22E-05		40000		0		0				1

		133		93		-3.04E-04		40000		0		0				1

		134		43		730.536		7311.68		0		0				0

		135		28		912.522		4568.57		0		0				0

		136		21		995.148		3323.13		0		0				0

		137		17		1000		2505.98		0		0				0

		138		14		1000		2005.98		0		0				0

		139		12		1000		1672.65		0		0				0

		140		11		1000		1434.56		0		0				0

		141		10		1000		1255.99		0		0				0

		142		10		1000		1117.1		0		0				0

		143		9		1000		1005.99		0		0				0

		144		9		-1000		1005.99		0		0				0

		145		10		-1000		1117.1		0		0				0

		146		10		-1000		1255.99		0		0				0

		147		11		-1000		1434.56		0		0				0

		148		12		-1000		1672.65		0		0				0

		149		14		-1000		2005.98		0		0				0

		150		17		-1000		2505.98		0		0				0

		151		21		-995.148		3323.13		0		0				0

		152		28		-912.522		4568.57		0		0				0

		153		43		-730.536		7311.68		0		0				0

		154		93		-3.04E-04		40000		0		0				1

		155		88		2.26E-04		40000		0		0				1

		156		39		730.538		7311.65		0		0				0

		157		25		912.522		4568.58		0		0				0

		158		19		995.148		3323.13		0		0				0

		159		15		1000		2505.98		0		0				0

		160		13		1000		2005.99		0		0				0

		161		11		1000		1672.65		0		0				0

		162		10		1000		1434.56		0		0				0

		163		10		1000		1255.99		0		0				0

		164		9		1000		1117.1		0		0				0

		165		9		1000		1005.99		0		0				0

		166		9		-1000		1005.99		0		0				0

		167		9		-1000		1117.1		0		0				0

		168		10		-1000		1255.99		0		0				0

		169		10		-1000		1434.56		0		0				0

		170		11		-1000		1672.65		0		0				0

		171		13		-1000		2005.99		0		0				0

		172		15		-1000		2505.98		0		0				0

		173		19		-995.148		3323.13		0		0				0

		174		25		-912.522		4568.58		0		0				0

		175		39		-730.538		7311.65		0		0				0

		176		88		2.26E-04		40000		0		0				1

		177		84		-2.55E-04		40000		0		0				1

		178		36		730.54		7311.62		0		0				0

		179		23		912.521		4568.58		0		0				0

		180		18		995.148		3323.14		0		0				0

		181		14		1000		2505.98		0		0				0

		182		12		1000		2005.99		0		0				0

		183		11		1000		1672.66		0		0				0

		184		10		1000		1434.56		0		0				0

		185		9		1000		1255.99		0		0				0

		186		8		1000		1117.1		0		0				0

		187		8		1000		1005.99		0		0				0

		188		8		-1000		1005.99		0		0				0

		189		8		-1000		1117.1		0		0				0

		190		9		-1000		1255.99		0		0				0

		191		10		-1000		1434.56		0		0				0

		192		11		-1000		1672.66		0		0				0

		193		12		-1000		2005.99		0		0				0

		194		14		-1000		2505.98		0		0				0

		195		18		-995.148		3323.14		0		0				0

		196		23		-912.521		4568.58		0		0				0

		197		36		-730.54		7311.62		0		0				0

		198		84		-2.55E-04		40000		0		0				1

		199		80		9.37E-05		40000		0		0				1

		200		33		730.541		7311.6		0		0				0

		201		22		912.521		4568.58		0		0				0

		202		17		995.148		3323.14		0		0				0

		203		13		1000		2505.99		0		0				0

		204		11		1000		2005.99		0		0				0

		205		10		1000		1672.66		0		0				0

		206		9		1000		1434.56		0		0				0

		207		8		1000		1255.99		0		0				0

		208		8		1000		1117.1		0		0				0

		209		8		1000		1005.99		0		0				0

		210		8		-1000		1005.99		0		0				0

		211		8		-1000		1117.1		0		0				0

		212		8		-1000		1255.99		0		0				0

		213		9		-1000		1434.56		0		0				0

		214		10		-1000		1672.66		0		0				0

		215		11		-1000		2005.99		0		0				0

		216		13		-1000		2505.99		0		0				0

		217		17		-995.148		3323.14		0		0				0

		218		22		-912.521		4568.58		0		0				0

		219		33		-730.541		7311.6		0		0				0

		220		80		9.37E-05		40000		0		0				1

																				%%

										Electrons Reaching Anode						20				9.1

										Electrons falling on Cupola						200				90.9





10 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-----------------------		-------		---------------------		----------------

		1		92		-8.53E-05		40000		0		0				1

		2		92		-754.009		40000		0		0				1

		3		45		687.613		7958.66		0		0				0

		4		36		800.141		6262.51		0		0				0

		5		31		869.873		5211.43		0		0				0

		6		28		916.654		4506.29		0		0				0

		7		25		949.766		4007.19		0		0				0

		8		23		974.14		3639.8		0		0				0

		9		22		992.634		3361.02		0		0				0

		10		21		1000		3121.73		0		0				0

		11		20		1000		2916.81		0		0				0

		12		20		-1000		2916.81		0		0				0

		13		21		-1000		3121.74		0		0				0

		14		22		-992.634		3361.02		0		0				0

		15		23		-974.14		3639.8		0		0				0

		16		25		-949.766		4007.19		0		0				0

		17		28		-916.654		4506.29		0		0				0

		18		31		-869.873		5211.43		0		0				0

		19		36		-800.141		6262.51		0		0				0

		20		45		-687.613		7958.66		0		0				0

		21		92		754.009		40000		0		0				1

		22		92		-8.53E-05		40000		0		0				1

		23		89		0.0001486		40000		0		0				1

		24		89		-993.722		40000		0		0				1

		25		37		759.807		6870.47		0		0				0

		26		30		865.349		5279.62		0		0				0

		27		25		928.832		4322.73		0		0				0

		28		22		970.719		3691.36		0		0				0

		29		20		1000		3248.04		0		0				0

		30		18		1000		2860.36		0		0				0

		31		17		1000		2577.01		0		0				0

		32		16		1000		2362.39		0		0				0

		33		15		1000		2195.26		0		0				0

		34		15		-1000		2195.26		0		0				0

		35		16		-1000		2362.39		0		0				0

		36		17		-1000		2577.01		0		0				0

		37		18		-1000		2860.36		0		0				0

		38		20		-1000		3248.04		0		0				0

		39		22		-970.719		3691.36		0		0				0

		40		25		-928.832		4322.73		0		0				0

		41		30		-865.349		5279.62		0		0				0

		42		37		-759.807		6870.47		0		0				0

		43		89		993.722		40000		0		0				1

		44		89		0.0001486		40000		0		0				1

		45		86		1.10E-04		40000		0		0				1

		46		87		-1147.75		40000		0		0				1

		47		33		794.667		6345.02		0		0				0

		48		26		895.91		4818.96		0		0				0

		49		22		955.98		3913.53		0		0				0

		50		19		995.348		3320.11		0		0				0

		51		17		1000		2838.97		0		0				0

		52		16		1000		2491.37		0		0				0

		53		15		1000		2236.25		0		0				0

		54		14		1000		2042.19		0		0				0

		55		13		1000		1890.29		0		0				0

		56		13		-1000		1890.29		0		0				0

		57		14		-1000		2042.19		0		0				0

		58		15		-1000		2236.25		0		0				0

		59		16		-1000		2491.37		0		0				0

		60		17		-1000		2838.97		0		0				0

		61		19		-995.348		3320.11		0		0				0

		62		22		-955.98		3913.53		0		0				0

		63		26		-895.91		4818.96		0		0				0

		64		33		-794.667		6345.02		0		0				0

		65		87		1147.75		40000		0		0				1

		66		86		1.10E-04		40000		0		0				1

		67		84		2.11E-06		40000		0		0				1

		68		44		607.921		9159.87		0		0				0

		69		30		815.901		6024.96		0		0				0

		70		24		914.235		4542.76		0		0				0

		71		20		972.117		3670.28		0		0				0

		72		17		1000		3069.75		0		0				0

		73		16		1000		2606.41		0		0				0

		74		14		1000		2281.26		0		0				0

		75		13		1000		2041.92		0		0				0

		76		13		1000		1859.38		0		0				0

		77		12		1000		1716.1		0		0				0

		78		12		-1000		1716.1		0		0				0

		79		13		-1000		1859.38		0		0				0

		80		13		-1000		2041.92		0		0				0

		81		14		-1000		2281.26		0		0				0

		82		16		-1000		2606.41		0		0				0

		83		17		-1000		3069.75		0		0				0

		84		20		-972.117		3670.28		0		0				0

		85		24		-914.235		4542.76		0		0				0

		86		30		-815.901		6024.96		0		0				0

		87		44		-607.921		9159.87		0		0				0

		88		84		2.11E-06		40000		0		0				1

		89		83		-6.05E-05		40000		0		0				1

		90		41		630.062		8826.14		0		0				0

		91		28		830.383		5806.67		0		0				0

		92		22		926.61		4356.22		0		0				0

		93		19		982.96		3506.85		0		0				0

		94		16		1000		2895.87		0		0				0

		95		14		1000		2454.25		0		0				0

		96		13		1000		2143.7		0		0				0

		97		12		1000		1914.57		0		0				0

		98		12		1000		1739.44		0		0				0

		99		11		1000		1601.45		0		0				0

		100		11		-1000		1601.56		0		0				0

		101		12		-1000		1739.45		0		0				0

		102		12		-1000		1914.57		0		0				0

		103		13		-1000		2143.7		0		0				0

		104		14		-1000		2454.25		0		0				0

		105		16		-1000		2895.87		0		0				0

		106		19		-982.96		3506.85		0		0				0

		107		22		-926.61		4356.22		0		0				0

		108		28		-830.383		5806.67		0		0				0

		109		41		-630.062		8826.14		0		0				0

		110		83		-6.05E-05		40000		0		0				1

		111		80		1.10E-05		40000		0		0				1

		112		37		653.129		8478.45		0		0				0

		113		25		849.06		5525.14		0		0				0

		114		20		942.43		4117.76		0		0				0

		115		17		996.758		3298.86		0		0				0

		116		14		1000		2680.01		0		0				0

		117		13		1000		2265.21		0		0				0

		118		12		1000		1972.64		0		0				0

		119		11		1000		1756.09		0		0				0

		120		10		1000		1589.78		0		0				0

		121		10		1000		1458.78		0		0				0

		122		10		-1000		1458.78		0		0				0

		123		10		-1000		1589.78		0		0				0

		124		11		-1000		1756.09		0		0				0

		125		12		-1000		1972.64		0		0				0

		126		13		-1000		2265.21		0		0				0

		127		14		-1000		2680.01		0		0				0

		128		17		-996.758		3298.86		0		0				0

		129		20		-942.43		4117.76		0		0				0

		130		25		-849.06		5525.14		0		0				0

		131		37		-653.129		8478.45		0		0				0

		132		80		1.10E-05		40000		0		0				1

		133		77		2991.15		40000		0		0				1

		134		34		667.09		8268.01		0		0				0

		135		23		860.686		5349.91		0		0				0

		136		18		952.201		3970.49		0		0				0

		137		15		1000		3154.29		0		0				0

		138		13		1000		2550.12		0		0				0

		139		11		1000		2151.28		0		0				0

		140		10		1000		1869.41		0		0				0

		141		10		1000		1660.34		0		0				0

		142		9		1000		1499.42		0		0				0

		143		9		1000		1372.43		0		0				0

		144		9		-1000		1372.43		0		0				0

		145		9		-1000		1499.42		0		0				0

		146		10		-1000		1660.34		0		0				0

		147		10		-1000		1869.41		0		0				0

		148		11		-1000		2151.28		0		0				0

		149		13		-1000		2550.12		0		0				0

		150		15		-1000		3154.29		0		0				0

		151		18		-952.201		3970.49		0		0				0

		152		23		-860.686		5349.9		0		0				0

		153		34		-667.09		8268.01		0		0				0

		154		77		2991.15		40000		0		0				1

		155		75		-3.39E-05		40000		0		0				1

		156		32		676.667		8123.66		0		0				0

		157		22		868.66		5229.71		0		0				0

		158		17		958.87		3869.96		0		0				0

		159		14		1000		3050.04		0		0				0

		160		12		1000		2462.86		0		0				0

		161		11		1000		2074.71		0		0				0

		162		10		1000		1799.98		0		0				0

		163		9		1000		1595.73		0		0				0

		164		9		1000		1438.68		0		0				0

		165		8		1000		1313.96		0		0				0

		166		8		-1000		1313.96		0		0				0

		167		9		-1000		1438.68		0		0				0

		168		9		-1000		1595.73		0		0				0

		169		10		-1000		1799.98		0		0				0

		170		11		-1000		2074.71		0		0				0

		171		12		-1000		2462.86		0		0				0

		172		14		-1000		3050.04		0		0				0

		173		17		-958.87		3869.96		0		0				0

		174		22		-868.66		5229.71		0		0				0

		175		32		-676.667		8123.66		0		0				0

		176		75		-3.39E-05		40000		0		0				1

		177		73		-7.88E-05		40000		0		0				1

		178		30		683.699		8017.66		0		0				0

		179		20		874.484		5141.92		0		0				0

		180		16		963.725		3796.78		0		0				0

		181		13		1000		2975.01		0		0				0

		182		11		1000		2399.99		0		0				0

		183		10		1000		2019.55		0		0				0

		184		9		1000		1749.8		0		0				0

		185		9		1000		1549.39		0		0				0

		186		8		1000		1394.54		0		0				0

		187		8		1000		1272.12		0		0				0

		188		8		-1000		1272.12		0		0				0

		189		8		-1000		1394.54		0		0				0

		190		9		-1000		1549.39		0		0				0

		191		9		-1000		1749.8		0		0				0

		192		10		-1000		2019.55		0		0				0

		193		11		-1000		2399.99		0		0				0

		194		13		-1000		2975.01		0		0				0

		195		16		-963.725		3796.78		0		0				0

		196		20		-874.484		5141.92		0		0				0

		197		30		-683.699		8017.66		0		0				0

		198		73		-7.88E-05		40000		0		0				1

		199		71		-8.13E-05		40000		0		0				1

		200		29		689.09		7936.4		0		0				0

		201		19		878.931		5074.9		0		0				0

		202		15		967.423		3741.04		0		0				0

		203		12		1000		2918.36		0		0				0

		204		11		1000		2352.54		0		0				0

		205		9		1000		1977.79		0		0				0

		206		9		1000		1712.05		0		0				0

		207		8		1000		1514.03		0		0				0

		208		8		1000		1361.49		0		0				0

		209		7		1000		1240.14		0		0				0

		210		7		-1000		1240.14		0		0				0

		211		8		-1000		1361.49		0		0				0

		212		8		-1000		1514.03		0		0				0

		213		9		-1000		1712.05		0		0				0

		214		9		-1000		1977.79		0		0				0

		215		11		-1000		2352.54		0		0				0

		216		12		-1000		2918.36		0		0				0

		217		15		-967.423		3741.04		0		0				0

		218		19		-878.931		5074.9		0		0				0

		219		29		-689.09		7936.4		0		0				0

		220		71		-8.13E-05		40000		0		0				1

																				%%

																26				11.8

																194				88.2





20 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		----------------------		-------		---------------------		----------------

		1		77		-0.000103666		40000		0		0				1

		2		78		-23.4583		40000		0		0				1

		3		78		-425.607		40000		0		0				1

		4		31		719.982		7470.76		0		0				0

		5		27		795.338		6334.9		0		0				0

		6		25		847.079		5555.01		0		0				0

		7		23		884.223		4995.12		0		0				0

		8		21		911.799		4579.48		0		0				0

		9		20		932.82		4262.61		0		0				0

		10		19		949.2		4015.72		0		0				0

		11		18		962.199		3819.78		0		0				0

		12		18		-962.199		3819.78		0		0				0

		13		19		-949.2		4015.72		0		0				0

		14		20		-932.82		4262.61		0		0				0

		15		21		-911.799		4579.48		0		0				0

		16		23		-884.223		4995.12		0		0				0

		17		25		-847.079		5555		0		0				0

		18		27		-795.338		6334.9		0		0				0

		19		31		-719.982		7470.76		0		0				0

		20		78		425.571		40000		0		0				1

		21		78		23.4584		40000		0		0				1

		22		77		-0.000103666		40000		0		0				1

		23		76		-0.0002234		40000		0		0				1

		24		76		-23.9643		40000		0		0				1

		25		32		687.673		7957.76		0		0				0

		26		26		800.147		6262.42		0		0				0

		27		22		869.877		5211.37		0		0				0

		28		20		916.658		4506.22		0		0				0

		29		18		949.77		4007.13		0		0				0

		30		17		974.144		3639.73		0		0				0

		31		16		992.639		3360.95		0		0				0

		32		15		1000		3121.64		0		0				0

		33		15		1000		2916.74		0		0				0

		34		15		-1000		2916.74		0		0				0

		35		15		-1000		3121.64		0		0				0

		36		16		-992.639		3360.95		0		0				0

		37		17		-974.144		3639.73		0		0				0

		38		18		-949.77		4007.13		0		0				0

		39		20		-916.658		4506.22		0		0				0

		40		22		-869.877		5211.37		0		0				0

		41		26		-800.147		6262.42		0		0				0

		42		32		-687.673		7957.76		0		0				0

		43		76		23.9643		40000		0		0				1

		44		76		-0.0002234		40000		0		0				1

		45		75		-4.02E-05		40000		0		0				1

		46		75		-13.4163		40000		0		0				1

		47		29		731.695		7294.21		0		0				0

		48		23		840.25		5657.94		0		0				0

		49		20		906.309		4662.22		0		0				0

		50		18		950.163		4001.21		0		0				0

		51		16		981.023		3536.04		0		0				0

		52		15		1000		3182.66		0		0				0

		53		14		1000		2874.31		0		0				0

		54		13		1000		2641.4		0		0				0

		55		13		1000		2460.59		0		0				0

		56		13		-1000		2460.59		0		0				0

		57		13		-1000		2641.39		0		0				0

		58		14		-1000		2874.31		0		0				0

		59		15		-1000		3182.66		0		0				0

		60		16		-981.023		3536.04		0		0				0

		61		18		-950.163		4001.21		0		0				0

		62		20		-906.309		4662.22		0		0				0

		63		23		-840.25		5657.94		0		0				0

		64		29		-731.695		7294.21		0		0				0

		65		75		13.4164		40000		0		0				1

		66		75		-4.02E-05		40000		0		0				1

		67		74		-3.21E-05		40000		0		0				1

		68		74		34.1508		40000		0		0				1

		69		27		759.814		6870.37		0		0				0

		70		21		865.352		5279.57		0		0				0

		71		18		928.835		4322.67		0		0				0

		72		16		970.723		3691.3		0		0				0

		73		15		1000		3247.98		0		0				0

		74		13		1000		2860.29		0		0				0

		75		13		1000		2576.97		0		0				0

		76		12		1000		2362.37		0		0				0

		77		11		1000		2195.16		0		0				0

		78		11		-1000		2195.16		0		0				0

		79		12		-1000		2362.37		0		0				0

		80		13		-1000		2576.97		0		0				0

		81		13		-1000		2860.29		0		0				0

		82		15		-1000		3247.98		0		0				0

		83		16		-970.723		3691.3		0		0				0

		84		18		-928.835		4322.67		0		0				0

		85		21		-865.352		5279.57		0		0				0

		86		27		-759.814		6870.37		0		0				0

		87		74		-34.1508		40000		0		0				1

		88		74		-3.21E-05		40000		0		0				1

		89		73		-7.77E-06		40000		0		0				1

		90		73		100.596		40000		0		0				1

		91		25		779.706		6570.54		0		0				0

		92		20		882.866		5015.58		0		0				0

		93		17		944.43		4087.62		0		0				0

		94		15		984.89		3477.76		0		0				0

		95		14		1000		3010.18		0		0				0

		96		12		1000		2645.85		0		0				0

		97		12		1000		2378.99		0		0				0

		98		11		1000		2176.42		0		0				0

		99		10		1000		2018.19		0		0				0

		100		10		-1000		2018.19		0		0				0

		101		11		-1000		2176.42		0		0				0

		102		12		-1000		2378.99		0		0				0

		103		12		-1000		2645.85		0		0				0

		104		14		-1000		3010.18		0		0				0

		105		15		-984.889		3477.76		0		0				0

		106		17		-944.43		4087.62		0		0				0

		107		20		-882.866		5015.58		0		0				0

		108		25		-779.706		6570.54		0		0				0

		109		73		-100.596		40000		0		0				1

		110		73		-7.77E-06		40000		0		0				1

		111		71		6.77E-05		40000		0		0				1

		112		33		599.05		9293.58		0		0				0

		113		23		806.411		6168.01		0		0				0

		114		18		906.071		4665.81		0		0				0

		115		15		964.941		3778.45		0		0				0

		116		13		1000		3186.96		0		0				0

		117		12		1000		2708.91		0		0				0

		118		11		1000		2373.92		0		0				0

		119		10		1000		2127.55		0		0				0

		120		10		1000		1939.99		0		0				0

		121		9		1000		1792.8		0		0				0

		122		9		-1000		1792.8		0		0				0

		123		10		-1000		1939.99		0		0				0

		124		10		-1000		2127.55		0		0				0

		125		11		-1000		2373.92		0		0				0

		126		12		-1000		2708.91		0		0				0

		127		13		-1000		3186.96		0		0				0

		128		15		-964.941		3778.45		0		0				0

		129		18		-906.071		4665.81		0		0				0

		130		23		-806.411		6168.01		0		0				0

		131		33		-599.05		9293.58		0		0				0

		132		71		6.77E-05		40000		0		0				1

		133		70		7.11E-05		40000		0		0				1

		134		30		623.17		8930.03		0		0				0

		135		21		823.761		5906.47		0		0				0

		136		17		920.964		4441.33		0		0				0

		137		14		978.019		3581.32		0		0				0

		138		12		1000		2974.61		0		0				0

		139		11		1000		2523.2		0		0				0

		140		10		1000		2205.98		0		0				0

		141		9		1000		1972.29		0		0				0

		142		9		1000		1793.83		0		0				0

		143		9		1000		1653.5		0		0				0

		144		9		-1000		1653.5		0		0				0

		145		9		-1000		1793.83		0		0				0

		146		9		-1000		1972.29		0		0				0

		147		10		-1000		2205.98		0		0				0

		148		11		-1000		2523.2		0		0				0

		149		12		-1000		2974.61		0		0				0

		150		14		-978.019		3581.32		0		0				0

		151		17		-920.964		4441.33		0		0				0

		152		21		-823.761		5906.47		0		0				0

		153		30		-623.17		8930.03		0		0				0

		154		70		7.11E-05		40000		0		0				1

		155		68		-1.56E-05		40000		0		0				1

		156		29		637.996		8706.54		0		0				0

		157		20		836.054		5721.19		0		0				0

		158		16		931.43		4283.57		0		0				0

		159		13		987.172		3443.36		0		0				0

		160		11		1000		2829.35		0		0				0

		161		10		1000		2395.97		0		0				0

		162		9		1000		2090.99		0		0				0

		163		9		1000		1865.83		0		0				0

		164		8		1000		1693.24		0		0				0

		165		8		1000		1557.77		0		0				0

		166		8		-1000		1557.77		0		0				0

		167		8		-1000		1693.24		0		0				0

		168		9		-1000		1865.83		0		0				0

		169		9		-1000		2090.99		0		0				0

		170		10		-1000		2395.97		0		0				0

		171		11		-1000		2829.35		0		0				0

		172		13		-987.172		3443.36		0		0				0

		173		16		-931.43		4283.57		0		0				0

		174		20		-836.054		5721.19		0		0				0

		175		29		-637.996		8706.54		0		0				0

		176		68		-1.56E-05		40000		0		0				1

		177		67		-9.79E-05		40000		0		0				1

		178		27		648.895		8542.27		0		0				0

		179		19		845.264		5582.36		0		0				0

		180		15		939.227		4166.04		0		0				0

		181		12		993.971		3340.88		0		0				0

		182		11		1000		2723.15		0		0				0

		183		10		1000		2303		0		0				0

		184		9		1000		2006.86		0		0				0

		185		8		1000		1787.62		0		0				0

		186		8		1000		1619.85		0		0				0

		187		8		1000		1487.41		0		0				0

		188		8		-1000		1487.41		0		0				0

		189		8		-1000		1619.85		0		0				0

		190		8		-1000		1787.62		0		0				0

		191		9		-1000		2006.86		0		0				0

		192		10		-1000		2303		0		0				0

		193		11		-1000		2723.15		0		0				0

		194		12		-993.971		3340.88		0		0				0

		195		15		-939.227		4166.04		0		0				0

		196		19		-845.264		5582.36		0		0				0

		197		27		-648.895		8542.27		0		0				0

		198		67		-9.79E-05		40000		0		0				1

		199		66		3.29E-05		40000		0		0				1

		200		26		657.387		8414.26		0		0				0

		201		18		852.444		5474.14		0		0				0

		202		14		945.28		4074.8		0		0				0

		203		12		999.238		3261.49		0		0				0

		204		10		1000		2641.82		0		0				0

		205		9		1000		2231.66		0		0				0

		206		8		1000		1942.24		0		0				0

		207		8		1000		1727.99		0		0				0

		208		7		1000		1563.33		0		0				0

		209		7		1000		1433.04		0		0				0

		210		7		-1000		1433.04		0		0				0

		211		7		-1000		1563.33		0		0				0

		212		8		-1000		1727.99		0		0				0

		213		8		-1000		1942.24		0		0				0

		214		9		-1000		2231.66		0		0				0

		215		10		-1000		2641.82		0		0				0

		216		12		-999.238		3261.49		0		0				0

		217		14		-945.28		4074.8		0		0				0

		218		18		-852.444		5474.14		0		0				0

		219		26		-657.387		8414.26		0		0				0

		220		66		3.29E-05		40000		0		0				1

																				%%

																32				14.5

																188				85.5





30 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-----------------------		-------		---------------------		----------------

		1		71		-5.53E-05		40000		0		0				1

		2		71		214.173		40000		0		0				1

		3		71		289.965		40000		0		0				1

		4		29		667.532		8261.34		0		0				0

		5		25		745.338		7088.56		0		0				0

		6		23		799.592		6270.78		0		0				0

		7		21		838.956		5677.44		0		0				0

		8		20		868.383		5233.88		0		0				0

		9		19		890.919		4894.2		0		0				0

		10		18		908.527		4628.79		0		0				0

		11		17		922.524		4417.81		0		0				0

		12		17		-922.524		4417.81		0		0				0

		13		18		-908.527		4628.79		0		0				0

		14		19		-890.919		4894.2		0		0				0

		15		20		-868.383		5233.88		0		0				0

		16		21		-838.956		5677.44		0		0				0

		17		23		-799.592		6270.78		0		0				0

		18		25		-745.338		7088.56		0		0				0

		19		29		-667.532		8261.34		0		0				0

		20		71		-289.965		40000		0		0				1

		21		71		-214.173		40000		0		0				1

		22		71		-5.53E-05		40000		0		0				1

		23		70		0.00023083		40000		0		0				1

		24		70		299.723		40000		0		0				1

		25		29		638.986		8691.62		0		0				0

		26		24		754.887		6944.63		0		0				0

		27		21		828.079		5841.39		0		0				0

		28		19		877.817		5091.69		0		0				0

		29		17		913.291		4556.98		0		0				0

		30		16		939.519		4161.64		0		0				0

		31		15		959.466		3860.98		0		0				0

		32		15		974.987		3627.02		0		0				0

		33		14		987.3		3441.43		0		0				0

		34		14		-987.3		3441.43		0		0				0

		35		15		-974.987		3627.02		0		0				0

		36		15		-959.466		3860.98		0		0				0

		37		16		-939.519		4161.64		0		0				0

		38		17		-913.291		4556.98		0		0				0

		39		19		-877.817		5091.69		0		0				0

		40		21		-828.079		5841.39		0		0				0

		41		24		-754.887		6944.63		0		0				0

		42		29		-638.986		8691.62		0		0				0

		43		70		-299.723		40000		0		0				1

		44		70		0.00023083		40000		0		0				1

		45		69		-0.000128196		40000		0		0				1

		46		69		364.513		40000		0		0				1

		47		27		687.702		7957.32		0		0				0

		48		22		800.154		6262.31		0		0				0

		49		19		869.884		5211.26		0		0				0

		50		17		916.665		4506.13		0		0				0

		51		15		949.776		4007.04		0		0				0

		52		14		974.15		3639.64		0		0				0

		53		13		992.644		3360.88		0		0				0

		54		13		1000		3121.56		0		0				0

		55		12		1000		2916.62		0		0				0

		56		12		-1000		2916.62		0		0				0

		57		13		-1000		3121.56		0		0				0

		58		13		-992.644		3360.88		0		0				0

		59		14		-974.15		3639.64		0		0				0

		60		15		-949.776		4007.04		0		0				0

		61		17		-916.665		4506.13		0		0				0

		62		19		-869.884		5211.26		0		0				0

		63		22		-800.154		6262.31		0		0				0

		64		27		-687.702		7957.32		0		0				0

		65		69		-364.513		40000		0		0				1

		66		69		-0.000128196		40000		0		0				1

		67		68		3.66E-05		40000		0		0				1

		68		68		401.244		40000		0		0				1

		69		25		719.425		7479.15		0		0				0

		70		20		829.167		5824.99		0		0				0

		71		17		896.297		4813.14		0		0				0

		72		15		940.986		4139.52		0		0				0

		73		14		972.481		3664.79		0		0				0

		74		13		995.615		3316.09		0		0				0

		75		12		1000		3011.03		0		0				0

		76		12		1000		2769.7		0		0				0

		77		11		1000		2582.41		0		0				0

		78		11		-1000		2582.47		0		0				0

		79		12		-1000		2769.7		0		0				0

		80		12		-1000		3011.03		0		0				0

		81		13		-995.615		3316.09		0		0				0

		82		14		-972.481		3664.79		0		0				0

		83		15		-940.986		4139.52		0		0				0

		84		17		-896.297		4813.13		0		0				0

		85		20		-829.167		5824.99		0		0				0

		86		25		-719.425		7479.15		0		0				0

		87		68		-401.244		40000		0		0				1

		88		68		3.66E-05		40000		0		0				1

		89		68		-4.77E-05		40000		0		0				1

		90		68		491.264		40000		0		0				1

		91		23		742.314		7134.14		0		0				0

		92		19		849.778		5514.32		0		0				0

		93		16		914.885		4532.96		0		0				0

		94		14		958.003		3883.03		0		0				0

		95		13		988.31		3426.21		0		0				0

		96		12		1000		3058.09		0		0				0

		97		11		1000		2759.42		0		0				0

		98		11		1000		2533.68		0		0				0

		99		10		1000		2358.05		0		0				0

		100		10		-1000		2358.05		0		0				0

		101		11		-1000		2533.68		0		0				0

		102		11		-1000		2759.42		0		0				0

		103		12		-1000		3058.11		0		0				0

		104		13		-988.309		3426.22		0		0				0

		105		14		-958.003		3883.03		0		0				0

		106		16		-914.885		4532.96		0		0				0

		107		19		-849.778		5514.32		0		0				0

		108		23		-742.314		7134.14		0		0				0

		109		68		-491.264		40000		0		0				1

		110		68		-4.77E-05		40000		0		0				1

		111		67		-5.89E-05		40000		0		0				1

		112		67		611.77		40000		0		0				1

		113		21		773.745		6660.38		0		0				0

		114		17		877.639		5094.36		0		0				0

		115		14		939.786		4157.62		0		0				0

		116		13		980.676		3541.28		0		0				0

		117		12		1000		3080.18		0		0				0

		118		11		1000		2709		0		0				0

		119		10		1000		2437.33		0		0				0

		120		9		1000		2231.15		0		0				0

		121		9		1000		2070.15		0		0				0

		122		9		-1000		2070.15		0		0				0

		123		9		-1000		2231.15		0		0				0

		124		10		-1000		2437.33		0		0				0

		125		11		-1000		2709		0		0				0

		126		12		-1000		3080.18		0		0				0

		127		13		-980.676		3541.28		0		0				0

		128		14		-939.786		4157.62		0		0				0

		129		17		-877.639		5094.36		0		0				0

		130		21		-773.745		6660.38		0		0				0

		131		67		-611.77		40000		0		0				1

		132		67		-5.89E-05		40000		0		0				1

		133		65		6.40E-05		40000		0		0				1

		134		65		-199.677		40000		0		0				1

		135		20		794.676		6344.88		0		0				0

		136		16		895.918		4818.85		0		0				0

		137		13		955.987		3913.42		0		0				0

		138		12		995.355		3320.01		0		0				0

		139		11		1000		2838.86		0		0				0

		140		10		1000		2491.26		0		0				0

		141		9		1000		2236		0		0				0

		142		9		1000		2042.07		0		0				0

		143		8		1000		1890.05		0		0				0

		144		8		-1000		1890.05		0		0				0

		145		9		-1000		2042.07		0		0				0

		146		9		-1000		2236		0		0				0

		147		10		-1000		2491.26		0		0				0

		148		11		-1000		2838.86		0		0				0

		149		12		-995.355		3320.01		0		0				0

		150		13		-955.987		3913.42		0		0				0

		151		16		-895.918		4818.85		0		0				0

		152		20		-794.676		6344.88		0		0				0

		153		65		199.678		40000		0		0				1

		154		65		6.40E-05		40000		0		0				1

		155		64		5.60E-06		40000		0		0				1

		156		27		606.514		9181.08		0		0				0

		157		19		809.794		6117.01		0		0				0

		158		15		908.986		4621.86		0		0				0

		159		13		967.506		3739.79		0		0				0

		160		11		1000		3144.84		0		0				0

		161		10		1000		2672.09		0		0				0

		162		9		1000		2340.53		0		0				0

		163		9		1000		2096.89		0		0				0

		164		8		1000		1910.77		0		0				0

		165		8		1000		1765.32		0		0				0

		166		8		-1000		1765.32		0		0				0

		167		8		-1000		1910.77		0		0				0

		168		9		-1000		2096.89		0		0				0

		169		9		-1000		2340.53		0		0				0

		170		10		-1000		2672.09		0		0				0

		171		11		-1000		3144.84		0		0				0

		172		13		-967.506		3739.79		0		0				0

		173		15		-908.986		4621.86		0		0				0

		174		19		-809.794		6117.01		0		0				0

		175		27		-606.514		9181.08		0		0				0

		176		64		5.60E-06		40000		0		0				1

		177		64		-9.04E-05		40000		0		0				1

		178		25		620.744		8966.59		0		0				0

		179		18		821.301		5943.56		0		0				0

		180		14		918.86		4473.03		0		0				0

		181		12		976.176		3609.11		0		0				0

		182		10		1000		3004.2		0		0				0

		183		9		1000		2549.02		0		0				0

		184		9		1000		2229.47		0		0				0

		185		8		1000		1993.79		0		0				0

		186		8		1000		1814.17		0		0				0

		187		7		1000		1672.69		0		0				0

		188		7		-1000		1672.69		0		0				0

		189		8		-1000		1814.17		0		0				0

		190		8		-1000		1993.79		0		0				0

		191		9		-1000		2229.47		0		0				0

		192		9		-1000		2549.02		0		0				0

		193		10		-1000		3004.2		0		0				0

		194		12		-976.176		3609.11		0		0				0

		195		14		-918.86		4473.03		0		0				0

		196		18		-821.301		5943.56		0		0				0

		197		25		-620.744		8966.59		0		0				0

		198		64		-9.04E-05		40000		0		0				1

		199		63		4.94E-05		40000		0		0				1

		200		24		631.534		8803.95		0		0				0

		201		17		830.39		5806.56		0		0				0

		202		13		926.616		4356.13		0		0				0

		203		11		982.966		3506.75		0		0				0

		204		10		1000		2895.78		0		0				0

		205		9		1000		2454.19		0		0				0

		206		8		1000		2143.5		0		0				0

		207		8		1000		1914.55		0		0				0

		208		7		1000		1739.06		0		0				0

		209		7		1000		1601.7		0		0				0

		210		7		-1000		1601.7		0		0				0

		211		7		-1000		1739.06		0		0				0

		212		8		-1000		1914.55		0		0				0

		213		8		-1000		2143.5		0		0				0

		214		9		-1000		2454.18		0		0				0

		215		10		-1000		2895.78		0		0				0

		216		11		-982.966		3506.75		0		0				0

		217		13		-926.616		4356.13		0		0				0

		218		17		-830.39		5806.56		0		0				0

		219		24		-631.534		8803.94		0		0				0

		220		63		4.94E-05		40000		0		0				1

																				%%

																36				16.4

																184				83.6





40 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-----------------------		-------		---------------------		----------------

		1		66		-2.67E-05		40000		0		0				1

		2		66		318.601		40000		0		0				1

		3		66		595.885		40000		0		0				1

		4		27		628.207		8854.11		0		0				0

		5		24		707.516		7658.66		0		0				0

		6		21		763.221		6819.01		0		0				0

		7		20		803.98		6204.65		0		0				0

		8		19		834.625		5742.72		0		0				0

		9		18		858.187		5387.58		0		0				0

		10		17		876.645		5109.35		0		0				0

		11		17		891.343		4887.8		0		0				0

		12		17		-891.343		4887.8		0		0				0

		13		17		-876.645		5109.35		0		0				0

		14		18		-858.187		5387.58		0		0				0

		15		19		-834.626		5742.72		0		0				0

		16		20		-803.98		6204.65		0		0				0

		17		21		-763.221		6819.01		0		0				0

		18		24		-707.516		7658.66		0		0				0

		19		27		-628.207		8854.1		0		0				0

		20		66		-595.885		40000		0		0				1

		21		66		-318.601		40000		0		0				1

		22		66		-2.67E-05		40000		0		0				1

		23		66		2.64E-05		40000		0		0				1

		24		66		441.486		40000		0		0				1

		25		27		600.077		9278.11		0		0				0

		26		23		720.005		7470.42		0		0				0

		27		20		795.355		6334.64		0		0				0

		28		18		847.095		5554.76		0		0				0

		29		16		884.239		4994.89		0		0				0

		30		15		911.813		4579.25		0		0				0

		31		15		932.834		4262.4		0		0				0

		32		14		949.213		4015.52		0		0				0

		33		13		962.212		3819.59		0		0				0

		34		13		-962.212		3819.59		0		0				0

		35		14		-949.213		4015.52		0		0				0

		36		15		-932.834		4262.4		0		0				0

		37		15		-911.813		4579.25		0		0				0

		38		16		-884.239		4994.89		0		0				0

		39		18		-847.095		5554.76		0		0				0

		40		20		-795.355		6334.64		0		0				0

		41		23		-720.005		7470.42		0		0				0

		42		27		-600.077		9278.11		0		0				0

		43		66		-441.486		40000		0		0				1

		44		66		2.64E-05		40000		0		0				1

		45		65		-0.00011944		40000		0		0				1

		46		65		524.631		40000		0		0				1

		47		25		653.748		8469.12		0		0				0

		48		20		768.544		6738.77		0		0				0

		49		18		840.769		5650.11		0		0				0

		50		16		889.656		4913.23		0		0				0

		51		15		924.439		4388.94		0		0				0

		52		14		950.12		4001.86		0		0				0

		53		13		969.634		3707.71		0		0				0

		54		12		984.814		3478.9		0		0				0

		55		12		996.854		3297.42		0		0				0

		56		12		-996.854		3297.42		0		0				0

		57		12		-984.814		3478.9		0		0				0

		58		13		-969.634		3707.71		0		0				0

		59		14		-950.12		4001.86		0		0				0

		60		15		-924.439		4388.94		0		0				0

		61		16		-889.656		4913.23		0		0				0

		62		18		-840.769		5650.11		0		0				0

		63		20		-768.544		6738.77		0		0				0

		64		25		-653.748		8469.12		0		0				0

		65		65		-524.631		40000		0		0				1

		66		65		-0.00011944		40000		0		0				1

		67		65		2.20E-05		40000		0		0				1

		68		65		594.573		40000		0		0				1

		69		23		687.718		7957.07		0		0				0

		70		19		800.163		6262.18		0		0				0

		71		16		869.892		5211.14		0		0				0

		72		15		916.672		4506.02		0		0				0

		73		13		949.783		4006.93		0		0				0

		74		12		974.156		3639.55		0		0				0

		75		12		992.65		3360.78		0		0				0

		76		11		1000		3121.42		0		0				0

		77		11		1000		2916.55		0		0				0

		78		11		-1000		2916.55		0		0				0

		79		11		-1000		3121.42		0		0				0

		80		12		-992.65		3360.78		0		0				0

		81		12		-974.156		3639.55		0		0				0

		82		13		-949.783		4006.93		0		0				0

		83		15		-916.672		4506.02		0		0				0

		84		16		-869.892		5211.14		0		0				0

		85		19		-800.163		6262.18		0		0				0

		86		23		-687.718		7957.07		0		0				0

		87		65		-594.573		40000		0		0				1

		88		65		2.20E-05		40000		0		0				1

		89		64		-2.48E-05		40000		0		0				1

		90		64		681.226		40000		0		0				1

		91		22		712.53		7583.09		0		0				0

		92		18		822.902		5919.42		0		0				0

		93		15		890.618		4898.73		0		0				0

		94		14		935.771		4218.13		0		0				0

		95		12		967.622		3738.05		0		0				0

		96		12		991.026		3385.26		0		0				0

		97		11		1000		3089.95		0		0				0

		98		10		1000		2843.64		0		0				0

		99		10		1000		2652.78		0		0				0

		100		10		-1000		2652.78		0		0				0

		101		10		-1000		2843.64		0		0				0

		102		11		-1000		3089.95		0		0				0

		103		12		-991.026		3385.26		0		0				0

		104		12		-967.622		3738.05		0		0				0

		105		14		-935.771		4218.13		0		0				0

		106		15		-890.618		4898.73		0		0				0

		107		18		-822.902		5919.42		0		0				0

		108		22		-712.53		7583.09		0		0				0

		109		64		-681.226		40000		0		0				1

		110		64		-2.48E-05		40000		0		0				1

		111		63		6.53E-05		40000		0		0				1

		112		63		768.215		40000		0		0				1

		113		20		747.116		7061.77		0		0				0

		114		16		854.066		5449.68		0		0				0

		115		14		918.734		4474.94		0		0				0

		116		12		961.517		3830.06		0		0				0

		117		11		991.572		3377.03		0		0				0

		118		10		1000		3002.94		0		0				0

		119		10		1000		2708.58		0		0				0

		120		9		1000		2485.83		0		0				0

		121		9		1000		2312.75		0		0				0

		122		9		-1000		2312.75		0		0				0

		123		9		-1000		2485.83		0		0				0

		124		10		-1000		2708.58		0		0				0

		125		10		-1000		3002.94		0		0				0

		126		11		-991.572		3377.03		0		0				0

		127		12		-961.517		3830.06		0		0				0

		128		14		-918.734		4474.94		0		0				0

		129		16		-854.066		5449.68		0		0				0

		130		20		-747.116		7061.77		0		0				0

		131		63		-768.215		40000		0		0				1

		132		63		6.53E-05		40000		0		0				1

		133		62		2.63E-05		40000		0		0				1

		134		63		889.463		40000		0		0				1

		135		19		770.541		6708.68		0		0				0

		136		15		874.822		5136.83		0		0				0

		137		13		937.279		4195.4		0		0				0

		138		11		978.4		3575.58		0		0				0

		139		10		1000		3118.23		0		0				0

		140		10		1000		2743.33		0		0				0

		141		9		1000		2469.02		0		0				0

		142		8		1000		2260.95		0		0				0

		143		8		1000		2098.44		0		0				0

		144		8		-1000		2098.44		0		0				0

		145		8		-1000		2260.95		0		0				0

		146		9		-1000		2469.02		0		0				0

		147		10		-1000		2743.33		0		0				0

		148		10		-1000		3118.23		0		0				0

		149		11		-978.4		3575.58		0		0				0

		150		13		-937.279		4195.4		0		0				0

		151		15		-874.822		5136.83		0		0				0

		152		19		-770.541		6708.67		0		0				0

		153		63		-889.463		40000		0		0				1

		154		62		2.63E-05		40000		0		0				1

		155		62		-8.23E-05		40000		0		0				1

		156		62		1116.39		40000		0		0				1

		157		18		787.687		6450.23		0		0				0

		158		14		889.839		4910.48		0		0				0

		159		12		950.61		3994.46		0		0				0

		160		11		990.49		3393.35		0		0				0

		161		10		1000		2918.05		0		0				0

		162		9		1000		2562.75		0		0				0

		163		8		1000		2302.09		0		0				0

		164		8		1000		2104.24		0		0				0

		165		8		1000		1949.4		0		0				0

		166		8		-1000		1949.4		0		0				0

		167		8		-1000		2104.24		0		0				0

		168		8		-1000		2302.09		0		0				0

		169		9		-1000		2562.75		0		0				0

		170		10		-1000		2918.05		0		0				0

		171		11		-990.49		3393.35		0		0				0

		172		12		-950.61		3994.46		0		0				0

		173		14		-889.839		4910.48		0		0				0

		174		18		-787.687		6450.23		0		0				0

		175		62		-1116.39		40000		0		0				1

		176		62		-8.23E-05		40000		0		0				1

		177		61		0.000112214		40000		0		0				1

		178		24		595.218		9351.35		0		0				0

		179		17		800.883		6251.32		0		0				0

		180		13		901.297		4737.77		0		0				0

		181		11		960.734		3841.86		0		0				0

		182		10		999.648		3255.31		0		0				0

		183		9		1000		2769.59		0		0				0

		184		8		1000		2428.72		0		0				0

		185		8		1000		2178.41		0		0				0

		186		8		1000		1987.75		0		0				0

		187		7		1000		1838.06		0		0				0

		188		7		-1000		1838.06		0		0				0

		189		8		-1000		1987.75		0		0				0

		190		8		-1000		2178.41		0		0				0

		191		8		-1000		2428.72		0		0				0

		192		9		-1000		2769.59		0		0				0

		193		10		-999.648		3255.31		0		0				0

		194		11		-960.734		3841.86		0		0				0

		195		13		-901.297		4737.77		0		0				0

		196		17		-800.883		6251.32		0		0				0

		197		24		-595.218		9351.35		0		0				0

		198		61		0.000112214		40000		0		0				1

		199		60		-0.000134958		40000		0		0				1

		200		23		608.934		9144.61		0		0				0

		201		16		811.405		6092.73		0		0				0

		202		13		910.372		4600.97		0		0				0

		203		11		968.725		3721.41		0		0				0

		204		10		1000		3124.94		0		0				0

		205		9		1000		2654.66		0		0				0

		206		8		1000		2324.74		0		0				0

		207		8		1000		2082.33		0		0				0

		208		7		1000		1896.96		0		0				0

		209		7		1000		1752.29		0		0				0

		210		7		-1000		1752.29		0		0				0

		211		7		-1000		1896.96		0		0				0

		212		8		-1000		2082.33		0		0				0

		213		8		-1000		2324.8		0		0				0

		214		9		-1000		2654.66		0		0				0

		215		10		-1000		3124.94		0		0				0

		216		11		-968.725		3721.41		0		0				0

		217		13		-910.372		4600.97		0		0				0

		218		16		-811.405		6092.73		0		0				0

		219		23		-608.934		9144.61		0		0				0

		220		60		-0.000134958		40000		0		0				1

																				%%

																38				17.3

																182				82.7





50 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		----------------------		-------		---------------------		----------------

		1		63		0.000208371		40000		0		0				1

		2		63		370.621		40000		0		0				1

		3		63		733.88		40000		0		0				1

		4		25		594.86		9356.74		0		0				0

		5		22		677.09		8117.28		0		0				0

		6		20		733.685		7264.21		0		0				0

		7		19		775.376		6635.8		0		0				0

		8		18		806.876		6160.99		0		0				0

		9		17		831.178		5794.69		0		0				0

		10		16		850.262		5507.03		0		0				0

		11		16		865.484		5277.59		0		0				0

		12		16		-865.484		5277.59		0		0				0

		13		16		-850.262		5507.03		0		0				0

		14		17		-831.178		5794.69		0		0				0

		15		18		-806.876		6160.99		0		0				0

		16		19		-775.376		6635.8		0		0				0

		17		20		-733.685		7264.21		0		0				0

		18		22		-677.09		8117.28		0		0				0

		19		25		-594.86		9356.74		0		0				0

		20		63		-733.88		40000		0		0				1

		21		63		-370.621		40000		0		0				1

		22		63		0.000208371		40000		0		0				1

		23		63		0.00010387		40000		0		0				1

		24		63		505.194		40000		0		0				1

		25		63		856.058		40000		0		0				1

		26		22		691.592		7898.68		0		0				0

		27		19		768.369		6741.42		0		0				0

		28		17		821.547		5939.85		0		0				0

		29		16		859.94		5361.14		0		0				0

		30		15		888.547		4929.95		0		0				0

		31		14		910.406		4600.47		0		0				0

		32		13		927.461		4343.39		0		0				0

		33		13		941.008		4139.2		0		0				0

		34		13		-941.008		4139.2		0		0				0

		35		13		-927.461		4343.39		0		0				0

		36		14		-910.406		4600.47		0		0				0

		37		15		-888.547		4929.95		0		0				0

		38		16		-859.94		5361.14		0		0				0

		39		17		-821.547		5939.85		0		0				0

		40		19		-768.369		6741.42		0		0				0

		41		22		-691.592		7898.68		0		0				0

		42		63		-856.058		40000		0		0				1

		43		63		-505.194		40000		0		0				1

		44		63		0.00010387		40000		0		0				1

		45		62		4.12E-05		40000		0		0				1

		46		62		610.635		40000		0		0				1

		47		24		625.951		8888.1		0		0				0

		48		20		742.378		7133.19		0		0				0

		49		17		816.394		6017.52		0		0				0

		50		15		866.879		5256.56		0		0				0

		51		14		902.968		4712.58		0		0				0

		52		13		929.688		4309.82		0		0				0

		53		12		950.025		4003.28		0		0				0

		54		12		965.857		3764.65		0		0				0

		55		12		978.417		3575.33		0		0				0

		56		12		-978.417		3575.32		0		0				0

		57		12		-965.857		3764.65		0		0				0

		58		12		-950.025		4003.28		0		0				0

		59		13		-929.688		4309.82		0		0				0

		60		14		-902.968		4712.58		0		0				0

		61		15		-866.879		5256.56		0		0				0

		62		17		-816.394		6017.52		0		0				0

		63		20		-742.378		7133.19		0		0				0

		64		24		-625.951		8888.1		0		0				0

		65		62		-610.636		40000		0		0				1

		66		62		4.12E-05		40000		0		0				1

		67		62		-1.36E-04		40000		0		0				1

		68		62		699.716		40000		0		0				1

		69		22		661.58		8351.07		0		0				0

		70		18		775.864		6628.43		0		0				0

		71		16		847.543		5548.01		0		0				0

		72		14		895.959		4818.23		0		0				0

		73		13		930.363		4299.64		0		0				0

		74		12		955.747		3917.04		0		0				0

		75		11		975.027		3626.42		0		0				0

		76		11		990.024		3400.38		0		0				0

		77		11		1000		3214.75		0		0				0

		78		11		-1000		3214.75		0		0				0

		79		11		-990.023		3400.39		0		0				0

		80		11		-975.027		3626.42		0		0				0

		81		12		-955.747		3917.04		0		0				0

		82		13		-930.363		4299.64		0		0				0

		83		14		-895.959		4818.23		0		0				0

		84		16		-847.543		5548.01		0		0				0

		85		18		-775.864		6628.43		0		0				0

		86		22		-661.58		8351.07		0		0				0

		87		62		-699.716		40000		0		0				1

		88		62		-1.36E-04		40000		0		0				1

		89		62		0.000180352		40000		0		0				1

		90		62		751.552		40000		0		0				1

		91		21		687.737		7956.79		0		0				0

		92		17		800.172		6262.04		0		0				0

		93		15		869.9		5211.01		0		0				0

		94		13		916.68		4505.9		0		0				0

		95		12		949.79		4006.82		0		0				0

		96		11		974.164		3639.43		0		0				0

		97		11		992.658		3360.67		0		0				0

		98		10		1000		3121.29		0		0				0

		99		10		1000		2916.42		0		0				0

		100		10		-1000		2916.42		0		0				0

		101		10		-1000		3121.29		0		0				0

		102		11		-992.658		3360.67		0		0				0

		103		11		-974.164		3639.43		0		0				0

		104		12		-949.79		4006.82		0		0				0

		105		13		-916.68		4505.9		0		0				0

		106		15		-869.9		5211.01		0		0				0

		107		17		-800.172		6262.04		0		0				0

		108		21		-687.737		7956.79		0		0				0

		109		62		-751.552		40000		0		0				1

		110		62		0.000180352		40000		0		0				1

		111		61		7.21E-05		40000		0		0				1

		112		61		902.732		40000		0		0				1

		113		19		724.583		7401.41		0		0				0

		114		16		833.833		5754.66		0		0				0

		115		13		900.517		4749.52		0		0				0

		116		12		944.857		4081.18		0		0				0

		117		11		976.086		3610.46		0		0				0

		118		10		999.02		3264.77		0		0				0

		119		10		1000		2953.04		0		0				0

		120		9		1000		2715.19		0		0				0

		121		9		1000		2530.77		0		0				0

		122		9		-1000		2530.77		0		0				0

		123		9		-1000		2715.19		0		0				0

		124		10		-1000		2953.04		0		0				0

		125		10		-999.02		3264.77		0		0				0

		126		11		-976.086		3610.46		0		0				0

		127		12		-944.857		4081.18		0		0				0

		128		13		-900.517		4749.52		0		0				0

		129		16		-833.833		5754.66		0		0				0

		130		19		-724.583		7401.41		0		0				0

		131		61		-902.732		40000		0		0				1

		132		61		7.21E-05		40000		0		0				1

		133		60		-0.000238309		40000		0		0				1

		134		60		970.525		40000		0		0				1

		135		18		749.856		7020.47		0		0				0

		136		15		856.508		5412.88		0		0				0

		137		12		920.923		4441.94		0		0				0

		138		11		963.514		3799.96		0		0				0

		139		10		993.426		3349.09		0		0				0

		140		9		1000		2971.75		0		0				0

		141		9		1000		2679.81		0		0				0

		142		8		1000		2458.82		0		0				0

		143		8		1000		2286.83		0		0				0

		144		8		-1000		2286.83		0		0				0

		145		8		-1000		2458.82		0		0				0

		146		9		-1000		2679.81		0		0				0

		147		9		-1000		2971.75		0		0				0

		148		10		-993.426		3349.09		0		0				0

		149		11		-963.514		3799.96		0		0				0

		150		12		-920.923		4441.94		0		0				0

		151		15		-856.508		5412.88		0		0				0

		152		18		-749.856		7020.47		0		0				0

		153		60		-970.525		40000		0		0				1

		154		60		-0.000238309		40000		0		0				1

		155		59		7.38E-05		40000		0		0				1

		156		59		1124.5		40000		0		0				1

		157		17		768.542		6738.81		0		0				0

		158		14		873.061		5163.38		0		0				0

		159		12		935.711		4219.04		0		0				0

		160		10		976.975		3597.05		0		0				0

		161		10		1000		3142.17		0		0				0

		162		9		1000		2764.9		0		0				0

		163		8		1000		2488.77		0		0				0

		164		8		1000		2279.64		0		0				0

		165		7		1000		2116.51		0		0				0

		166		7		-1000		2116.51		0		0				0

		167		8		-1000		2279.64		0		0				0

		168		8		-1000		2488.77		0		0				0

		169		9		-1000		2764.89		0		0				0

		170		10		-1000		3142.17		0		0				0

		171		10		-976.975		3597.05		0		0				0

		172		12		-935.711		4219.04		0		0				0

		173		14		-873.061		5163.38		0		0				0

		174		17		-768.542		6738.81		0		0				0

		175		59		-1124.5		40000		0		0				1

		176		59		7.38E-05		40000		0		0				1

		177		59		9.66E-06		40000		0		0				1

		178		59		1237.02		40000		0		0				1

		179		16		783.044		6520.22		0		0				0

		180		13		885.786		4971.57		0		0				0

		181		11		947.02		4048.58		0		0				0

		182		10		987.237		3442.38		0		0				0

		183		9		1000		2971.38		0		0				0

		184		8		1000		2610.82		0		0				0

		185		8		1000		2346.68		0		0				0

		186		7		1000		2145.88		0		0				0

		187		7		1000		1988.86		0		0				0

		188		7		-1000		1988.86		0		0				0

		189		7		-1000		2145.88		0		0				0

		190		8		-1000		2346.68		0		0				0

		191		8		-1000		2610.82		0		0				0

		192		9		-1000		2971.42		0		0				0

		193		10		-987.237		3442.38		0		0				0

		194		11		-947.02		4048.58		0		0				0

		195		13		-885.786		4971.57		0		0				0

		196		16		-783.044		6520.22		0		0				0

		197		59		-1237.02		40000		0		0				1

		198		59		9.66E-06		40000		0		0				1

		199		58		5.23E-05		40000		0		0				1

		200		59		226.169		40000		0		0				1

		201		16		794.69		6344.67		0		0				0

		202		13		895.93		4818.67		0		0				0

		203		11		955.997		3913.26		0		0				0

		204		10		995.365		3319.87		0		0				0

		205		9		1000		2838.72		0		0				0

		206		8		1000		2491.14		0		0				0

		207		7		1000		2235.93		0		0				0

		208		7		1000		2041.6		0		0				0

		209		7		1000		1890.1		0		0				0

		210		7		-1000		1890.1		0		0				0

		211		7		-1000		2041.6		0		0				0

		212		7		-1000		2235.93		0		0				0

		213		8		-1000		2491.14		0		0				0

		214		9		-1000		2838.72		0		0				0

		215		10		-995.365		3319.87		0		0				0

		216		11		-955.997		3913.26		0		0				0

		217		13		-895.93		4818.67		0		0				0

		218		16		-794.69		6344.67		0		0				0

		219		59		-226.166		40000		0		0				1

		220		58		5.23E-05		40000		0		0				1

																				%%

																44				20.0

																176				80.0





60 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-----------------------		-------		---------------------		----------------

		1		61		-6.70E-05		40000		0		0				1

		2		61		395.314		40000		0		0				1

		3		61		783.189		40000		0		0				1

		4		61		1121.22		40000		0		0				1

		5		22		651.652		8500.71		0		0				0

		6		20		708.814		7639.09		0		0				0

		7		18		751.149		7000.97		0		0				0

		8		17		783.272		6516.78		0		0				0

		9		17		808.128		6142.11		0		0				0

		10		16		827.691		5847.25		0		0				0

		11		15		843.318		5611.69		0		0				0

		12		15		-843.318		5611.69		0		0				0

		13		16		-827.691		5847.25		0		0				0

		14		17		-808.128		6142.11		0		0				0

		15		17		-783.272		6516.78		0		0				0

		16		18		-751.149		7000.97		0		0				0

		17		20		-708.814		7639.09		0		0				0

		18		22		-651.652		8500.7		0		0				0

		19		61		-1121.22		40000		0		0				1

		20		61		-783.189		40000		0		0				1

		21		61		-395.321		40000		0		0				1

		22		61		-6.70E-05		40000		0		0				1

		23		61		-0.000141174		40000		0		0				1

		24		61		540.736		40000		0		0				1

		25		61		1037.27		40000		0		0				1

		26		21		667.628		8259.9		0		0				0

		27		18		745.374		7088.02		0		0				0

		28		16		799.624		6270.3		0		0				0

		29		15		838.987		5676.98		0		0				0

		30		14		868.413		5233.44		0		0				0

		31		14		890.947		4893.78		0		0				0

		32		13		908.554		4628.38		0		0				0

		33		13		922.551		4417.41		0		0				0

		34		13		-922.551		4417.4		0		0				0

		35		13		-908.554		4628.38		0		0				0

		36		14		-890.947		4893.78		0		0				0

		37		14		-868.413		5233.44		0		0				0

		38		15		-838.987		5676.98		0		0				0

		39		16		-799.624		6270.3		0		0				0

		40		18		-745.374		7088.02		0		0				0

		41		21		-667.628		8259.9		0		0				0

		42		61		-1037.27		40000		0		0				1

		43		61		-540.736		40000		0		0				1

		44		61		-0.000141174		40000		0		0				1

		45		60		-7.34E-05		40000		0		0				1

		46		60		648.212		40000		0		0				1

		47		23		601.978		9249.45		0		0				0

		48		19		720.031		7470.03		0		0				0

		49		16		795.377		6334.31		0		0				0

		50		15		847.116		5554.45		0		0				0

		51		14		884.258		4994.6		0		0				0

		52		13		911.832		4578.98		0		0				0

		53		12		932.853		4262.12		0		0				0

		54		12		949.231		4015.25		0		0				0

		55		11		962.23		3819.31		0		0				0

		56		11		-962.23		3819.31		0		0				0

		57		12		-949.231		4015.25		0		0				0

		58		12		-932.853		4262.12		0		0				0

		59		13		-911.832		4578.98		0		0				0

		60		14		-884.258		4994.6		0		0				0

		61		15		-847.116		5554.45		0		0				0

		62		16		-795.377		6334.31		0		0				0

		63		19		-720.031		7470.03		0		0				0

		64		23		-601.978		9249.45		0		0				0

		65		60		-648.212		40000		0		0				1

		66		60		-7.34E-05		40000		0		0				1

		67		60		-0.000260377		40000		0		0				1

		68		60		744.323		40000		0		0				1

		69		21		639.319		8686.6		0		0				0

		70		17		754.919		6944.15		0		0				0

		71		15		828.107		5840.97		0		0				0

		72		14		877.843		5091.3		0		0				0

		73		13		913.316		4556.61		0		0				0

		74		12		939.543		4161.28		0		0				0

		75		11		959.49		3860.62		0		0				0

		76		11		975.01		3626.68		0		0				0

		77		10		987.323		3441.09		0		0				0

		78		10		-987.323		3441.09		0		0				0

		79		11		-975.01		3626.68		0		0				0

		80		11		-959.49		3860.62		0		0				0

		81		12		-939.543		4161.28		0		0				0

		82		13		-913.316		4556.61		0		0				0

		83		14		-877.843		5091.3		0		0				0

		84		15		-828.107		5840.97		0		0				0

		85		17		-754.919		6944.15		0		0				0

		86		21		-639.319		8686.6		0		0				0

		87		60		-744.323		40000		0		0				1

		88		60		-0.000260377		40000		0		0				1

		89		59		0.000120677		40000		0		0				1

		90		60		820.95		40000		0		0				1

		91		20		666.486		8277.12		0		0				0

		92		16		780.434		6559.56		0		0				0

		93		14		851.761		5484.43		0		0				0

		94		13		899.878		4759.16		0		0				0

		95		12		934.044		4244.17		0		0				0

		96		11		959.24		3864.38		0		0				0

		97		10		978.374		3575.98		0		0				0

		98		10		993.256		3351.66		0		0				0

		99		10		1000		3157.29		0		0				0

		100		10		-1000		3157.29		0		0				0

		101		10		-993.256		3351.66		0		0				0

		102		10		-978.374		3575.98		0		0				0

		103		11		-959.24		3864.38		0		0				0

		104		12		-934.044		4244.17		0		0				0

		105		13		-899.878		4759.16		0		0				0

		106		14		-851.761		5484.43		0		0				0

		107		16		-780.434		6559.56		0		0				0

		108		20		-666.486		8277.12		0		0				0

		109		60		-820.95		40000		0		0				1

		110		59		0.000120677		40000		0		0				1

		111		59		1.76E-06		40000		0		0				1

		112		59		910.948		40000		0		0				1

		113		18		705.034		7696.08		0		0				0

		114		15		816.061		6022.55		0		0				0

		115		13		884.401		4992.45		0		0				0

		116		12		930.053		4304.32		0		0				0

		117		11		962.288		3818.45		0		0				0

		118		10		985.988		3461.21		0		0				0

		119		9		1000		3177.4		0		0				0

		120		9		1000		2925.68		0		0				0

		121		9		1000		2730.62		0		0				0

		122		9		-1000		2730.62		0		0				0

		123		9		-1000		2925.68		0		0				0

		124		9		-1000		3177.4		0		0				0

		125		10		-985.987		3461.22		0		0				0

		126		11		-962.288		3818.45		0		0				0

		127		12		-930.053		4304.32		0		0				0

		128		13		-884.401		4992.45		0		0				0

		129		15		-816.061		6022.55		0		0				0

		130		18		-705.034		7696.08		0		0				0

		131		59		-910.948		40000		0		0				1

		132		59		1.76E-06		40000		0		0				1

		133		58		-2.86E-05		40000		0		0				1

		134		58		1065.33		40000		0		0				1

		135		17		731.731		7293.67		0		0				0

		136		14		840.279		5657.51		0		0				0

		137		12		906.335		4661.83		0		0				0

		138		11		950.186		4000.85		0		0				0

		139		10		981.045		3535.72		0		0				0

		140		9		1000		3182.23		0		0				0

		141		9		1000		2873.93		0		0				0

		142		8		1000		2640.97		0		0				0

		143		8		1000		2460.28		0		0				0

		144		8		-1000		2460.28		0		0				0

		145		8		-1000		2640.97		0		0				0

		146		9		-1000		2873.93		0		0				0

		147		9		-1000		3182.23		0		0				0

		148		10		-981.045		3535.72		0		0				0

		149		11		-950.186		4000.85		0		0				0

		150		12		-906.335		4661.83		0		0				0

		151		14		-840.279		5657.51		0		0				0

		152		17		-731.731		7293.67		0		0				0

		153		58		-1065.33		40000		0		0				1

		154		58		-2.86E-05		40000		0		0				1

		155		58		-9.11E-05		40000		0		0				1

		156		58		1156.65		40000		0		0				1

		157		16		751.629		6993.74		0		0				0

		158		13		858.085		5389.1		0		0				0

		159		11		922.336		4420.65		0		0				0

		160		10		964.803		3780.54		0		0				0

		161		9		994.621		3331.07		0		0				0

		162		9		1000		2951.73		0		0				0

		163		8		1000		2661.2		0		0				0

		164		8		1000		2441.55		0		0				0

		165		7		1000		2270.43		0		0				0

		166		7		-1000		2270.43		0		0				0

		167		8		-1000		2441.55		0		0				0

		168		8		-1000		2661.2		0		0				0

		169		9		-1000		2951.73		0		0				0

		170		9		-994.621		3331.07		0		0				0

		171		10		-964.803		3780.54		0		0				0

		172		11		-922.336		4420.64		0		0				0

		173		13		-858.085		5389.1		0		0				0

		174		16		-751.629		6993.74		0		0				0

		175		58		-1156.65		40000		0		0				1

		176		58		-9.11E-05		40000		0		0				1

		177		57		0.000237682		40000		0		0				1

		178		57		1078.82		40000		0		0				1

		179		16		767.176		6759.39		0		0				0

		180		13		871.856		5181.53		0		0				0

		181		11		934.638		4235.21		0		0				0

		182		10		976		3611.76		0		0				0

		183		9		1000		3158.58		0		0				0

		184		8		1000		2779.6		0		0				0

		185		8		1000		2502.56		0		0				0

		186		7		1000		2292.29		0		0				0

		187		7		1000		2128.35		0		0				0

		188		7		-1000		2128.35		0		0				0

		189		7		-1000		2292.29		0		0				0

		190		8		-1000		2502.56		0		0				0

		191		8		-1000		2779.6		0		0				0

		192		9		-1000		3158.58		0		0				0

		193		10		-976		3611.75		0		0				0

		194		11		-934.638		4235.21		0		0				0

		195		13		-871.856		5181.53		0		0				0

		196		16		-767.176		6759.39		0		0				0

		197		57		-1078.82		40000		0		0				1

		198		57		0.000237682		40000		0		0				1

		199		57		4.25E-05		40000		0		0				1

		200		57		1362.48		40000		0		0				1

		201		15		779.734		6570.11		0		0				0

		202		12		882.89		5015.21		0		0				0

		203		10		944.451		4087.3		0		0				0

		204		9		984.909		3477.46		0		0				0

		205		8		1000		3009.86		0		0				0

		206		8		1000		2645.58		0		0				0

		207		7		1000		2378.57		0		0				0

		208		7		1000		2176.23		0		0				0

		209		7		1000		2017.91		0		0				0

		210		7		-1000		2017.91		0		0				0

		211		7		-1000		2176.23		0		0				0

		212		7		-1000		2378.57		0		0				0

		213		8		-1000		2645.58		0		0				0

		214		8		-1000		3009.86		0		0				0

		215		9		-984.909		3477.46		0		0				0

		216		10		-944.451		4087.3		0		0				0

		217		12		-882.89		5015.21		0		0				0

		218		15		-779.734		6570.11		0		0				0

		219		57		-1362.48		40000		0		0				1

		220		57		4.25E-05		40000		0		0				1

																				%%

																46				20.9

																174				79.1





70 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		----------------------		-------		---------------------		----------------

		1		59		0.000122626		40000		0		0				1

		2		59		405.146		40000		0		0				1

		3		59		794.823		40000		0		0				1

		4		59		1134.27		40000		0		0				1

		5		21		629.788		8830.27		0		0				0

		6		19		687.331		7962.91		0		0				0

		7		18		730.12		7317.94		0		0				0

		8		17		762.706		6826.77		0		0				0

		9		16		787.989		6445.67		0		0				0

		10		15		807.926		6145.16		0		0				0

		11		15		823.876		5904.75		0		0				0

		12		15		-823.876		5904.75		0		0				0

		13		15		-807.926		6145.16		0		0				0

		14		16		-787.989		6445.67		0		0				0

		15		17		-762.706		6826.77		0		0				0

		16		18		-730.12		7317.94		0		0				0

		17		19		-687.332		7962.9		0		0				0

		18		21		-629.788		8830.27		0		0				0

		19		59		-1134.27		40000		0		0				1

		20		59		-794.823		40000		0		0				1

		21		59		-405.146		40000		0		0				1

		22		59		0.000122626		40000		0		0				1

		23		59		4.41E-06		40000		0		0				1

		24		59		560.676		40000		0		0				1

		25		59		1121.59		40000		0		0				1

		26		20		646.917		8572.08		0		0				0

		27		18		725.331		7390.14		0		0				0

		28		16		780.397		6560.11		0		0				0

		29		15		820.53		5955.18		0		0				0

		30		14		850.623		5501.59		0		0				0

		31		13		873.714		5153.52		0		0				0

		32		13		891.782		4881.19		0		0				0

		33		12		906.156		4664.52		0		0				0

		34		12		-906.156		4664.52		0		0				0

		35		13		-891.782		4881.19		0		0				0

		36		13		-873.714		5153.52		0		0				0

		37		14		-850.623		5501.59		0		0				0

		38		15		-820.53		5955.18		0		0				0

		39		16		-780.397		6560.11		0		0				0

		40		18		-725.331		7390.14		0		0				0

		41		20		-646.917		8572.08		0		0				0

		42		59		-1121.59		40000		0		0				1

		43		59		-560.676		40000		0		0				1

		44		59		4.41E-06		40000		0		0				1

		45		58		-6.45E-06		40000		0		0				1

		46		58		673.161		40000		0		0				1

		47		58		1371.18		40000		0		0				1

		48		18		700.525		7764.04		0		0				0

		49		16		776.881		6613.11		0		0				0

		50		14		829.627		5818.07		0		0				0

		51		13		867.638		5245.11		0		0				0

		52		12		895.927		4818.71		0		0				0

		53		12		917.527		4493.13		0		0				0

		54		11		934.373		4239.21		0		0				0

		55		11		947.749		4037.59		0		0				0

		56		11		-947.749		4037.59		0		0				0

		57		11		-934.373		4239.21		0		0				0

		58		12		-917.527		4493.13		0		0				0

		59		12		-895.927		4818.7		0		0				0

		60		13		-867.638		5245.11		0		0				0

		61		14		-829.627		5818.07		0		0				0

		62		16		-776.881		6613.11		0		0				0

		63		18		-700.525		7764.04		0		0				0

		64		58		-1371.18		40000		0		0				1

		65		58		-673.161		40000		0		0				1

		66		58		-6.45E-06		40000		0		0				1

		67		58		4.15E-05		40000		0		0				1

		68		58		759.035		40000		0		0				1

		69		20		619.945		8978.64		0		0				0

		70		17		736.5		7221.78		0		0				0

		71		15		810.884		6100.58		0		0				0

		72		13		861.708		5334.49		0		0				0

		73		12		898.08		4786.25		0		0				0

		74		12		925.029		4380.06		0		0				0

		75		11		945.548		4070.77		0		0				0

		76		10		961.524		3829.95		0		0				0

		77		10		974.202		3638.86		0		0				0

		78		10		-974.202		3638.86		0		0				0

		79		10		-961.524		3829.95		0		0				0

		80		11		-945.548		4070.77		0		0				0

		81		12		-925.029		4380.06		0		0				0

		82		12		-898.08		4786.25		0		0				0

		83		13		-861.708		5334.49		0		0				0

		84		15		-810.884		6100.58		0		0				0

		85		17		-736.5		7221.78		0		0				0

		86		20		-619.945		8978.64		0		0				0

		87		58		-759.035		40000		0		0				1

		88		58		4.15E-05		40000		0		0				1

		89		58		5.33E-05		40000		0		0				1

		90		58		830.172		40000		0		0				1

		91		19		647.909		8557.13		0		0				0

		92		16		762.972		6822.76		0		0				0

		93		14		835.599		5728.04		0		0				0

		94		13		884.837		4985.86		0		0				0

		95		12		919.905		4457.28		0		0				0

		96		11		945.811		4066.8		0		0				0

		97		10		965.503		3769.97		0		0				0

		98		10		980.823		3539.06		0		0				0

		99		10		992.974		3355.9		0		0				0

		100		10		-992.974		3355.9		0		0				0

		101		10		-980.823		3539.06		0		0				0

		102		10		-965.503		3769.97		0		0				0

		103		11		-945.811		4066.8		0		0				0

		104		12		-919.905		4457.28		0		0				0

		105		13		-884.838		4985.86		0		0				0

		106		14		-835.599		5728.04		0		0				0

		107		16		-762.972		6822.76		0		0				0

		108		19		-647.909		8557.13		0		0				0

		109		58		-830.172		40000		0		0				1

		110		58		5.33E-05		40000		0		0				1

		111		57		7.18E-05		40000		0		0				1

		112		57		976.387		40000		0		0				1

		113		18		687.768		7956.32		0		0				0

		114		15		800.194		6261.71		0		0				0

		115		13		869.92		5210.72		0		0				0

		116		11		916.698		4505.63		0		0				0

		117		10		949.807		4006.57		0		0				0

		118		10		974.179		3639.2		0		0				0

		119		9		992.674		3360.42		0		0				0

		120		9		1000		3121.05		0		0				0

		121		9		1000		2916.17		0		0				0

		122		9		-1000		2916.17		0		0				0

		123		9		-1000		3121.05		0		0				0

		124		9		-992.674		3360.42		0		0				0

		125		10		-974.179		3639.2		0		0				0

		126		10		-949.807		4006.57		0		0				0

		127		11		-916.698		4505.63		0		0				0

		128		13		-869.92		5210.72		0		0				0

		129		15		-800.194		6261.71		0		0				0

		130		18		-687.768		7956.32		0		0				0

		131		57		-976.387		40000		0		0				1

		132		57		7.18E-05		40000		0		0				1

		133		57		0.000364478		40000		0		0				1

		134		57		1085.17		40000		0		0				1

		135		17		715.596		7536.87		0		0				0

		136		14		825.687		5877.45		0		0				0

		137		12		893.143		4860.67		0		0				0

		138		11		938.091		4183.17		0		0				0

		139		10		969.784		3705.45		0		0				0

		140		9		993.07		3354.45		0		0				0

		141		9		1000		3054.74		0		0				0

		142		8		1000		2810.55		0		0				0

		143		8		1000		2621.4		0		0				0

		144		8		-1000		2621.4		0		0				0

		145		8		-1000		2810.55		0		0				0

		146		9		-1000		3054.74		0		0				0

		147		9		-993.07		3354.45		0		0				0

		148		10		-969.784		3705.45		0		0				0

		149		11		-938.091		4183.17		0		0				0

		150		12		-893.143		4860.67		0		0				0

		151		14		-825.687		5877.45		0		0				0

		152		17		-715.596		7536.87		0		0				0

		153		57		-1085.17		40000		0		0				1

		154		57		0.000364478		40000		0		0				1

		155		56		-0.000203981		40000		0		0				1

		156		56		1119.75		40000		0		0				1

		157		16		736.475		7222.16		0		0				0

		158		13		844.543		5593.23		0		0				0

		159		11		910.177		4603.92		0		0				0

		160		10		953.7		3947.88		0		0				0

		161		9		984.312		3486.46		0		0				0

		162		9		1000		3126.23		0		0				0

		163		8		1000		2822.17		0		0				0

		164		8		1000		2592.61		0		0				0

		165		7		1000		2414.02		0		0				0

		166		7		-1000		2414.02		0		0				0

		167		8		-1000		2592.61		0		0				0

		168		8		-1000		2822.17		0		0				0

		169		9		-1000		3126.23		0		0				0

		170		9		-984.312		3486.46		0		0				0

		171		10		-953.7		3947.89		0		0				0

		172		11		-910.177		4603.92		0		0				0

		173		13		-844.543		5593.23		0		0				0

		174		16		-736.475		7222.16		0		0				0

		175		56		-1119.75		40000		0		0				1

		176		56		-0.000203981		40000		0		0				1

		177		56		8.24E-05		40000		0		0				1

		178		56		1263.23		40000		0		0				1

		179		15		752.879		6974.9		0		0				0

		180		12		859.198		5372.34		0		0				0

		181		11		923.332		4405.64		0		0				0

		182		10		965.709		3766.87		0		0				0

		183		9		995.461		3318.41		0		0				0

		184		8		1000		2937.59		0		0				0

		185		8		1000		2648.36		0		0				0

		186		7		1000		2429.19		0		0				0

		187		7		1000		2259.06		0		0				0

		188		7		-1000		2259.06		0		0				0

		189		7		-1000		2429.19		0		0				0

		190		8		-1000		2648.36		0		0				0

		191		8		-1000		2937.59		0		0				0

		192		9		-995.461		3318.41		0		0				0

		193		10		-965.709		3766.87		0		0				0

		194		11		-923.332		4405.64		0		0				0

		195		12		-859.198		5372.34		0		0				0

		196		15		-752.879		6974.9		0		0				0

		197		56		-1263.23		40000		0		0				1

		198		56		8.24E-05		40000		0		0				1

		199		55		0.000153606		40000		0		0				1

		200		55		1319.92		40000		0		0				1

		201		15		766.193		6774.22		0		0				0

		202		12		870.989		5194.6		0		0				0

		203		10		933.866		4246.85		0		0				0

		204		9		975.299		3622.33		0		0				0

		205		8		1000		3170.4		0		0				0

		206		8		1000		2790.39		0		0				0

		207		7		1000		2512.28		0		0				0

		208		7		1000		2301.81		0		0				0

		209		7		1000		2137.55		0		0				0

		210		7		-1000		2137.55		0		0				0

		211		7		-1000		2301.81		0		0				0

		212		7		-1000		2512.28		0		0				0

		213		8		-1000		2790.39		0		0				0

		214		8		-1000		3170.4		0		0				0

		215		9		-975.299		3622.33		0		0				0

		216		10		-933.866		4246.85		0		0				0

		217		12		-870.989		5194.6		0		0				0

		218		15		-766.193		6774.22		0		0				0

		219		55		-1319.92		40000		0		0				1

		220		55		0.000153606		40000		0		0				1

																				%%

																48				21.8

																172				78.2





80 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		----------------------		-------		---------------------		----------------

		1		58		1.75E-06		40000		0		0				1

		2		58		410.131		40000		0		0				1

		3		58		810.803		40000		0		0				1

		4		58		1199.11		40000		0		0				1

		5		20		610.498		9121.04		0		0				0

		6		18		668.454		8247.45		0		0				0

		7		17		711.559		7597.72		0		0				0

		8		16		744.495		7101.27		0		0				0

		9		16		770.112		6715.15		0		0				0

		10		15		790.348		6410.12		0		0				0

		11		15		806.558		6165.78		0		0				0

		12		15		-806.558		6165.78		0		0				0

		13		15		-790.348		6410.12		0		0				0

		14		16		-770.112		6715.15		0		0				0

		15		16		-744.495		7101.27		0		0				0

		16		17		-711.559		7597.72		0		0				0

		17		18		-668.454		8247.45		0		0				0

		18		20		-610.498		9121.04		0		0				0

		19		58		-1199.11		40000		0		0				1

		20		58		-810.803		40000		0		0				1

		21		58		-410.131		40000		0		0				1

		22		58		1.75E-06		40000		0		0				1

		23		57		9.48E-05		40000		0		0				1

		24		57		565.69		40000		0		0				1

		25		57		1123.5		40000		0		0				1

		26		19		628.671		8847.11		0		0				0

		27		17		707.569		7657.86		0		0				0

		28		16		763.268		6818.3		0		0				0

		29		14		804.025		6203.97		0		0				0

		30		14		834.669		5742.06		0		0				0

		31		13		858.229		5386.94		0		0				0

		32		13		876.686		5108.73		0		0				0

		33		12		891.384		4887.18		0		0				0

		34		12		-891.384		4887.18		0		0				0

		35		13		-876.686		5108.73		0		0				0

		36		13		-858.229		5386.94		0		0				0

		37		14		-834.669		5742.06		0		0				0

		38		14		-804.025		6203.97		0		0				0

		39		16		-763.268		6818.3		0		0				0

		40		17		-707.569		7657.86		0		0				0

		41		19		-628.671		8847.11		0		0				0

		42		57		-1123.5		40000		0		0				1

		43		57		-565.69		40000		0		0				1

		44		57		9.48E-05		40000		0		0				1

		45		57		-0.000160053		40000		0		0				1

		46		57		677.316		40000		0		0				1

		47		57		1350.9		40000		0		0				1

		48		18		683.222		8024.85		0		0				0

		49		16		760.352		6862.25		0		0				0

		50		14		813.921		6054.8		0		0				0

		51		13		852.663		5470.84		0		0				0

		52		12		881.562		5035.24		0		0				0

		53		12		903.66		4702.15		0		0				0

		54		11		920.911		4442.12		0		0				0

		55		11		934.616		4235.54		0		0				0

		56		11		-934.616		4235.54		0		0				0

		57		11		-920.911		4442.12		0		0				0

		58		12		-903.66		4702.15		0		0				0

		59		12		-881.562		5035.24		0		0				0

		60		13		-852.663		5470.84		0		0				0

		61		14		-813.921		6054.8		0		0				0

		62		16		-760.352		6862.25		0		0				0

		63		18		-683.222		8024.85		0		0				0

		64		57		-1350.9		40000		0		0				1

		65		57		-677.316		40000		0		0				1

		66		57		-0.000160053		40000		0		0				1

		67		57		0.000111802		40000		0		0				1

		68		57		766.884		40000		0		0				1

		69		20		602.791		9237.2		0		0				0

		70		16		720.054		7469.67		0		0				0

		71		14		795.4		6333.97		0		0				0

		72		13		847.138		5554.12		0		0				0

		73		12		884.279		4994.28		0		0				0

		74		11		911.852		4578.67		0		0				0

		75		11		932.872		4261.84		0		0				0

		76		10		949.25		4014.97		0		0				0

		77		10		962.247		3819.05		0		0				0

		78		10		-962.247		3819.05		0		0				0

		79		10		-949.25		4014.97		0		0				0

		80		11		-932.872		4261.84		0		0				0

		81		11		-911.852		4578.67		0		0				0

		82		12		-884.279		4994.28		0		0				0

		83		13		-847.138		5554.12		0		0				0

		84		14		-795.4		6333.97		0		0				0

		85		16		-720.054		7469.67		0		0				0

		86		20		-602.791		9237.2		0		0				0

		87		57		-766.884		40000		0		0				1

		88		57		0.000111802		40000		0		0				1

		89		56		1.85E-05		40000		0		0				1

		90		56		846.908		40000		0		0				1

		91		19		631.409		8805.83		0		0				0

		92		16		747.303		7058.95		0		0				0

		93		14		821.001		5948.08		0		0				0

		94		12		871.196		5191.49		0		0				0

		95		11		907.045		4651.13		0		0				0

		96		11		933.572		4251.28		0		0				0

		97		10		953.757		3947.03		0		0				0

		98		10		969.466		3710.24		0		0				0

		99		9		981.93		3522.37		0		0				0

		100		9		-981.93		3522.37		0		0				0

		101		10		-969.466		3710.24		0		0				0

		102		10		-953.757		3947.03		0		0				0

		103		11		-933.572		4251.28		0		0				0

		104		11		-907.045		4651.13		0		0				0

		105		12		-871.196		5191.49		0		0				0

		106		14		-821.001		5948.08		0		0				0

		107		16		-747.303		7058.95		0		0				0

		108		19		-631.409		8805.83		0		0				0

		109		56		-846.908		40000		0		0				1

		110		56		1.85E-05		40000		0		0				1

		111		56		9.28E-05		40000		0		0				1

		112		56		986.493		40000		0		0				1

		113		17		672.302		8189.45		0		0				0

		114		14		785.843		6478.03		0		0				0

		115		12		856.747		5409.28		0		0				0

		116		11		904.503		4689.44		0		0				0

		117		10		938.383		4178.76		0		0				0

		118		10		963.354		3802.37		0		0				0

		119		9		982.315		3516.57		0		0				0

		120		9		997.057		3294.37		0		0				0

		121		8		1000		3090.16		0		0				0

		122		8		-1000		3090.16		0		0				0

		123		9		-997.057		3294.37		0		0				0

		124		9		-982.315		3516.57		0		0				0

		125		10		-963.354		3802.37		0		0				0

		126		10		-938.383		4178.76		0		0				0

		127		11		-904.503		4689.44		0		0				0

		128		12		-856.747		5409.28		0		0				0

		129		14		-785.843		6478.03		0		0				0

		130		17		-672.302		8189.45		0		0				0

		131		56		-986.493		40000		0		0				1

		132		56		9.28E-05		40000		0		0				1

		133		55		0.00016013		40000		0		0				1

		134		55		1091.98		40000		0		0				1

		135		16		701.046		7756.19		0		0				0

		136		13		812.41		6077.58		0		0				0

		137		12		881.076		5042.56		0		0				0

		138		10		926.991		4350.48		0		0				0

		139		10		959.427		3861.56		0		0				0

		140		9		983.281		3502.01		0		0				0

		141		8		1000		3224.75		0		0				0

		142		8		1000		2969.96		0		0				0

		143		8		1000		2772.79		0		0				0

		144		8		-1000		2772.79		0		0				0

		145		8		-1000		2969.96		0		0				0

		146		8		-1000		3224.75		0		0				0

		147		9		-983.281		3502.01		0		0				0

		148		10		-959.427		3861.56		0		0				0

		149		10		-926.991		4350.48		0		0				0

		150		12		-881.076		5042.56		0		0				0

		151		13		-812.41		6077.58		0		0				0

		152		16		-701.046		7756.19		0		0				0

		153		55		-1091.98		40000		0		0				1

		154		55		0.00016013		40000		0		0				1

		155		55		0.000174139		40000		0		0				1

		156		55		1154.46		40000		0		0				1

		157		16		722.734		7429.29		0		0				0

		158		13		832.16		5779.88		0		0				0

		159		11		899.004		4772.33		0		0				0

		160		10		943.469		4102.11		0		0				0

		161		9		974.794		3629.94		0		0				0

		162		8		997.797		3283.21		0		0				0

		163		8		1000		2973.87		0		0				0

		164		8		1000		2734.83		0		0				0

		165		7		1000		2549.17		0		0				0

		166		7		-1000		2549.17		0		0				0

		167		8		-1000		2734.83		0		0				0

		168		8		-1000		2973.87		0		0				0

		169		8		-997.797		3283.21		0		0				0

		170		9		-974.794		3629.94		0		0				0

		171		10		-943.469		4102.11		0		0				0

		172		11		-899.004		4772.33		0		0				0

		173		13		-832.16		5779.88		0		0				0

		174		16		-722.734		7429.28		0		0				0

		175		55		-1154.46		40000		0		0				1

		176		55		0.000174139		40000		0		0				1

		177		55		-1.75E-06		40000		0		0				1

		178		55		1263.45		40000		0		0				1

		179		15		739.853		7171.23		0		0				0

		180		12		847.574		5547.54		0		0				0

		181		10		912.903		4562.82		0		0				0

		182		9		956.192		3910.32		0		0				0

		183		9		986.626		3451.59		0		0				0

		184		8		1000		3086.67		0		0				0

		185		8		1000		2785.89		0		0				0

		186		7		1000		2558.28		0		0				0

		187		7		1000		2381.78		0		0				0

		188		7		-1000		2381.78		0		0				0

		189		7		-1000		2558.28		0		0				0

		190		8		-1000		2785.89		0		0				0

		191		8		-1000		3086.72		0		0				0

		192		9		-986.626		3451.59		0		0				0

		193		9		-956.192		3910.32		0		0				0

		194		10		-912.903		4562.82		0		0				0

		195		12		-847.574		5547.54		0		0				0

		196		15		-739.854		7171.23		0		0				0

		197		55		-1263.45		40000		0		0				1

		198		55		-1.75E-06		40000		0		0				1

		199		54		0.000144891		40000		0		0				1

		200		54		1348.36		40000		0		0				1

		201		14		753.805		6960.94		0		0				0

		202		12		860.022		5359.91		0		0				0

		203		10		924.07		4394.51		0		0				0

		204		9		966.381		3756.74		0		0				0

		205		8		996.087		3308.99		0		0				0

		206		8		1000		2927.26		0		0				0

		207		7		1000		2638.6		0		0				0

		208		7		1000		2420.43		0		0				0

		209		7		1000		2250.47		0		0				0

		210		7		-1000		2250.47		0		0				0

		211		7		-1000		2420.43		0		0				0

		212		7		-1000		2638.6		0		0				0

		213		8		-1000		2927.26		0		0				0

		214		8		-996.087		3308.99		0		0				0

		215		9		-966.381		3756.74		0		0				0

		216		10		-924.07		4394.51		0		0				0

		217		12		-860.022		5359.91		0		0				0

		218		14		-753.805		6960.94		0		0				0

		219		54		-1348.36		40000		0		0				1

		220		54		0.000144891		40000		0		0				1

																				%%

																48				21.8

																172				78.2





90 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		56		0.00011052		40000		0		0				1

		2		56		409.258		40000		0		0				1

		3		56		808.209		40000		0		0				1

		4		56		1188.35		40000		0		0				1

		5		19		593.683		9374.48		0		0				0

		6		18		651.66		8500.58		0		0				0

		7		17		694.984		7847.55		0		0				0

		8		16		728.186		7347.1		0		0				0

		9		15		754.067		6956.99		0		0				0

		10		15		774.546		6648.31		0		0				0

		11		14		790.97		6400.74		0		0				0

		12		14		-790.97		6400.74		0		0				0

		13		15		-774.546		6648.31		0		0				0

		14		15		-754.067		6956.99		0		0				0

		15		16		-728.186		7347.1		0		0				0

		16		17		-694.984		7847.55		0		0				0

		17		18		-651.66		8500.58		0		0				0

		18		19		-593.683		9374.48		0		0				0

		19		56		-1188.35		40000		0		0				1

		20		56		-808.209		40000		0		0				1

		21		56		-409.258		40000		0		0				1

		22		56		0.00011052		40000		0		0				1

		23		56		-0.000359625		40000		0		0				1

		24		56		564.933		40000		0		0				1

		25		56		1117.74		40000		0		0				1

		26		19		612.462		9091.42		0		0				0

		27		17		691.641		7897.96		0		0				0

		28		15		747.83		7051		0		0				0

		29		14		789.097		6428.97		0		0				0

		30		13		820.205		5960.09		0		0				0

		31		13		844.163		5598.96		0		0				0

		32		12		862.956		5315.68		0		0				0

		33		12		877.934		5089.92		0		0				0

		34		12		-877.934		5089.92		0		0				0

		35		12		-862.956		5315.68		0		0				0

		36		13		-844.163		5598.96		0		0				0

		37		13		-820.205		5960.09		0		0				0

		38		14		-789.097		6428.97		0		0				0

		39		15		-747.83		7051		0		0				0

		40		17		-691.641		7897.95		0		0				0

		41		19		-612.462		9091.42		0		0				0

		42		56		-1117.74		40000		0		0				1

		43		56		-564.933		40000		0		0				1

		44		56		-0.000359625		40000		0		0				1

		45		56		3.00E-06		40000		0		0				1

		46		56		678.081		40000		0		0				1

		47		56		1340.06		40000		0		0				1

		48		17		667.685		8259.04		0		0				0

		49		15		745.413		7087.44		0		0				0

		50		14		799.661		6269.74		0		0				0

		51		13		839.022		5676.44		0		0				0

		52		12		868.448		5232.91		0		0				0

		53		11		890.982		4893.25		0		0				0

		54		11		908.588		4627.86		0		0				0

		55		11		922.585		4416.89		0		0				0

		56		11		-922.585		4416.89		0		0				0

		57		11		-908.588		4627.86		0		0				0

		58		11		-890.982		4893.25		0		0				0

		59		12		-868.448		5232.91		0		0				0

		60		13		-839.022		5676.44		0		0				0

		61		14		-799.661		6269.74		0		0				0

		62		15		-745.413		7087.44		0		0				0

		63		17		-667.685		8259.03		0		0				0

		64		56		-1340.06		40000		0		0				1

		65		56		-678.081		40000		0		0				1

		66		56		3.00E-06		40000		0		0				1

		67		55		5.57E-05		40000		0		0				1

		68		55		769.546		40000		0		0				1

		69		56		1438.19		40000		0		0				1

		70		16		705.202		7693.54		0		0				0

		71		14		781.328		6546.09		0		0				0

		72		13		833.839		5754.57		0		0				0

		73		12		871.647		5184.69		0		0				0

		74		11		899.767		4760.82		0		0				0

		75		11		921.23		4437.31		0		0				0

		76		10		937.965		4185.06		0		0				0

		77		10		951.251		3984.8		0		0				0

		78		10		-951.251		3984.8		0		0				0

		79		10		-937.965		4185.06		0		0				0

		80		11		-921.23		4437.31		0		0				0

		81		11		-899.767		4760.82		0		0				0

		82		12		-871.647		5184.69		0		0				0

		83		13		-833.839		5754.57		0		0				0

		84		14		-781.328		6546.09		0		0				0

		85		16		-705.202		7693.54		0		0				0

		86		56		-1438.19		40000		0		0				1

		87		55		-769.546		40000		0		0				1

		88		55		5.57E-05		40000		0		0				1

		89		55		-8.34E-05		40000		0		0				1

		90		55		848.893		40000		0		0				1

		91		18		616.562		9029.63		0		0				0

		92		15		733.086		7273.24		0		0				0

		93		13		807.677		6148.92		0		0				0

		94		12		858.695		5379.91		0		0				0

		95		11		895.229		4829.23		0		0				0

		96		10		922.308		4421.06		0		0				0

		97		10		942.932		4110.2		0		0				0

		98		10		958.992		3868.12		0		0				0

		99		9		971.737		3676.02		0		0				0

		100		9		-971.737		3676.02		0		0				0

		101		10		-958.992		3868.12		0		0				0

		102		10		-942.932		4110.2		0		0				0

		103		10		-922.308		4421.06		0		0				0

		104		11		-895.229		4829.23		0		0				0

		105		12		-858.695		5379.91		0		0				0

		106		13		-807.677		6148.92		0		0				0

		107		15		-733.086		7273.24		0		0				0

		108		18		-616.562		9029.63		0		0				0

		109		55		-848.893		40000		0		0				1

		110		55		-8.34E-05		40000		0		0				1

		111		55		4.05E-05		40000		0		0				1

		112		55		985.123		40000		0		0				1

		113		17		658.3		8400.51		0		0				0

		114		14		772.731		6675.67		0		0				0

		115		12		844.644		5591.7		0		0				0

		116		11		893.262		4858.88		0		0				0

		117		10		927.829		4337.84		0		0				0

		118		9		953.339		3953.33		0		0				0

		119		9		972.72		3661.19		0		0				0

		120		9		987.793		3434		0		0				0

		121		8		999.751		3253.75		0		0				0

		122		8		-999.751		3253.75		0		0				0

		123		9		-987.793		3434		0		0				0

		124		9		-972.719		3661.2		0		0				0

		125		9		-953.339		3953.33		0		0				0

		126		10		-927.829		4337.84		0		0				0

		127		11		-893.262		4858.88		0		0				0

		128		12		-844.645		5591.7		0		0				0

		129		14		-772.731		6675.67		0		0				0

		130		17		-658.3		8400.51		0		0				0

		131		55		-985.123		40000		0		0				1

		132		55		4.05E-05		40000		0		0				1

		133		54		-4.99E-05		40000		0		0				1

		134		54		1094.13		40000		0		0				1

		135		16		687.793		7955.95		0		0				0

		136		13		800.213		6261.42		0		0				0

		137		11		869.937		5210.46		0		0				0

		138		10		916.713		4505.4		0		0				0

		139		9		949.82		4006.38		0		0				0

		140		9		974.19		3639.04		0		0				0

		141		8		992.683		3360.3		0		0				0

		142		8		1000		3120.86		0		0				0

		143		8		1000		2916.06		0		0				0

		144		8		-1000		2916.06		0		0				0

		145		8		-1000		3120.91		0		0				0

		146		8		-992.683		3360.3		0		0				0

		147		9		-974.19		3639.04		0		0				0

		148		9		-949.82		4006.38		0		0				0

		149		10		-916.713		4505.4		0		0				0

		150		11		-869.937		5210.46		0		0				0

		151		13		-800.213		6261.42		0		0				0

		152		16		-687.793		7955.94		0		0				0

		153		54		-1094.13		40000		0		0				1

		154		54		-4.99E-05		40000		0		0				1

		155		54		0.0001646		40000		0		0				1

		156		54		1180.93		40000		0		0				1

		157		15		710.156		7618.87		0		0				0

		158		12		820.735		5952.09		0		0				0

		159		11		888.649		4928.41		0		0				0

		160		10		933.961		4245.42		0		0				0

		161		9		965.932		3763.52		0		0				0

		162		8		989.43		3409.32		0		0				0

		163		8		1000		3117.6		0		0				0

		164		7		1000		2869.62		0		0				0

		165		7		1000		2677.25		0		0				0

		166		7		-1000		2677.25		0		0				0

		167		7		-1000		2869.62		0		0				0

		168		8		-1000		3117.6		0		0				0

		169		8		-989.43		3409.32		0		0				0

		170		9		-965.932		3763.52		0		0				0

		171		10		-933.961		4245.42		0		0				0

		172		11		-888.65		4928.41		0		0				0

		173		12		-820.735		5952.09		0		0				0

		174		15		-710.156		7618.87		0		0				0

		175		54		-1180.93		40000		0		0				1

		176		54		0.0001646		40000		0		0				1

		177		54		-0.000110553		40000		0		0				1

		178		54		1248.19		40000		0		0				1

		179		15		727.881		7351.69		0		0				0

		180		12		836.81		5709.79		0		0				0

		181		10		903.206		4708.99		0		0				0

		182		9		947.32		4044.05		0		0				0

		183		8		978.377		3575.92		0		0				0

		184		8		1000		3228.36		0		0				0

		185		7		1000		2916.43		0		0				0

		186		7		1000		2680.82		0		0				0

		187		7		1000		2498.24		0		0				0

		188		7		-1000		2498.24		0		0				0

		189		7		-1000		2680.82		0		0				0

		190		7		-1000		2916.43		0		0				0

		191		8		-1000		3228.36		0		0				0

		192		8		-978.377		3575.92		0		0				0

		193		9		-947.32		4044.05		0		0				0

		194		10		-903.206		4708.99		0		0				0

		195		12		-836.81		5709.79		0		0				0

		196		15		-727.881		7351.69		0		0				0

		197		54		-1248.19		40000		0		0				1

		198		54		-0.000110553		40000		0		0				1

		199		53		-0.000119295		40000		0		0				1

		200		53		1343.01		40000		0		0				1

		201		14		742.377		7133.2		0		0				0

		202		11		849.832		5513.5		0		0				0

		203		10		914.933		4532.22		0		0				0

		204		9		958.047		3882.36		0		0				0

		205		8		988.352		3425.58		0		0				0

		206		8		1000		3057.45		0		0				0

		207		7		1000		2758.72		0		0				0

		208		7		1000		2533.15		0		0				0

		209		7		1000		2357.72		0		0				0

		210		7		-1000		2357.72		0		0				0

		211		7		-1000		2533.15		0		0				0

		212		7		-1000		2758.72		0		0				0

		213		8		-1000		3057.45		0		0				0

		214		8		-988.352		3425.58		0		0				0

		215		9		-958.047		3882.36		0		0				0

		216		10		-914.933		4532.22		0		0				0

		217		11		-849.832		5513.5		0		0				0

		218		14		-742.377		7133.2		0		0				0

		219		53		-1343.01		40000		0		0				1

		220		53		-0.000119295		40000		0		0				1

																				%%

																50				22.7

																170				77.3





100 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		55		0.000165832		40000		0		0				1

		2		55		405.532		40000		0		0				1

		3		55		799.9		40000		0		0				1

		4		55		1173.61		40000		0		0				1

		5		55		1520.69		40000		0		0				1

		6		18		636.502		8729.07		0		0				0

		7		16		679.959		8074.03		0		0				0

		8		16		713.36		7570.57		0		0				0

		9		15		739.45		7177.31		0		0				0

		10		14		760.126		6865.66		0		0				0

		11		14		776.727		6615.43		0		0				0

		12		14		-776.727		6615.43		0		0				0

		13		14		-760.126		6865.66		0		0				0

		14		15		-739.45		7177.31		0		0				0

		15		16		-713.36		7570.57		0		0				0

		16		16		-679.959		8074.03		0		0				0

		17		18		-636.502		8729.07		0		0				0

		18		55		-1520.69		40000		0		0				1

		19		55		-1173.61		40000		0		0				1

		20		55		-799.9		40000		0		0				1

		21		55		-405.532		40000		0		0				1

		22		55		0.000165832		40000		0		0				1

		23		55		-0.000268051		40000		0		0				1

		24		55		560.726		40000		0		0				1

		25		55		1106.4		40000		0		0				1

		26		18		597.913		9310.72		0		0				0

		27		16		677.186		8115.84		0		0				0

		28		15		733.758		7263.11		0		0				0

		29		14		775.446		6634.74		0		0				0

		30		13		806.945		6159.95		0		0				0

		31		13		831.245		5793.67		0		0				0

		32		12		850.329		5506.02		0		0				0

		33		12		865.55		5276.58		0		0				0

		34		12		-865.55		5276.58		0		0				0

		35		12		-850.329		5506.02		0		0				0

		36		13		-831.245		5793.67		0		0				0

		37		13		-806.945		6159.95		0		0				0

		38		14		-775.446		6634.74		0		0				0

		39		15		-733.758		7263.11		0		0				0

		40		16		-677.186		8115.84		0		0				0

		41		18		-597.913		9310.72		0		0				0

		42		55		-1106.4		40000		0		0				1

		43		55		-560.726		40000		0		0				1

		44		55		-0.000268051		40000		0		0				1

		45		55		-7.76E-05		40000		0		0				1

		46		55		674.583		40000		0		0				1

		47		55		1331.81		40000		0		0				1

		48		17		653.596		8471.41		0		0				0

		49		15		731.776		7292.98		0		0				0

		50		13		786.592		6466.74		0		0				0

		51		13		826.485		5865.43		0		0				0

		52		12		856.368		5414.99		0		0				0

		53		11		879.284		5069.58		0		0				0

		54		11		897.205		4799.44		0		0				0

		55		10		911.46		4584.57		0		0				0

		56		10		-911.46		4584.57		0		0				0

		57		11		-897.205		4799.44		0		0				0

		58		11		-879.284		5069.58		0		0				0

		59		12		-856.368		5414.99		0		0				0

		60		13		-826.485		5865.43		0		0				0

		61		13		-786.592		6466.74		0		0				0

		62		15		-731.776		7292.98		0		0				0

		63		17		-653.596		8471.41		0		0				0

		64		55		-1331.81		40000		0		0				1

		65		55		-674.583		40000		0		0				1

		66		55		-7.76E-05		40000		0		0				1

		67		54		-0.000119339		40000		0		0				1

		68		54		767.091		40000		0		0				1

		69		54		1514.71		40000		0		0				1

		70		16		691.667		7897.56		0		0				0

		71		14		768.429		6740.51		0		0				0

		72		13		821.604		5939		0		0				0

		73		12		859.993		5360.34		0		0				0

		74		11		888.597		4929.19		0		0				0

		75		10		910.454		4599.74		0		0				0

		76		10		927.507		4342.69		0		0				0

		77		10		941.053		4138.52		0		0				0

		78		10		-941.053		4138.52		0		0				0

		79		10		-927.507		4342.69		0		0				0

		80		10		-910.454		4599.74		0		0				0

		81		11		-888.597		4929.19		0		0				0

		82		12		-859.994		5360.34		0		0				0

		83		13		-821.604		5939		0		0				0

		84		14		-768.429		6740.51		0		0				0

		85		16		-691.667		7897.56		0		0				0

		86		54		-1514.71		40000		0		0				1

		87		54		-767.092		40000		0		0				1

		88		54		-0.000119339		40000		0		0				1

		89		54		4.17E-05		40000		0		0				1

		90		54		846.034		40000		0		0				1

		91		18		603.217		9230.78		0		0				0

		92		15		720.082		7469.25		0		0				0

		93		13		795.424		6333.6		0		0				0

		94		12		847.159		5553.8		0		0				0

		95		11		884.298		4994		0		0				0

		96		10		911.869		4578.42		0		0				0

		97		10		932.887		4261.61		0		0				0

		98		9		949.264		4014.76		0		0				0

		99		9		962.261		3818.84		0		0				0

		100		9		-962.261		3818.84		0		0				0

		101		9		-949.264		4014.76		0		0				0

		102		10		-932.887		4261.61		0		0				0

		103		10		-911.869		4578.42		0		0				0

		104		11		-884.298		4994		0		0				0

		105		12		-847.159		5553.8		0		0				0

		106		13		-795.424		6333.6		0		0				0

		107		15		-720.082		7469.25		0		0				0

		108		18		-603.217		9230.78		0		0				0

		109		54		-846.034		40000		0		0				1

		110		54		4.17E-05		40000		0		0				1

		111		54		0.00012026		40000		0		0				1

		112		54		978.066		40000		0		0				1

		113		17		645.529		8593		0		0				0

		114		14		760.671		6857.45		0		0				0

		115		12		833.459		5760.3		0		0				0

		116		11		882.839		5015.99		0		0				0

		117		10		918.021		4485.68		0		0				0

		118		9		944.018		4093.82		0		0				0

		119		9		963.781		3795.94		0		0				0

		120		9		979.157		3564.18		0		0				0

		121		8		991.355		3380.3		0		0				0

		122		8		-991.355		3380.3		0		0				0

		123		9		-979.157		3564.18		0		0				0

		124		9		-963.781		3795.94		0		0				0

		125		9		-944.018		4093.82		0		0				0

		126		10		-918.022		4485.68		0		0				0

		127		11		-882.839		5015.99		0		0				0

		128		12		-833.459		5760.3		0		0				0

		129		14		-760.671		6857.45		0		0				0

		130		17		-645.529		8593		0		0				0

		131		54		-978.066		40000		0		0				1

		132		54		0.00012026		40000		0		0				1

		133		53		-0.000201025		40000		0		0				1

		134		53		1086.97		40000		0		0				1

		135		16		675.643		8139.09		0		0				0

		136		13		788.947		6431.24		0		0				0

		137		11		859.6		5366.27		0		0				0

		138		10		907.148		4649.58		0		0				0

		139		9		940.861		4141.41		0		0				0

		140		9		965.702		3766.98		0		0				0

		141		8		984.561		3482.71		0		0				0

		142		8		999.222		3261.73		0		0				0

		143		8		1000		3052.2		0		0				0

		144		8		-1000		3052.2		0		0				0

		145		8		-999.222		3261.73		0		0				0

		146		8		-984.561		3482.71		0		0				0

		147		9		-965.702		3766.98		0		0				0

		148		9		-940.861		4141.41		0		0				0

		149		10		-907.148		4649.58		0		0				0

		150		11		-859.6		5366.27		0		0				0

		151		13		-788.947		6431.24		0		0				0

		152		16		-675.643		8139.09		0		0				0

		153		53		-1086.97		40000		0		0				1

		154		53		-0.000201025		40000		0		0				1

		155		53		0.000138059		40000		0		0				1

		156		53		1179.81		40000		0		0				1

		157		15		698.577		7793.4		0		0				0

		158		12		810.143		6111.75		0		0				0

		159		11		879.01		5073.71		0		0				0

		160		10		925.085		4379.2		0		0				0

		161		9		957.647		3888.39		0		0				0

		162		8		981.598		3527.37		0		0				0

		163		8		999.762		3253.58		0		0				0

		164		7		1000		2997.82		0		0				0

		165		7		1000		2799.07		0		0				0

		166		7		-1000		2799.07		0		0				0

		167		7		-1000		2997.82		0		0				0

		168		8		-999.762		3253.58		0		0				0

		169		8		-981.598		3527.37		0		0				0

		170		9		-957.647		3888.39		0		0				0

		171		10		-925.085		4379.2		0		0				0

		172		11		-879.01		5073.71		0		0				0

		173		12		-810.143		6111.75		0		0				0

		174		15		-698.577		7793.4		0		0				0

		175		53		-1179.81		40000		0		0				1

		176		53		0.000138059		40000		0		0				1

		177		53		0.00011369		40000		0		0				1

		178		53		1261.35		40000		0		0				1

		179		14		716.821		7518.4		0		0				0

		180		12		826.8		5860.68		0		0				0

		181		10		894.152		4845.46		0		0				0

		182		9		939.017		4169.2		0		0				0

		183		8		970.647		3692.45		0		0				0

		184		8		993.881		3342.24		0		0				0

		185		7		1000		3040.78		0		0				0

		186		7		1000		2797.49		0		0				0

		187		7		1000		2609.07		0		0				0

		188		7		-1000		2609.07		0		0				0

		189		7		-1000		2797.49		0		0				0

		190		7		-1000		3040.78		0		0				0

		191		8		-993.881		3342.24		0		0				0

		192		8		-970.647		3692.45		0		0				0

		193		9		-939.017		4169.2		0		0				0

		194		10		-894.152		4845.46		0		0				0

		195		12		-826.8		5860.68		0		0				0

		196		14		-716.821		7518.4		0		0				0

		197		53		-1261.35		40000		0		0				1

		198		53		0.00011369		40000		0		0				1

		199		52		-0.000134187		40000		0		0				1

		200		52		1332.68		40000		0		0				1

		201		14		731.787		7292.82		0		0				0

		202		11		840.33		5656.73		0		0				0

		203		10		906.382		4661.11		0		0				0

		204		9		950.23		4000.2		0		0				0

		205		8		981.086		3535.09		0		0				0

		206		8		1000		3181.61		0		0				0

		207		7		1000		2873.23		0		0				0

		208		7		1000		2640.57		0		0				0

		209		6		1000		2459.79		0		0				0

		210		6		-1000		2459.79		0		0				0

		211		7		-1000		2640.57		0		0				0

		212		7		-1000		2873.23		0		0				0

		213		8		-1000		3181.61		0		0				0

		214		8		-981.086		3535.09		0		0				0

		215		9		-950.23		4000.2		0		0				0

		216		10		-906.382		4661.11		0		0				0

		217		11		-840.33		5656.73		0		0				0

		218		14		-731.787		7292.82		0		0				0

		219		52		-1332.68		40000		0		0				1

		220		52		-0.000134187		40000		0		0				1

																				%%

																52				23.6

																168				76.4





110 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		54		0.000193106		40000		0		0				1

		2		54		400.688		40000		0		0				1

		3		54		790.65		40000		0		0				1

		4		54		1161.27		40000		0		0				1

		5		54		1502.8		40000		0		0				1

		6		17		622.719		8936.81		0		0				0

		7		16		666.208		8281.31		0		0				0

		8		15		699.752		7775.68		0		0				0

		9		15		726.006		7379.96		0		0				0

		10		14		746.841		7065.91		0		0				0

		11		14		763.588		6813.48		0		0				0

		12		14		-763.588		6813.48		0		0				0

		13		14		-746.841		7065.91		0		0				0

		14		15		-726.006		7379.95		0		0				0

		15		15		-699.752		7775.68		0		0				0

		16		16		-666.208		8281.31		0		0				0

		17		17		-622.72		8936.81		0		0				0

		18		54		-1502.8		40000		0		0				1

		19		54		-1161.27		40000		0		0				1

		20		54		-790.65		40000		0		0				1

		21		54		-400.688		40000		0		0				1

		22		54		0.000193106		40000		0		0				1

		23		54		1.88E-05		40000		0		0				1

		24		54		553.999		40000		0		0				1

		25		54		1090.96		40000		0		0				1

		26		54		1573.84		40000		0		0				1

		27		16		663.958		8315.23		0		0				0

		28		15		720.822		7458.1		0		0				0

		29		14		762.859		6824.46		0		0				0

		30		13		794.692		6344.65		0		0				0

		31		12		819.287		5973.92		0		0				0

		32		12		838.623		5682.46		0		0				0

		33		12		854.058		5449.81		0		0				0

		34		12		-854.058		5449.8		0		0				0

		35		12		-838.623		5682.46		0		0				0

		36		12		-819.287		5973.92		0		0				0

		37		13		-794.692		6344.65		0		0				0

		38		14		-762.859		6824.46		0		0				0

		39		15		-720.822		7458.1		0		0				0

		40		16		-663.956		8315.24		0		0				0

		41		54		-1573.84		40000		0		0				1

		42		54		-1090.96		40000		0		0				1

		43		54		-553.999		40000		0		0				1

		44		54		1.88E-05		40000		0		0				1

		45		54		9.67E-05		40000		0		0				1

		46		54		666.487		40000		0		0				1

		47		54		1310.9		40000		0		0				1

		48		16		640.716		8665.55		0		0				0

		49		15		719.233		7482.05		0		0				0

		50		13		774.524		6648.64		0		0				0

		51		12		814.876		6040.41		0		0				0

		52		12		845.16		5583.93		0		0				0

		53		11		868.414		5233.42		0		0				0

		54		11		886.616		4959.06		0		0				0

		55		10		901.102		4740.71		0		0				0

		56		10		-901.102		4740.71		0		0				0

		57		11		-886.616		4959.06		0		0				0

		58		11		-868.414		5233.42		0		0				0

		59		12		-845.16		5583.93		0		0				0

		60		12		-814.876		6040.41		0		0				0

		61		13		-774.524		6648.64		0		0				0

		62		15		-719.233		7482.05		0		0				0

		63		16		-640.716		8665.55		0		0				0

		64		54		-1310.9		40000		0		0				1

		65		54		-666.488		40000		0		0				1

		66		54		9.67E-05		40000		0		0				1

		67		53		0.000355124		40000		0		0				1

		68		54		758.355		40000		0		0				1

		69		54		1497.76		40000		0		0				1

		70		15		679.238		8084.9		0		0				0

		71		14		756.523		6919.98		0		0				0

		72		12		810.269		6109.85		0		0				0

		73		11		849.17		5523.48		0		0				0

		74		11		878.204		5085.85		0		0				0

		75		10		900.414		4751.08		0		0				0

		76		10		917.755		4489.69		0		0				0

		77		10		931.534		4282		0		0				0

		78		10		-931.534		4282		0		0				0

		79		10		-917.755		4489.69		0		0				0

		80		10		-900.414		4751.08		0		0				0

		81		11		-878.204		5085.85		0		0				0

		82		11		-849.17		5523.48		0		0				0

		83		12		-810.269		6109.85		0		0				0

		84		14		-756.523		6919.98		0		0				0

		85		15		-679.238		8084.9		0		0				0

		86		54		-1497.76		40000		0		0				1

		87		54		-758.355		40000		0		0				1

		88		53		0.000355124		40000		0		0				1

		89		53		-7.84E-05		40000		0		0				1

		90		53		837.412		40000		0		0				1

		91		53		1622.1		40000		0		0				1

		92		15		708.105		7649.79		0		0				0

		93		13		784.082		6504.57		0		0				0

		94		12		836.444		5715.31		0		0				0

		95		11		874.121		5147.39		0		0				0

		96		10		902.134		4725.16		0		0				0

		97		10		923.508		4402.97		0		0				0

		98		9		940.172		4151.8		0		0				0

		99		9		953.4		3952.42		0		0				0

		100		9		-953.4		3952.42		0		0				0

		101		9		-940.172		4151.8		0		0				0

		102		10		-923.508		4402.97		0		0				0

		103		10		-902.134		4725.16		0		0				0

		104		11		-874.121		5147.39		0		0				0

		105		12		-836.444		5715.31		0		0				0

		106		13		-784.082		6504.57		0		0				0

		107		15		-708.105		7649.79		0		0				0

		108		17		-591.272		9410.82		0		0				0

		109		53		-837.413		40000		0		0				1

		110		53		-7.84E-05		40000		0		0				1

		111		53		7.24E-05		40000		0		0				1

		112		53		971.856		40000		0		0				1

		113		16		633.819		8769.51		0		0				0

		114		13		749.507		7025.72		0		0				0

		115		12		823.058		5917.08		0		0				0

		116		11		873.118		5162.51		0		0				0

		117		10		908.857		4623.82		0		0				0

		118		9		935.296		4225.29		0		0				0

		119		9		955.409		3922.13		0		0				0

		120		8		971.062		3686.18		0		0				0

		121		8		983.481		3498.99		0		0				0

		122		8		-983.481		3498.99		0		0				0

		123		8		-971.062		3686.18		0		0				0

		124		9		-955.409		3922.13		0		0				0

		125		9		-935.296		4225.29		0		0				0

		126		10		-908.857		4623.82		0		0				0

		127		11		-873.118		5162.51		0		0				0

		128		12		-823.058		5917.08		0		0				0

		129		13		-749.507		7025.72		0		0				0

		130		16		-633.819		8769.51		0		0				0

		131		53		-971.856		40000		0		0				1

		132		53		7.24E-05		40000		0		0				1

		133		53		6.86E-06		40000		0		0				1

		134		53		1083.15		40000		0		0				1

		135		15		664.435		8308.04		0		0				0

		136		13		778.478		6589.04		0		0				0

		137		11		849.955		5511.65		0		0				0

		138		10		898.198		4784.48		0		0				0

		139		9		932.465		4267.97		0		0				0

		140		9		957.738		3887.02		0		0				0

		141		8		976.934		3597.67		0		0				0

		142		8		991.859		3372.71		0		0				0

		143		8		1000		3182.1		0		0				0

		144		8		-1000		3182.1		0		0				0

		145		8		-991.859		3372.71		0		0				0

		146		8		-976.934		3597.67		0		0				0

		147		9		-957.738		3887.02		0		0				0

		148		9		-932.465		4267.97		0		0				0

		149		10		-898.198		4784.48		0		0				0

		150		11		-849.955		5511.65		0		0				0

		151		13		-778.478		6589.04		0		0				0

		152		15		-664.435		8308.03		0		0				0

		153		53		-1083.15		40000		0		0				1

		154		53		6.86E-06		40000		0		0				1

		155		52		-8.51E-05		40000		0		0				1

		156		52		1170.77		40000		0		0				1

		157		15		687.855		7955.01		0		0				0

		158		12		800.269		6260.57		0		0				0

		159		10		869.988		5209.69		0		0				0

		160		9		916.76		4504.69		0		0				0

		161		9		949.863		4005.72		0		0				0

		162		8		974.232		3638.4		0		0				0

		163		8		992.719		3359.74		0		0				0

		164		7		1000		3120.24		0		0				0

		165		7		1000		2915.36		0		0				0

		166		7		-1000		2915.36		0		0				0

		167		7		-1000		3120.24		0		0				0

		168		8		-992.719		3359.74		0		0				0

		169		8		-974.232		3638.4		0		0				0

		170		9		-949.863		4005.72		0		0				0

		171		9		-916.76		4504.69		0		0				0

		172		10		-869.988		5209.69		0		0				0

		173		12		-800.269		6260.57		0		0				0

		174		15		-687.855		7955.01		0		0				0

		175		52		-1170.77		40000		0		0				1

		176		52		-8.51E-05		40000		0		0				1

		177		52		0.00013496		40000		0		0				1

		178		52		1247.77		40000		0		0				1

		179		14		706.548		7673.26		0		0				0

		180		11		817.443		6001.72		0		0				0

		181		10		885.658		4973.49		0		0				0

		182		9		931.21		4286.89		0		0				0

		183		8		963.368		3802.17		0		0				0

		184		8		987.005		3445.88		0		0				0

		185		7		1000		3159.64		0		0				0

		186		7		1000		2908.97		0		0				0

		187		7		1000		2714.95		0		0				0

		188		7		-1000		2714.95		0		0				0

		189		7		-1000		2909.06		0		0				0

		190		7		-1000		3159.64		0		0				0

		191		8		-987.005		3445.88		0		0				0

		192		8		-963.368		3802.17		0		0				0

		193		9		-931.21		4286.89		0		0				0

		194		10		-885.658		4973.49		0		0				0

		195		11		-817.443		6001.72		0		0				0

		196		14		-706.548		7673.26		0		0				0

		197		52		-1247.77		40000		0		0				1

		198		52		0.00013496		40000		0		0				1

		199		51		7.24E-06		40000		0		0				1

		200		52		1321.56		40000		0		0				1

		201		14		721.923		7441.5		0		0				0

		202		11		831.426		5790.94		0		0				0

		203		10		898.342		4782.3		0		0				0

		204		9		942.864		4111.22		0		0				0

		205		8		974.231		3638.42		0		0				0

		206		7		997.268		3291.18		0		0				0

		207		7		1000		2982.75		0		0				0

		208		7		1000		2743.29		0		0				0

		209		6		1000		2557.33		0		0				0

		210		6		-1000		2557.33		0		0				0

		211		7		-1000		2743.29		0		0				0

		212		7		-1000		2982.75		0		0				0

		213		7		-997.268		3291.18		0		0				0

		214		8		-974.231		3638.42		0		0				0

		215		9		-942.864		4111.22		0		0				0

		216		10		-898.342		4782.3		0		0				0

		217		11		-831.427		5790.94		0		0				0

		218		14		-721.923		7441.5		0		0				0

		219		52		-1321.56		40000		0		0				1

		220		51		7.24E-06		40000		0		0				1

																				%%

																55				25.0

																165				75.0





120 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		53		-0.000113351		40000		0		0				1

		2		53		395.137		40000		0		0				1

		3		53		779.713		40000		0		0				1

		4		53		1143.83		40000		0		0				1

		5		53		1468.09		40000		0		0				1

		6		17		610.172		9125.94		0		0				0

		7		16		653.608		8471.23		0		0				0

		8		15		687.256		7964.04		0		0				0

		9		14		713.639		7566.36		0		0				0

		10		14		734.605		7250.34		0		0				0

		11		14		751.474		6996.08		0		0				0

		12		14		-751.474		6996.08		0		0				0

		13		14		-734.605		7250.34		0		0				0

		14		14		-713.639		7566.36		0		0				0

		15		15		-687.256		7964.05		0		0				0

		16		16		-653.608		8471.23		0		0				0

		17		17		-610.172		9125.94		0		0				0

		18		53		-1468.09		40000		0		0				1

		19		53		-1143.83		40000		0		0				1

		20		53		-779.713		40000		0		0				1

		21		53		-395.138		40000		0		0				1

		22		53		-0.000113351		40000		0		0				1

		23		53		-0.000140625		40000		0		0				1

		24		53		548.39		40000		0		0				1

		25		53		1088.47		40000		0		0				1

		26		53		1666.93		40000		0		0				1

		27		16		651.797		8498.52		0		0				0

		28		14		708.885		7638.03		0		0				0

		29		13		751.213		7000		0		0				0

		30		13		783.332		6515.87		0		0				0

		31		12		808.185		6141.26		0		0				0

		32		12		827.744		5846.45		0		0				0

		33		11		843.368		5610.94		0		0				0

		34		11		-843.368		5610.94		0		0				0

		35		12		-827.744		5846.45		0		0				0

		36		12		-808.185		6141.26		0		0				0

		37		13		-783.332		6515.87		0		0				0

		38		13		-751.213		7000		0		0				0

		39		14		-708.885		7638.03		0		0				0

		40		16		-651.797		8498.52		0		0				0

		41		53		-1666.93		40000		0		0				1

		42		53		-1088.47		40000		0		0				1

		43		53		-548.39		40000		0		0				1

		44		53		-0.000140625		40000		0		0				1

		45		53		-5.87E-05		40000		0		0				1

		46		53		660.766		40000		0		0				1

		47		53		1318.17		40000		0		0				1

		48		16		628.921		8843.34		0		0				0

		49		14		707.634		7656.88		0		0				0

		50		13		763.326		6817.43		0		0				0

		51		12		804.077		6203.18		0		0				0

		52		11		834.716		5741.35		0		0				0

		53		11		858.272		5386.29		0		0				0

		54		10		876.725		5108.14		0		0				0

		55		10		891.42		4886.65		0		0				0

		56		10		-891.42		4886.65		0		0				0

		57		10		-876.725		5108.14		0		0				0

		58		11		-858.272		5386.29		0		0				0

		59		11		-834.716		5741.35		0		0				0

		60		12		-804.077		6203.18		0		0				0

		61		13		-763.326		6817.43		0		0				0

		62		14		-707.634		7656.88		0		0				0

		63		16		-628.921		8843.34		0		0				0

		64		53		-1318.17		40000		0		0				1

		65		53		-660.766		40000		0		0				1

		66		53		-5.87E-05		40000		0		0				1

		67		53		0.00018416		40000		0		0				1

		68		53		750.842		40000		0		0				1

		69		53		1468.97		40000		0		0				1

		70		15		667.759		8257.93		0		0				0

		71		13		745.467		7086.61		0		0				0

		72		12		799.71		6269		0		0				0

		73		11		839.066		5675.79		0		0				0

		74		11		868.485		5232.35		0		0				0

		75		10		891.014		4892.76		0		0				0

		76		10		908.616		4627.44		0		0				0

		77		9		922.609		4416.53		0		0				0

		78		9		-922.609		4416.53		0		0				0

		79		10		-908.616		4627.44		0		0				0

		80		10		-891.014		4892.76		0		0				0

		81		11		-868.485		5232.34		0		0				0

		82		11		-839.066		5675.79		0		0				0

		83		12		-799.71		6269		0		0				0

		84		13		-745.467		7086.61		0		0				0

		85		15		-667.759		8257.93		0		0				0

		86		53		-1468.97		40000		0		0				1

		87		53		-750.842		40000		0		0				1

		88		53		0.00018416		40000		0		0				1

		89		52		0.000212899		40000		0		0				1

		90		52		826.701		40000		0		0				1

		91		53		1577.69		40000		0		0				1

		92		14		697.001		7817.16		0		0				0

		93		13		773.517		6663.81		0		0				0

		94		11		826.433		5866.21		0		0				0

		95		11		864.593		5291.01		0		0				0

		96		10		893.005		4862.75		0		0				0

		97		9		914.705		4535.66		0		0				0

		98		9		931.632		4280.53		0		0				0

		99		9		945.072		4077.94		0		0				0

		100		9		-945.072		4077.94		0		0				0

		101		9		-931.632		4280.53		0		0				0

		102		9		-914.705		4535.66		0		0				0

		103		10		-893.005		4862.75		0		0				0

		104		11		-864.593		5291.01		0		0				0

		105		11		-826.433		5866.21		0		0				0

		106		13		-773.517		6663.81		0		0				0

		107		14		-697.001		7817.16		0		0				0

		108		53		-1577.69		40000		0		0				1

		109		52		-826.701		40000		0		0				1

		110		52		0.000212899		40000		0		0				1

		111		52		-3.40E-05		40000		0		0				1

		112		52		953.955		40000		0		0				1

		113		16		622.989		8932.75		0		0				0

		114		13		739.105		7182.52		0		0				0

		115		12		813.325		6063.79		0		0				0

		116		10		863.996		5300.01		0		0				0

		117		10		900.24		4753.69		0		0				0

		118		9		927.085		4349.05		0		0				0

		119		9		947.521		4041.03		0		0				0

		120		8		963.431		3801.2		0		0				0

		121		8		976.054		3610.94		0		0				0

		122		8		-976.054		3610.94		0		0				0

		123		8		-963.431		3801.2		0		0				0

		124		9		-947.521		4041.03		0		0				0

		125		9		-927.085		4349.05		0		0				0

		126		10		-900.24		4753.69		0		0				0

		127		10		-863.996		5300.01		0		0				0

		128		12		-813.325		6063.79		0		0				0

		129		13		-739.105		7182.52		0		0				0

		130		16		-622.989		8932.75		0		0				0

		131		52		-953.955		40000		0		0				1

		132		52		-3.40E-05		40000		0		0				1

		133		52		0.000486442		40000		0		0				1

		134		52		1061.92		40000		0		0				1

		135		15		654.019		8465.03		0		0				0

		136		12		768.684		6736.67		0		0				0

		137		11		840.896		5648.2		0		0				0

		138		10		889.772		4911.49		0		0				0

		139		9		924.546		4387.32		0		0				0

		140		8		950.219		4000.36		0		0				0

		141		8		969.728		3706.29		0		0				0

		142		8		984.9		3477.61		0		0				0

		143		8		996.934		3296.21		0		0				0

		144		8		-996.934		3296.21		0		0				0

		145		8		-984.9		3477.61		0		0				0

		146		8		-969.728		3706.29		0		0				0

		147		8		-950.219		4000.36		0		0				0

		148		9		-924.546		4387.32		0		0				0

		149		10		-889.772		4911.49		0		0				0

		150		11		-840.896		5648.2		0		0				0

		151		12		-768.684		6736.67		0		0				0

		152		15		-654.019		8465.02		0		0				0

		153		52		-1061.92		40000		0		0				1

		154		52		0.000486442		40000		0		0				1

		155		51		0.000125522		40000		0		0				1

		156		51		1160.15		40000		0		0				1

		157		14		677.856		8105.73		0		0				0

		158		12		791.003		6400.25		0		0				0

		159		10		861.49		5337.78		0		0				0

		160		9		908.899		4623.18		0		0				0

		161		9		942.503		4116.66		0		0				0

		162		8		967.261		3743.48		0		0				0

		163		8		986.05		3460.27		0		0				0

		164		7		1000		3237.8		0		0				0

		165		7		1000		3026.98		0		0				0

		166		7		-1000		3026.98		0		0				0

		167		7		-1000		3237.8		0		0				0

		168		8		-986.05		3460.27		0		0				0

		169		8		-967.261		3743.48		0		0				0

		170		9		-942.503		4116.66		0		0				0

		171		9		-908.899		4623.18		0		0				0

		172		10		-861.49		5337.78		0		0				0

		173		12		-791.003		6400.25		0		0				0

		174		14		-677.856		8105.73		0		0				0

		175		51		-1160.15		40000		0		0				1

		176		51		0.000125522		40000		0		0				1

		177		51		-8.42E-05		40000		0		0				1

		178		51		1257.25		40000		0		0				1

		179		14		696.936		7818.14		0		0				0

		180		11		808.636		6134.46		0		0				0

		181		10		877.637		5094.41		0		0				0

		182		9		923.822		4398.25		0		0				0

		183		8		956.47		3906.14		0		0				0

		184		8		980.483		3544.18		0		0				0

		185		7		998.702		3269.56		0		0				0

		186		7		1000		3016.25		0		0				0

		187		7		1000		2816.67		0		0				0

		188		7		-1000		2816.67		0		0				0

		189		7		-1000		3016.25		0		0				0

		190		7		-998.702		3269.56		0		0				0

		191		8		-980.483		3544.18		0		0				0

		192		8		-956.47		3906.14		0		0				0

		193		9		-923.822		4398.25		0		0				0

		194		10		-877.637		5094.41		0		0				0

		195		11		-808.636		6134.46		0		0				0

		196		14		-696.936		7818.14		0		0				0

		197		51		-1257.25		40000		0		0				1

		198		51		-8.42E-05		40000		0		0				1

		199		51		-0.000154068		40000		0		0				1

		200		51		1333.79		40000		0		0				1

		201		13		712.669		7580.98		0		0				0

		202		11		823.027		5917.55		0		0				0

		203		9		890.733		4897		0		0				0

		204		9		935.879		4216.5		0		0				0

		205		8		967.724		3736.5		0		0				0

		206		7		991.126		3383.75		0		0				0

		207		7		1000		3088.28		0		0				0

		208		7		1000		2842.09		0		0				0

		209		6		1000		2651.16		0		0				0

		210		6		-1000		2651.16		0		0				0

		211		7		-1000		2842.09		0		0				0

		212		7		-1000		3088.28		0		0				0

		213		7		-991.126		3383.75		0		0				0

		214		8		-967.724		3736.5		0		0				0

		215		9		-935.879		4216.5		0		0				0

		216		9		-890.733		4897		0		0				0

		217		11		-823.027		5917.55		0		0				0

		218		13		-712.669		7580.98		0		0				0

		219		51		-1333.8		40000		0		0				1

		220		51		-0.000154068		40000		0		0				1

																				%%

																56				25.5

																164				74.5





130 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		53		-8.81E-05		40000		0		0				1

		2		53		386.59		40000		0		0				1

		3		53		760.495		40000		0		0				1

		4		53		1109.9		40000		0		0				1

		5		53		1426.91		40000		0		0				1

		6		16		598.842		9296.72		0		0				0

		7		15		641.979		8646.5		0		0				0

		8		15		675.685		8138.45		0		0				0

		9		14		702.17		7739.25		0		0				0

		10		14		723.242		7421.62		0		0				0

		11		13		740.212		7165.82		0		0				0

		12		13		-740.212		7165.82		0		0				0

		13		14		-723.242		7421.62		0		0				0

		14		14		-702.17		7739.25		0		0				0

		15		15		-675.685		8138.45		0		0				0

		16		15		-641.98		8646.5		0		0				0

		17		16		-598.842		9296.72		0		0				0

		18		53		-1426.91		40000		0		0				1

		19		53		-1109.9		40000		0		0				1

		20		53		-760.495		40000		0		0				1

		21		53		-386.59		40000		0		0				1

		22		53		-8.81E-05		40000		0		0				1

		23		52		-0.000110537		40000		0		0				1

		24		52		535.874		40000		0		0				1

		25		52		1051.57		40000		0		0				1

		26		52		1552.96		40000		0		0				1

		27		15		640.564		8667.84		0		0				0

		28		14		697.8		7805.11		0		0				0

		29		13		740.371		7163.44		0		0				0

		30		12		772.736		6675.59		0		0				0

		31		12		797.814		6297.58		0		0				0

		32		12		817.57		5999.79		0		0				0

		33		11		833.364		5761.74		0		0				0

		34		11		-833.364		5761.74		0		0				0

		35		12		-817.57		5999.79		0		0				0

		36		12		-797.814		6297.58		0		0				0

		37		12		-772.736		6675.59		0		0				0

		38		13		-740.371		7163.44		0		0				0

		39		14		-697.8		7805.11		0		0				0

		40		15		-640.564		8667.84		0		0				0

		41		52		-1552.96		40000		0		0				1

		42		52		-1051.57		40000		0		0				1

		43		52		-535.874		40000		0		0				1

		44		52		-0.000110537		40000		0		0				1

		45		52		-3.05E-05		40000		0		0				1

		46		52		647.267		40000		0		0				1

		47		52		1264.04		40000		0		0				1

		48		16		617.99		9008.11		0		0				0

		49		14		696.85		7819.44		0		0				0

		50		13		752.879		6974.9		0		0				0

		51		12		793.98		6355.38		0		0				0

		52		11		824.935		5888.79		0		0				0

		53		11		848.761		5529.65		0		0				0

		54		10		867.442		5248.07		0		0				0

		55		10		882.326		5023.73		0		0				0

		56		10		-882.326		5023.73		0		0				0

		57		10		-867.442		5248.07		0		0				0

		58		11		-848.761		5529.65		0		0				0

		59		11		-824.935		5888.79		0		0				0

		60		12		-793.98		6355.38		0		0				0

		61		13		-752.879		6974.9		0		0				0

		62		14		-696.85		7819.44		0		0				0

		63		16		-617.99		9008.11		0		0				0

		64		52		-1264.04		40000		0		0				1

		65		52		-647.267		40000		0		0				1

		66		52		-3.05E-05		40000		0		0				1

		67		52		0.000110427		40000		0		0				1

		68		52		739.903		40000		0		0				1

		69		52		1434.02		40000		0		0				1

		70		15		657.096		8418.64		0		0				0

		71		13		735.152		7242.1		0		0				0

		72		12		789.829		6417.95		0		0				0

		73		11		829.589		5818.63		0		0				0

		74		10		859.357		5369.94		0		0				0

		75		10		882.175		5025.99		0		0				0

		76		10		900.016		4757.07		0		0				0

		77		9		914.204		4543.21		0		0				0

		78		9		-914.204		4543.21		0		0				0

		79		10		-900.016		4757.07		0		0				0

		80		10		-882.175		5025.99		0		0				0

		81		10		-859.357		5369.94		0		0				0

		82		11		-829.589		5818.63		0		0				0

		83		12		-789.829		6417.95		0		0				0

		84		13		-735.152		7242.1		0		0				0

		85		15		-657.096		8418.64		0		0				0

		86		52		-1434.02		40000		0		0				1

		87		52		-739.904		40000		0		0				1

		88		52		0.000110427		40000		0		0				1

		89		52		-8.74E-05		40000		0		0				1

		90		52		820.854		40000		0		0				1

		91		52		1600.68		40000		0		0				1

		92		14		686.659		7973.04		0		0				0

		93		12		763.633		6812.8		0		0				0

		94		11		817.041		6007.78		0		0				0

		95		10		855.637		5426.01		0		0				0

		96		10		884.413		4992.26		0		0				0

		97		9		906.41		4660.69		0		0				0

		98		9		923.578		4401.92		0		0				0

		99		9		937.215		4196.37		0		0				0

		100		9		-937.215		4196.37		0		0				0

		101		9		-923.578		4401.92		0		0				0

		102		9		-906.41		4660.69		0		0				0

		103		10		-884.413		4992.26		0		0				0

		104		10		-855.637		5426.01		0		0				0

		105		11		-817.041		6007.78		0		0				0

		106		12		-763.633		6812.8		0		0				0

		107		14		-686.659		7973.04		0		0				0

		108		52		-1600.68		40000		0		0				1

		109		52		-820.854		40000		0		0				1

		110		52		-8.74E-05		40000		0		0				1

		111		51		0.000296893		40000		0		0				1

		112		51		956.918		40000		0		0				1

		113		16		613.005		9083.24		0		0				0

		114		13		729.37		7329.25		0		0				0

		115		11		804.181		6201.62		0		0				0

		116		10		855.405		5429.5		0		0				0

		117		10		892.112		4876.22		0		0				0

		118		9		919.33		4465.96		0		0				0

		119		9		940.064		4153.43		0		0				0

		120		8		956.213		3910.01		0		0				0

		121		8		969.024		3716.91		0		0				0

		122		8		-969.024		3716.91		0		0				0

		123		8		-956.213		3910.01		0		0				0

		124		9		-940.064		4153.43		0		0				0

		125		9		-919.33		4465.96		0		0				0

		126		10		-892.112		4876.22		0		0				0

		127		10		-855.405		5429.5		0		0				0

		128		11		-804.181		6201.62		0		0				0

		129		13		-729.37		7329.25		0		0				0

		130		16		-613.005		9083.24		0		0				0

		131		51		-956.918		40000		0		0				1

		132		51		0.000296893		40000		0		0				1

		133		51		9.24E-05		40000		0		0				1

		134		51		1058.35		40000		0		0				1

		135		15		644.309		8611.39		0		0				0

		136		12		759.487		6875.3		0		0				0

		137		11		832.359		5776.88		0		0				0

		138		10		881.813		5031.46		0		0				0

		139		9		917.055		4500.25		0		0				0

		140		8		943.098		4107.69		0		0				0

		141		8		962.899		3809.23		0		0				0

		142		8		978.301		3577.07		0		0				0

		143		7		990.52		3392.9		0		0				0

		144		7		-990.52		3392.9		0		0				0

		145		8		-978.301		3577.07		0		0				0

		146		8		-962.898		3809.25		0		0				0

		147		8		-943.098		4107.69		0		0				0

		148		9		-917.055		4500.25		0		0				0

		149		10		-881.813		5031.46		0		0				0

		150		11		-832.359		5776.88		0		0				0

		151		12		-759.487		6875.3		0		0				0

		152		15		-644.309		8611.39		0		0				0

		153		51		-1058.35		40000		0		0				1

		154		51		9.24E-05		40000		0		0				1

		155		51		0.000237548		40000		0		0				1

		156		51		1137.03		40000		0		0				1

		157		14		668.5		8246.76		0		0				0

		158		12		782.276		6531.8		0		0				0

		159		10		853.461		5458.81		0		0				0

		160		9		901.456		4735.37		0		0				0

		161		8		935.525		4221.84		0		0				0

		162		8		960.643		3843.23		0		0				0

		163		8		979.716		3555.74		0		0				0

		164		7		994.551		3332.13		0		0				0

		165		7		1000		3134.35		0		0				0

		166		7		-1000		3134.35		0		0				0

		167		7		-994.551		3332.13		0		0				0

		168		8		-979.716		3555.74		0		0				0

		169		8		-960.643		3843.23		0		0				0

		170		8		-935.525		4221.84		0		0				0

		171		9		-901.456		4735.37		0		0				0

		172		10		-853.461		5458.81		0		0				0

		173		12		-782.276		6531.8		0		0				0

		174		14		-668.5		8246.76		0		0				0

		175		51		-1137.03		40000		0		0				1

		176		51		0.000237548		40000		0		0				1

		177		50		0.000135442		40000		0		0				1

		178		50		1213.91		40000		0		0				1

		179		14		687.912		7954.15		0		0				0

		180		11		800.323		6259.77		0		0				0

		181		10		870.041		5208.9		0		0				0

		182		9		916.812		4503.91		0		0				0

		183		8		949.916		4004.92		0		0				0

		184		8		974.283		3637.64		0		0				0

		185		7		992.775		3358.91		0		0				0

		186		7		1000		3119.36		0		0				0

		187		7		1000		2914.58		0		0				0

		188		7		-1000		2914.58		0		0				0

		189		7		-1000		3119.36		0		0				0

		190		7		-992.775		3358.91		0		0				0

		191		8		-974.283		3637.64		0		0				0

		192		8		-949.916		4004.92		0		0				0

		193		9		-916.812		4503.91		0		0				0

		194		10		-870.041		5208.9		0		0				0

		195		11		-800.323		6259.77		0		0				0

		196		14		-687.912		7954.15		0		0				0

		197		50		-1213.91		40000		0		0				1

		198		50		0.000135442		40000		0		0				1

		199		50		6.24E-07		40000		0		0				1

		200		50		1286.78		40000		0		0				1

		201		13		703.961		7712.24		0		0				0

		202		11		815.08		6037.33		0		0				0

		203		9		883.513		5005.83		0		0				0

		204		8		929.24		4316.58		0		0				0

		205		8		961.532		3829.84		0		0				0

		206		7		985.277		3471.93		0		0				0

		207		7		1000		3189.82		0		0				0

		208		7		1000		2937.28		0		0				0

		209		6		1000		2741.53		0		0				0

		210		6		-1000		2741.53		0		0				0

		211		7		-1000		2937.28		0		0				0

		212		7		-1000		3189.82		0		0				0

		213		7		-985.277		3471.93		0		0				0

		214		8		-961.532		3829.84		0		0				0

		215		8		-929.24		4316.58		0		0				0

		216		9		-883.513		5005.83		0		0				0

		217		11		-815.08		6037.33		0		0				0

		218		13		-703.961		7712.24		0		0				0

		219		50		-1286.78		40000		0		0				1

		220		50		6.24E-07		40000		0		0				1

																				%%

																56				25.5

																164				74.5





140 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		52		-0.000103706		40000		0		0				1

		2		52		383.808		40000		0		0				1

		3		52		759.521		40000		0		0				1

		4		52		1124.71		40000		0		0				1

		5		52		1497.59		40000		0		0				1

		6		52		1886.27		40000		0		0				1

		7		15		631.254		8808.18		0		0				0

		8		14		664.931		8300.55		0		0				0

		9		14		691.489		7900.23		0		0				0

		10		13		712.649		7581.29		0		0				0

		11		13		729.705		7324.2		0		0				0

		12		13		-729.705		7324.2		0		0				0

		13		13		-712.649		7581.29		0		0				0

		14		14		-691.489		7900.23		0		0				0

		15		14		-664.931		8300.55		0		0				0

		16		15		-631.254		8808.18		0		0				0

		17		16		-590.36		9424.57		0		0				0

		18		52		-1497.59		40000		0		0				1

		19		52		-1124.71		40000		0		0				1

		20		52		-759.522		40000		0		0				1

		21		52		-383.808		40000		0		0				1

		22		52		-0.000103706		40000		0		0				1

		23		52		-0.000238948		40000		0		0				1

		24		52		532.383		40000		0		0				1

		25		52		1058.4		40000		0		0				1

		26		52		1614.76		40000		0		0				1

		27		15		630.199		8824.08		0		0				0

		28		14		687.463		7960.92		0		0				0

		29		13		730.237		7316.18		0		0				0

		30		12		762.816		6825.12		0		0				0

		31		12		788.092		6444.12		0		0				0

		32		11		808.024		6143.69		0		0				0

		33		11		823.968		5903.37		0		0				0

		34		11		-823.968		5903.37		0		0				0

		35		11		-808.024		6143.69		0		0				0

		36		12		-788.092		6444.12		0		0				0

		37		12		-762.816		6825.12		0		0				0

		38		13		-730.237		7316.18		0		0				0

		39		14		-687.463		7960.92		0		0				0

		40		15		-630.199		8824.08		0		0				0

		41		52		-1614.76		40000		0		0				1

		42		52		-1058.4		40000		0		0				1

		43		52		-532.384		40000		0		0				1

		44		52		-0.000238948		40000		0		0				1

		45		51		0.000291365		40000		0		0				1

		46		51		640.05		40000		0		0				1

		47		52		1261.52		40000		0		0				1

		48		15		607.96		9159.28		0		0				0

		49		14		686.779		7971.23		0		0				0

		50		13		743.096		7122.36		0		0				0

		51		12		784.503		6498.22		0		0				0

		52		11		815.74		6027.38		0		0				0

		53		11		839.811		5664.56		0		0				0

		54		10		858.698		5379.87		0		0				0

		55		10		873.754		5152.93		0		0				0

		56		10		-873.754		5152.93		0		0				0

		57		10		-858.698		5379.87		0		0				0

		58		11		-839.811		5664.56		0		0				0

		59		11		-815.74		6027.38		0		0				0

		60		12		-784.503		6498.22		0		0				0

		61		13		-743.096		7122.36		0		0				0

		62		14		-686.78		7971.23		0		0				0

		63		15		-607.96		9159.28		0		0				0

		64		52		-1261.52		40000		0		0				1

		65		51		-640.05		40000		0		0				1

		66		51		0.000291365		40000		0		0				1

		67		51		6.62E-05		40000		0		0				1

		68		51		726.035		40000		0		0				1

		69		51		1397.97		40000		0		0				1

		70		15		647.16		8568.41		0		0				0

		71		13		725.49		7387.74		0		0				0

		72		12		780.546		6557.87		0		0				0

		73		11		820.669		5953.09		0		0				0

		74		10		850.752		5499.64		0		0				0

		75		10		873.835		5151.7		0		0				0

		76		9		891.895		4879.49		0		0				0

		77		9		906.263		4662.92		0		0				0

		78		9		-906.263		4662.92		0		0				0

		79		9		-891.895		4879.49		0		0				0

		80		10		-873.835		5151.7		0		0				0

		81		10		-850.752		5499.64		0		0				0

		82		11		-820.669		5953.09		0		0				0

		83		12		-780.546		6557.87		0		0				0

		84		13		-725.49		7387.74		0		0				0

		85		15		-647.16		8568.41		0		0				0

		86		51		-1397.97		40000		0		0				1

		87		51		-726.035		40000		0		0				1

		88		51		6.62E-05		40000		0		0				1

		89		51		-3.30E-05		40000		0		0				1

		90		51		799.005		40000		0		0				1

		91		51		1516.31		40000		0		0				1

		92		14		676.983		8118.89		0		0				0

		93		12		754.351		6952.71		0		0				0

		94		11		808.196		6141.09		0		0				0

		95		10		847.187		5553.38		0		0				0

		96		10		876.296		5114.61		0		0				0

		97		9		898.567		4778.92		0		0				0

		98		9		915.958		4516.79		0		0				0

		99		9		929.776		4308.49		0		0				0

		100		9		-929.776		4308.49		0		0				0

		101		9		-915.958		4516.79		0		0				0

		102		9		-898.567		4778.92		0		0				0

		103		10		-876.296		5114.61		0		0				0

		104		10		-847.187		5553.38		0		0				0

		105		11		-808.196		6141.09		0		0				0

		106		12		-754.351		6952.71		0		0				0

		107		14		-676.983		8118.89		0		0				0

		108		51		-1516.31		40000		0		0				1

		109		51		-799.005		40000		0		0				1

		110		51		-3.30E-05		40000		0		0				1

		111		51		-0.000202432		40000		0		0				1

		112		51		923.364		40000		0		0				1

		113		15		603.968		9219.45		0		0				0

		114		13		720.226		7467.08		0		0				0

		115		11		795.558		6331.58		0		0				0

		116		10		847.285		5551.91		0		0				0

		117		9		884.416		4992.22		0		0				0

		118		9		911.979		4576.75		0		0				0

		119		8		932.991		4260.04		0		0				0

		120		8		949.362		4013.27		0		0				0

		121		8		962.352		3817.48		0		0				0

		122		8		-962.352		3817.48		0		0				0

		123		8		-949.362		4013.27		0		0				0

		124		8		-932.991		4260.04		0		0				0

		125		9		-911.979		4576.75		0		0				0

		126		9		-884.416		4992.22		0		0				0

		127		10		-847.285		5551.9		0		0				0

		128		11		-795.558		6331.58		0		0				0

		129		13		-720.226		7467.08		0		0				0

		130		15		-603.968		9219.45		0		0				0

		131		51		-923.365		40000		0		0				1

		132		51		-0.000202432		40000		0		0				1

		133		50		-0.000244649		40000		0		0				1

		134		50		1030.11		40000		0		0				1

		135		14		635.224		8748.33		0		0				0

		136		12		750.819		7005.95		0		0				0

		137		11		824.284		5898.6		0		0				0

		138		10		874.268		5145.18		0		0				0

		139		9		909.943		4607.44		0		0				0

		140		8		936.332		4209.67		0		0				0

		141		8		956.404		3907.13		0		0				0

		142		8		972.024		3671.68		0		0				0

		143		7		984.416		3484.9		0		0				0

		144		7		-984.416		3484.9		0		0				0

		145		8		-972.024		3671.68		0		0				0

		146		8		-956.404		3907.13		0		0				0

		147		8		-936.332		4209.67		0		0				0

		148		9		-909.944		4607.44		0		0				0

		149		10		-874.268		5145.18		0		0				0

		150		11		-824.284		5898.6		0		0				0

		151		12		-750.819		7005.95		0		0				0

		152		14		-635.224		8748.33		0		0				0

		153		50		-1030.11		40000		0		0				1

		154		50		-0.000244649		40000		0		0				1

		155		50		-0.000242314		40000		0		0				1

		156		50		1129.27		40000		0		0				1

		157		14		659.701		8379.39		0		0				0

		158		11		774.027		6656.12		0		0				0

		159		10		845.847		5573.57		0		0				0

		160		9		894.384		4841.97		0		0				0

		161		8		928.886		4321.92		0		0				0

		162		8		954.344		3938.19		0		0				0

		163		7		973.683		3646.69		0		0				0

		164		7		988.726		3419.93		0		0				0

		165		7		1000		3238.02		0		0				0

		166		7		-1000		3238.02		0		0				0

		167		7		-988.726		3419.93		0		0				0

		168		7		-973.683		3646.69		0		0				0

		169		8		-954.344		3938.19		0		0				0

		170		8		-928.886		4321.92		0		0				0

		171		9		-894.384		4841.97		0		0				0

		172		10		-845.847		5573.57		0		0				0

		173		11		-774.027		6656.12		0		0				0

		174		14		-659.701		8379.39		0		0				0

		175		50		-1129.27		40000		0		0				1

		176		50		-0.000242314		40000		0		0				1

		177		50		6.38E-05		40000		0		0				1

		178		50		1231.93		40000		0		0				1

		179		13		679.411		8082.3		0		0				0

		180		11		792.447		6378.48		0		0				0

		181		10		862.822		5317.71		0		0				0

		182		9		910.137		4604.52		0		0				0

		183		8		943.669		4099.08		0		0				0

		184		8		968.368		3726.8		0		0				0

		185		7		987.118		3444.17		0		0				0

		186		7		1000		3218.92		0		0				0

		187		7		1000		3009.07		0		0				0

		188		7		-1000		3009.07		0		0				0

		189		7		-1000		3218.92		0		0				0

		190		7		-987.118		3444.17		0		0				0

		191		8		-968.368		3726.8		0		0				0

		192		8		-943.669		4099.09		0		0				0

		193		9		-910.137		4604.52		0		0				0

		194		10		-862.822		5317.71		0		0				0

		195		11		-792.447		6378.48		0		0				0

		196		13		-679.411		8082.3		0		0				0

		197		50		-1231.93		40000		0		0				1

		198		50		6.38E-05		40000		0		0				1

		199		50		-0.000175168		40000		0		0				1

		200		50		1326.71		40000		0		0				1

		201		13		695.735		7836.23		0		0				0

		202		11		807.536		6151.04		0		0				0

		203		9		876.638		5109.45		0		0				0

		204		8		922.907		4412.03		0		0				0

		205		8		955.619		3918.96		0		0				0

		206		7		979.688		3556.17		0		0				0

		207		7		997.941		3281.03		0		0				0

		208		7		1000		3029.25		0		0				0

		209		6		1000		2828.82		0		0				0

		210		6		-1000		2828.82		0		0				0

		211		7		-1000		3029.25		0		0				0

		212		7		-997.941		3281.03		0		0				0

		213		7		-979.688		3556.17		0		0				0

		214		8		-955.619		3918.96		0		0				0

		215		8		-922.907		4412.03		0		0				0

		216		9		-876.638		5109.45		0		0				0

		217		11		-807.536		6151.04		0		0				0

		218		13		-695.735		7836.23		0		0				0

		219		50		-1326.71		40000		0		0				1

		220		50		-0.000175168		40000		0		0				1

																				%%

																57				25.9

																163				74.1





150 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		51		0.000248358		40000		0		0				1

		2		51		371.633		40000		0		0				1

		3		51		730.136		40000		0		0				1

		4		51		1063.87		40000		0		0				1

		5		51		1368.4		40000		0		0				1

		6		51		1638.36		40000		0		0				1

		7		15		621.189		8959.88		0		0				0

		8		14		654.892		8451.87		0		0				0

		9		14		681.494		8050.9		0		0				0

		10		13		702.723		7730.91		0		0				0

		11		13		719.849		7472.77		0		0				0

		12		13		-719.849		7472.77		0		0				0

		13		13		-702.723		7730.91		0		0				0

		14		14		-681.494		8050.9		0		0				0

		15		14		-654.892		8451.87		0		0				0

		16		15		-621.189		8959.88		0		0				0

		17		51		-1638.36		40000		0		0				1

		18		51		-1368.41		40000		0		0				1

		19		51		-1063.87		40000		0		0				1

		20		51		-730.137		40000		0		0				1

		21		51		-371.633		40000		0		0				1

		22		51		0.000248358		40000		0		0				1

		23		51		-0.000289702		40000		0		0				1

		24		51		516.068		40000		0		0				1

		25		51		1010.45		40000		0		0				1

		26		51		1482.78		40000		0		0				1

		27		15		620.43		8971.32		0		0				0

		28		14		677.781		8106.86		0		0				0

		29		13		720.722		7459.61		0		0				0

		30		12		753.485		6965.76		0		0				0

		31		12		778.937		6582.12		0		0				0

		32		11		799.025		6279.34		0		0				0

		33		11		815.104		6036.97		0		0				0

		34		11		-815.104		6036.97		0		0				0

		35		11		-799.025		6279.34		0		0				0

		36		12		-778.937		6582.12		0		0				0

		37		12		-753.485		6965.76		0		0				0

		38		13		-720.722		7459.61		0		0				0

		39		14		-677.781		8106.86		0		0				0

		40		15		-620.43		8971.32		0		0				0

		41		51		-1482.78		40000		0		0				1

		42		51		-1010.45		40000		0		0				1

		43		51		-516.068		40000		0		0				1

		44		51		-0.000289702		40000		0		0				1

		45		51		5.30E-05		40000		0		0				1

		46		51		624.497		40000		0		0				1

		47		51		1215.58		40000		0		0				1

		48		15		599.162		9291.9		0		0				0

		49		14		677.337		8113.55		0		0				0

		50		12		733.894		7261.06		0		0				0

		51		12		775.572		6632.84		0		0				0

		52		11		807.062		6158.19		0		0				0

		53		10		831.353		5792.04		0		0				0

		54		10		850.428		5504.52		0		0				0

		55		10		865.642		5275.21		0		0				0

		56		10		-865.642		5275.21		0		0				0

		57		10		-850.428		5504.52		0		0				0

		58		10		-831.353		5792.04		0		0				0

		59		11		-807.062		6158.19		0		0				0

		60		12		-775.572		6632.84		0		0				0

		61		12		-733.894		7261.06		0		0				0

		62		14		-677.337		8113.55		0		0				0

		63		15		-599.162		9291.9		0		0				0

		64		51		-1215.58		40000		0		0				1

		65		51		-624.497		40000		0		0				1

		66		51		5.30E-05		40000		0		0				1

		67		51		0.000225363		40000		0		0				1

		68		51		715.285		40000		0		0				1

		69		51		1376.86		40000		0		0				1

		70		14		637.875		8708.37		0		0				0

		71		13		716.401		7524.74		0		0				0

		72		12		771.79		6689.84		0		0				0

		73		11		812.239		6080.15		0		0				0

		74		10		842.61		5622.37		0		0				0

		75		10		865.935		5270.78		0		0				0

		76		9		884.196		4995.53		0		0				0

		77		9		898.73		4776.46		0		0				0

		78		9		-898.73		4776.45		0		0				0

		79		9		-884.196		4995.53		0		0				0

		80		10		-865.935		5270.78		0		0				0

		81		10		-842.61		5622.37		0		0				0

		82		11		-812.239		6080.15		0		0				0

		83		12		-771.79		6689.84		0		0				0

		84		13		-716.401		7524.74		0		0				0

		85		14		-637.875		8708.37		0		0				0

		86		51		-1376.86		40000		0		0				1

		87		51		-715.285		40000		0		0				1

		88		51		0.000225363		40000		0		0				1

		89		50		0.000407467		40000		0		0				1

		90		50		795.529		40000		0		0				1

		91		50		1525.86		40000		0		0				1

		92		14		667.9		8255.8		0		0				0

		93		12		745.599		7084.63		0		0				0

		94		11		799.836		6267.11		0		0				0

		95		10		839.185		5673.99		0		0				0

		96		10		868.6		5230.62		0		0				0

		97		9		891.124		4891.11		0		0				0

		98		9		908.721		4625.86		0		0				0

		99		9		922.709		4415.02		0		0				0

		100		9		-922.709		4415.01		0		0				0

		101		9		-908.721		4625.86		0		0				0

		102		9		-891.124		4891.11		0		0				0

		103		10		-868.6		5230.62		0		0				0

		104		10		-839.185		5673.99		0		0				0

		105		11		-799.836		6267.11		0		0				0

		106		12		-745.599		7084.63		0		0				0

		107		14		-667.9		8255.8		0		0				0

		108		50		-1525.86		40000		0		0				1

		109		50		-795.529		40000		0		0				1

		110		50		0.000407467		40000		0		0				1

		111		50		0.00027474		40000		0		0				1

		112		50		934.827		40000		0		0				1

		113		15		595.147		9352.42		0		0				0

		114		13		711.602		7597.07		0		0				0

		115		11		787.399		6454.57		0		0				0

		116		10		839.582		5668		0		0				0

		117		9		877.105		5102.42		0		0				0

		118		9		904.989		4682.12		0		0				0

		119		8		926.26		4361.5		0		0				0

		120		8		942.839		4111.6		0		0				0

		121		8		955.997		3913.26		0		0				0

		122		8		-955.997		3913.26		0		0				0

		123		8		-942.839		4111.6		0		0				0

		124		8		-926.26		4361.5		0		0				0

		125		9		-904.989		4682.12		0		0				0

		126		9		-877.105		5102.42		0		0				0

		127		10		-839.582		5668		0		0				0

		128		11		-787.399		6454.57		0		0				0

		129		13		-711.602		7597.07		0		0				0

		130		15		-595.147		9352.42		0		0				0

		131		50		-934.827		40000		0		0				1

		132		50		0.00027474		40000		0		0				1

		133		50		6.18E-05		40000		0		0				1

		134		50		1044.33		40000		0		0				1

		135		14		626.763		8875.87		0		0				0

		136		12		742.619		7129.55		0		0				0

		137		10		816.621		6014.1		0		0				0

		138		9		867.093		5253.33		0		0				0

		139		9		903.171		4709.52		0		0				0

		140		8		929.883		4306.89		0		0				0

		141		8		950.211		4000.47		0		0				0

		142		8		966.036		3761.95		0		0				0

		143		7		978.591		3572.7		0		0				0

		144		7		-978.591		3572.7		0		0				0

		145		8		-966.036		3761.95		0		0				0

		146		8		-950.211		4000.47		0		0				0

		147		8		-929.883		4306.89		0		0				0

		148		9		-903.171		4709.52		0		0				0

		149		9		-867.093		5253.33		0		0				0

		150		10		-816.621		6014.1		0		0				0

		151		12		-742.619		7129.55		0		0				0

		152		14		-626.763		8875.87		0		0				0

		153		50		-1044.33		40000		0		0				1

		154		50		6.18E-05		40000		0		0				1

		155		50		0.000107344		40000		0		0				1

		156		50		1119.31		40000		0		0				1

		157		14		651.407		8504.41		0		0				0

		158		11		766.206		6774.01		0		0				0

		159		10		838.605		5682.74		0		0				0

		160		9		887.643		4943.57		0		0				0

		161		8		922.55		4417.42		0		0				0

		162		8		948.327		4028.87		0		0				0

		163		7		967.918		3733.57		0		0				0

		164		7		983.16		3503.84		0		0				0

		165		7		995.242		3321.72		0		0				0

		166		7		-995.242		3321.72		0		0				0

		167		7		-983.16		3503.84		0		0				0

		168		7		-967.918		3733.57		0		0				0

		169		8		-948.327		4028.87		0		0				0

		170		8		-922.55		4417.42		0		0				0

		171		9		-887.643		4943.57		0		0				0

		172		10		-838.605		5682.74		0		0				0

		173		11		-766.206		6774.01		0		0				0

		174		14		-651.407		8504.41		0		0				0

		175		50		-1119.31		40000		0		0				1

		176		50		0.000107344		40000		0		0				1

		177		49		-0.000145749		40000		0		0				1

		178		49		1178.14		40000		0		0				1

		179		13		671.376		8203.41		0		0				0

		180		11		784.962		6491.3		0		0				0

		181		9		855.942		5421.41		0		0				0

		182		9		903.764		4700.58		0		0				0

		183		8		937.697		4189.11		0		0				0

		184		7		962.71		3812.07		0		0				0

		185		7		981.704		3525.77		0		0				0

		186		7		996.469		3303.23		0		0				0

		187		7		1000		3100.49		0		0				0

		188		7		-1000		3100.49		0		0				0

		189		7		-996.469		3303.23		0		0				0

		190		7		-981.704		3525.77		0		0				0

		191		7		-962.71		3812.07		0		0				0

		192		8		-937.697		4189.11		0		0				0

		193		9		-903.764		4700.58		0		0				0

		194		9		-855.942		5421.41		0		0				0

		195		11		-784.962		6491.3		0		0				0

		196		13		-671.376		8203.41		0		0				0

		197		49		-1178.14		40000		0		0				1

		198		49		-0.000145749		40000		0		0				1

		199		49		-2.11E-04		40000		0		0				1

		200		49		1242.87		40000		0		0				1

		201		13		687.942		7953.7		0		0				0

		202		10		800.354		6259.31		0		0				0

		203		9		870.075		5208.38		0		0				0

		204		8		916.85		4503.33		0		0				0

		205		8		949.958		4004.28		0		0				0

		206		7		974.333		3636.88		0		0				0

		207		7		992.824		3358.16		0		0				0

		208		7		1000		3118.32		0		0				0

		209		6		1000		2913.33		0		0				0

		210		6		-1000		2913.33		0		0				0

		211		7		-1000		3118.32		0		0				0

		212		7		-992.824		3358.16		0		0				0

		213		7		-974.333		3636.88		0		0				0

		214		8		-949.958		4004.28		0		0				0

		215		8		-916.85		4503.33		0		0				0

		216		9		-870.075		5208.38		0		0				0

		217		10		-800.354		6259.31		0		0				0

		218		13		-687.942		7953.7		0		0				0

		219		49		-1242.87		40000		0		0				1

		220		49		-2.11E-04		40000		0		0				1

																				%%

																58				26.4

																162				73.6





160 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		51		-0.000142831		40000		0		0				1

		2		51		371.088		40000		0		0				1

		3		51		733.802		40000		0		0				1

		4		51		1082.09		40000		0		0				1

		5		51		1410.28		40000		0		0				1

		6		50		1676.3		40000		0		0				1

		7		15		611.961		9098.97		0		0				0

		8		14		645.541		8592.82		0		0				0

		9		14		672.148		8191.76		0		0				0

		10		13		693.434		7870.92		0		0				0

		11		13		710.618		7611.9		0		0				0

		12		13		-710.618		7611.9		0		0				0

		13		13		-693.434		7870.92		0		0				0

		14		14		-672.148		8191.76		0		0				0

		15		14		-645.541		8592.82		0		0				0

		16		15		-611.961		9098.97		0		0				0

		17		50		-1676.3		40000		0		0				1

		18		51		-1410.28		40000		0		0				1

		19		51		-1082.09		40000		0		0				1

		20		51		-733.802		40000		0		0				1

		21		51		-371.088		40000		0		0				1

		22		51		-0.000142831		40000		0		0				1

		23		50		-0.000257122		40000		0		0				1

		24		50		517.563		40000		0		0				1

		25		50		1033.82		40000		0		0				1

		26		50		1607.51		40000		0		0				1

		27		15		611.479		9106.24		0		0				0

		28		13		668.7		8243.75		0		0				0

		29		13		711.777		7594.43		0		0				0

		30		12		744.703		7098.14		0		0				0

		31		11		770.311		6712.15		0		0				0

		32		11		790.539		6407.24		0		0				0

		33		11		806.742		6163.02		0		0				0

		34		11		-806.742		6163.01		0		0				0

		35		11		-790.539		6407.24		0		0				0

		36		11		-770.311		6712.15		0		0				0

		37		12		-744.703		7098.14		0		0				0

		38		13		-711.777		7594.43		0		0				0

		39		13		-668.7		8243.75		0		0				0

		40		15		-611.479		9106.24		0		0				0

		41		50		-1607.51		40000		0		0				1

		42		50		-1033.82		40000		0		0				1

		43		50		-517.563		40000		0		0				1

		44		50		-0.000257122		40000		0		0				1

		45		50		0.000472784		40000		0		0				1

		46		50		624.738		40000		0		0				1

		47		50		1258.57		40000		0		0				1

		48		15		592.433		9393.33		0		0				0

		49		13		668.464		8247.3		0		0				0

		50		12		725.221		7391.79		0		0				0

		51		11		767.139		6759.95		0		0				0

		52		11		798.857		6281.86		0		0				0

		53		10		823.35		5912.68		0		0				0

		54		10		842.597		5622.56		0		0				0

		55		10		857.957		5391.04		0		0				0

		56		10		-857.957		5391.04		0		0				0

		57		10		-842.597		5622.56		0		0				0

		58		10		-823.35		5912.68		0		0				0

		59		11		-798.857		6281.86		0		0				0

		60		11		-767.139		6759.95		0		0				0

		61		12		-725.221		7391.79		0		0				0

		62		13		-668.464		8247.3		0		0				0

		63		15		-592.433		9393.33		0		0				0

		64		50		-1258.57		40000		0		0				1

		65		50		-624.738		40000		0		0				1

		66		50		0.000472784		40000		0		0				1

		67		50		0.000236485		40000		0		0				1

		68		50		709.643		40000		0		0				1

		69		50		1409.38		40000		0		0				1

		70		14		629.244		8838.46		0		0				0

		71		13		707.826		7653.98		0		0				0

		72		11		763.51		6814.65		0		0				0

		73		11		804.255		6200.5		0		0				0

		74		10		834.888		5738.77		0		0				0

		75		10		858.437		5383.8		0		0				0

		76		9		876.885		5105.73		0		0				0

		77		9		891.574		4884.32		0		0				0

		78		9		-891.574		4884.32		0		0				0

		79		9		-876.885		5105.73		0		0				0

		80		10		-858.437		5383.8		0		0				0

		81		10		-834.888		5738.77		0		0				0

		82		11		-804.255		6200.5		0		0				0

		83		11		-763.511		6814.65		0		0				0

		84		13		-707.826		7653.98		0		0				0

		85		14		-629.245		8838.46		0		0				0

		86		50		-1409.38		40000		0		0				1

		87		50		-709.643		40000		0		0				1

		88		50		0.000236485		40000		0		0				1

		89		50		4.82E-05		40000		0		0				1

		90		50		779.649		40000		0		0				1

		91		50		1483.43		40000		0		0				1

		92		13		659.34		8384.82		0		0				0

		93		12		737.32		7209.42		0		0				0

		94		11		791.909		6386.59		0		0				0

		95		10		831.587		5788.52		0		0				0

		96		10		861.284		5340.89		0		0				0

		97		9		884.043		4997.84		0		0				0

		98		9		901.834		4729.67		0		0				0

		99		8		915.98		4516.44		0		0				0

		100		8		-915.98		4516.44		0		0				0

		101		9		-901.834		4729.67		0		0				0

		102		9		-884.043		4997.84		0		0				0

		103		10		-861.284		5340.89		0		0				0

		104		10		-831.587		5788.52		0		0				0

		105		11		-791.909		6386.59		0		0				0

		106		12		-737.32		7209.42		0		0				0

		107		13		-659.34		8384.82		0		0				0

		108		50		-1483.43		40000		0		0				1

		109		50		-779.649		40000		0		0				1

		110		50		4.82E-05		40000		0		0				1

		111		50		-0.000290309		40000		0		0				1

		112		50		893.39		40000		0		0				1

		113		50		1602.38		40000		0		0				1

		114		12		703.44		7720.11		0		0				0

		115		11		779.65		6571.37		0		0				0

		116		10		832.252		5778.5		0		0				0

		117		9		870.137		5207.45		0		0				0

		118		9		898.321		4782.63		0		0				0

		119		8		919.834		4458.35		0		0				0

		120		8		936.609		4205.5		0		0				0

		121		8		949.926		4004.77		0		0				0

		122		8		-949.926		4004.77		0		0				0

		123		8		-936.609		4205.51		0		0				0

		124		8		-919.834		4458.35		0		0				0

		125		9		-898.321		4782.63		0		0				0

		126		9		-870.137		5207.45		0		0				0

		127		10		-832.252		5778.5		0		0				0

		128		11		-779.65		6571.37		0		0				0

		129		12		-703.44		7720.11		0		0				0

		130		50		-1602.38		40000		0		0				1

		131		50		-893.39		40000		0		0				1

		132		50		-0.000290309		40000		0		0				1

		133		49		0.00037362		40000		0		0				1

		134		49		991.472		40000		0		0				1

		135		14		618.896		8994.44		0		0				0

		136		12		734.833		7246.91		0		0				0

		137		10		809.322		6124.13		0		0				0

		138		9		860.245		5356.56		0		0				0

		139		9		896.699		4807.07		0		0				0

		140		8		923.714		4399.88		0		0				0

		141		8		944.284		4089.81		0		0				0

		142		7		960.302		3848.37		0		0				0

		143		7		973.013		3656.79		0		0				0

		144		7		-973.013		3656.79		0		0				0

		145		7		-960.302		3848.37		0		0				0

		146		8		-944.284		4089.81		0		0				0

		147		8		-923.714		4399.88		0		0				0

		148		9		-896.699		4807.07		0		0				0

		149		9		-860.245		5356.56		0		0				0

		150		10		-809.322		6124.13		0		0				0

		151		12		-734.833		7246.91		0		0				0

		152		14		-618.896		8994.44		0		0				0

		153		49		-991.472		40000		0		0				1

		154		49		0.00037362		40000		0		0				1

		155		49		-8.29E-05		40000		0		0				1

		156		49		1080.15		40000		0		0				1

		157		13		643.563		8622.64		0		0				0

		158		11		758.76		6886.25		0		0				0

		159		10		831.69		5786.97		0		0				0

		160		9		881.196		5040.76		0		0				0

		161		8		916.482		4508.88		0		0				0

		162		8		942.56		4115.8		0		0				0

		163		7		962.39		3816.91		0		0				0

		164		7		977.819		3584.34		0		0				0

		165		7		990.052		3399.95		0		0				0

		166		7		-990.052		3399.95		0		0				0

		167		7		-977.819		3584.34		0		0				0

		168		7		-962.39		3816.9		0		0				0

		169		8		-942.56		4115.8		0		0				0

		170		8		-916.482		4508.88		0		0				0

		171		9		-881.196		5040.76		0		0				0

		172		10		-831.69		5786.97		0		0				0

		173		11		-758.76		6886.25		0		0				0

		174		13		-643.563		8622.64		0		0				0

		175		49		-1080.15		40000		0		0				1

		176		49		-8.29E-05		40000		0		0				1

		177		49		0.000123139		40000		0		0				1

		178		49		1160.19		40000		0		0				1

		179		13		663.737		8318.55		0		0				0

		180		11		777.822		6598.93		0		0				0

		181		9		849.36		5520.63		0		0				0

		182		8		897.656		4792.64		0		0				0

		183		8		931.967		4275.47		0		0				0

		184		7		957.278		3893.96		0		0				0

		185		7		976.503		3604.17		0		0				0

		186		7		991.451		3378.86		0		0				0

		187		7		1000		3189.24		0		0				0

		188		7		-1000		3189.24		0		0				0

		189		7		-991.451		3378.86		0		0				0

		190		7		-976.503		3604.17		0		0				0

		191		7		-957.278		3893.96		0		0				0

		192		8		-931.967		4275.47		0		0				0

		193		8		-897.656		4792.64		0		0				0

		194		9		-849.36		5520.63		0		0				0

		195		11		-777.822		6598.93		0		0				0

		196		13		-663.737		8318.55		0		0				0

		197		49		-1160.19		40000		0		0				1

		198		49		0.000123139		40000		0		0				1

		199		48		-0.000301766		40000		0		0				1

		200		48		1243.26		40000		0		0				1

		201		13		680.527		8065.47		0		0				0

		202		10		793.488		6362.79		0		0				0

		203		9		863.785		5303.2		0		0				0

		204		8		911.036		4590.97		0		0				0

		205		8		944.518		4086.29		0		0				0

		206		7		969.184		3714.5		0		0				0

		207		7		987.901		3432.37		0		0				0

		208		7		1000		3204.88		0		0				0

		209		6		1000		2995.46		0		0				0

		210		6		-1000		2995.46		0		0				0

		211		7		-1000		3204.88		0		0				0

		212		7		-987.901		3432.37		0		0				0

		213		7		-969.184		3714.5		0		0				0

		214		8		-944.518		4086.29		0		0				0

		215		8		-911.036		4590.97		0		0				0

		216		9		-863.785		5303.2		0		0				0

		217		10		-793.488		6362.79		0		0				0

		218		13		-680.527		8065.47		0		0				0

		219		48		-1243.26		40000		0		0				1

		220		48		-0.000301766		40000		0		0				1

																				%%

																60				27.3

																160				72.7





170 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-------------------		---------------		-----------------		----------------

		1		50		-2.68E-05		40000		0		0				1

		2		50		359.24		40000		0		0				1

		3		50		706.328		40000		0		0				1

		4		50		1027.66		40000		0		0				1

		5		50		1301.32		40000		0		0				1

		6		50		1488.17		40000		0		0				1

		7		15		603.362		9228.59		0		0				0

		8		14		636.665		8726.6		0		0				0

		9		13		663.324		8324.78		0		0				0

		10		13		684.642		8003.45		0		0				0

		11		13		701.877		7743.65		0		0				0

		12		13		-701.877		7743.65		0		0				0

		13		13		-684.642		8003.45		0		0				0

		14		13		-663.324		8324.78		0		0				0

		15		14		-636.666		8726.6		0		0				0

		16		15		-603.358		9228.66		0		0				0

		17		50		-1488.17		40000		0		0				1

		18		50		-1301.32		40000		0		0				1

		19		50		-1027.66		40000		0		0				1

		20		50		-706.328		40000		0		0				1

		21		50		-359.24		40000		0		0				1

		22		50		-2.68E-05		40000		0		0				1

		23		50		-0.000277596		40000		0		0				1

		24		50		496.317		40000		0		0				1

		25		50		969.248		40000		0		0				1

		26		50		1383.13		40000		0		0				1

		27		14		603.195		9231.11		0		0				0

		28		13		660.133		8372.87		0		0				0

		29		12		703.313		7722.02		0		0				0

		30		12		736.378		7223.62		0		0				0

		31		11		762.124		6835.54		0		0				0

		32		11		782.479		6528.74		0		0				0

		33		11		798.792		6282.84		0		0				0

		34		11		-798.792		6282.84		0		0				0

		35		11		-782.479		6528.74		0		0				0

		36		11		-762.124		6835.54		0		0				0

		37		12		-736.378		7223.62		0		0				0

		38		12		-703.313		7722.02		0		0				0

		39		13		-660.133		8372.87		0		0				0

		40		14		-603.195		9231.11		0		0				0

		41		50		-1383.13		40000		0		0				1

		42		50		-969.248		40000		0		0				1

		43		50		-496.318		40000		0		0				1

		44		50		-0.000277596		40000		0		0				1

		45		50		9.10E-05		40000		0		0				1

		46		50		597.821		40000		0		0				1

		47		50		1164.62		40000		0		0				1

		48		50		1711		40000		0		0				1

		49		13		660.091		8373.5		0		0				0

		50		12		717.006		7515.62		0		0				0

		51		11		759.136		6880.59		0		0				0

		52		11		791.059		6399.4		0		0				0

		53		10		815.736		6027.44		0		0				0

		54		10		835.142		5734.94		0		0				0

		55		10		850.635		5501.41		0		0				0

		56		10		-850.635		5501.41		0		0				0

		57		10		-835.142		5734.94		0		0				0

		58		10		-815.736		6027.44		0		0				0

		59		11		-791.059		6399.4		0		0				0

		60		11		-759.136		6880.59		0		0				0

		61		12		-717.006		7515.62		0		0				0

		62		13		-660.091		8373.5		0		0				0

		63		50		-1711		40000		0		0				1

		64		50		-1164.62		40000		0		0				1

		65		50		-597.821		40000		0		0				1

		66		50		9.10E-05		40000		0		0				1

		67		50		0.000431444		40000		0		0				1

		68		50		681.919		40000		0		0				1

		69		50		1332.94		40000		0		0				1

		70		14		620.934		8963.72		0		0				0

		71		12		699.707		7776.37		0		0				0

		72		11		755.65		6933.12		0		0				0

		73		11		796.662		6314.96		0		0				0

		74		10		827.535		5849.59		0		0				0

		75		10		851.291		5491.52		0		0				0

		76		9		869.912		5210.84		0		0				0

		77		9		884.745		4987.26		0		0				0

		78		9		-884.745		4987.26		0		0				0

		79		9		-869.912		5210.84		0		0				0

		80		10		-851.291		5491.52		0		0				0

		81		10		-827.535		5849.59		0		0				0

		82		11		-796.662		6314.96		0		0				0

		83		11		-755.65		6933.12		0		0				0

		84		12		-699.707		7776.37		0		0				0

		85		14		-620.934		8963.72		0		0				0

		86		50		-1332.94		40000		0		0				1

		87		50		-681.919		40000		0		0				1

		88		50		0.000431444		40000		0		0				1

		89		49		-0.000108691		40000		0		0				1

		90		49		755.486		40000		0		0				1

		91		49		1465.21		40000		0		0				1

		92		13		651.266		8506.53		0		0				0

		93		12		729.465		7327.82		0		0				0

		94		11		784.371		6500.21		0		0				0

		95		10		824.35		5897.61		0		0				0

		96		9		854.307		5446.05		0		0				0

		97		9		877.285		5099.71		0		0				0

		98		9		895.256		4828.82		0		0				0

		99		8		909.552		4613.35		0		0				0

		100		8		-909.552		4613.35		0		0				0

		101		9		-895.256		4828.82		0		0				0

		102		9		-877.285		5099.71		0		0				0

		103		9		-854.307		5446.05		0		0				0

		104		10		-824.35		5897.6		0		0				0

		105		11		-784.371		6500.21		0		0				0

		106		12		-729.465		7327.82		0		0				0

		107		13		-651.266		8506.53		0		0				0

		108		49		-1465.21		40000		0		0				1

		109		49		-755.487		40000		0		0				1

		110		49		-0.000108691		40000		0		0				1

		111		49		-0.000316129		40000		0		0				1

		112		49		884.24		40000		0		0				1

		113		49		1575.37		40000		0		0				1

		114		12		695.697		7836.81		0		0				0

		115		11		772.276		6682.52		0		0				0

		116		10		825.261		5883.87		0		0				0

		117		9		863.483		5307.75		0		0				0

		118		9		891.946		4878.72		0		0				0

		119		8		913.688		4551		0		0				0

		120		8		930.646		4295.38		0		0				0

		121		8		944.114		4092.38		0		0				0

		122		8		-944.114		4092.38		0		0				0

		123		8		-930.646		4295.38		0		0				0

		124		8		-913.688		4551		0		0				0

		125		9		-891.946		4878.72		0		0				0

		126		9		-863.483		5307.75		0		0				0

		127		10		-825.261		5883.87		0		0				0

		128		11		-772.276		6682.52		0		0				0

		129		12		-695.697		7836.81		0		0				0

		130		49		-1575.37		40000		0		0				1

		131		49		-884.24		40000		0		0				1

		132		49		-0.000316129		40000		0		0				1

		133		49		0.000192897		40000		0		0				1

		134		49		1002.16		40000		0		0				1

		135		14		611.136		9111.41		0		0				0

		136		12		727.428		7358.53		0		0				0

		137		10		802.359		6229.08		0		0				0

		138		9		853.7		5455.21		0		0				0

		139		9		890.505		4900.43		0		0				0

		140		8		917.805		4488.94		0		0				0

		141		8		938.604		4175.43		0		0				0

		142		7		954.806		3931.22		0		0				0

		143		7		967.663		3737.42		0		0				0

		144		7		-967.663		3737.42		0		0				0

		145		7		-954.806		3931.22		0		0				0

		146		8		-938.604		4175.43		0		0				0

		147		8		-917.805		4488.94		0		0				0

		148		9		-890.505		4900.43		0		0				0

		149		9		-853.7		5455.21		0		0				0

		150		10		-802.359		6229.08		0		0				0

		151		12		-727.428		7358.53		0		0				0

		152		14		-611.136		9111.41		0		0				0

		153		49		-1002.16		40000		0		0				1

		154		49		0.000192897		40000		0		0				1

		155		49		-6.61E-05		40000		0		0				1

		156		49		1115.2		40000		0		0				1

		157		13		636.159		8734.23		0		0				0

		158		11		751.663		6993.23		0		0				0

		159		10		825.08		5886.6		0		0				0

		160		9		875.022		5133.82		0		0				0

		161		8		910.665		4596.56		0		0				0

		162		8		937.028		4199.19		0		0				0

		163		7		957.083		3896.89		0		0				0

		164		7		972.691		3661.64		0		0				0

		165		7		985.068		3475.07		0		0				0

		166		7		-985.068		3475.07		0		0				0

		167		7		-972.691		3661.64		0		0				0

		168		7		-957.083		3896.89		0		0				0

		169		8		-937.028		4199.19		0		0				0

		170		8		-910.665		4596.56		0		0				0

		171		9		-875.022		5133.82		0		0				0

		172		10		-825.08		5886.6		0		0				0

		173		11		-751.663		6993.22		0		0				0

		174		13		-636.159		8734.23		0		0				0

		175		49		-1115.2		40000		0		0				1

		176		49		-6.61E-05		40000		0		0				1

		177		48		-0.000197004		40000		0		0				1

		178		48		1211.52		40000		0		0				1

		179		13		656.487		8427.83		0		0				0

		180		11		771.003		6701.7		0		0				0

		181		9		843.058		5615.62		0		0				0

		182		8		891.799		4880.94		0		0				0

		183		8		926.466		4358.38		0		0				0

		184		7		952.058		3972.63		0		0				0

		185		7		971.503		3679.54		0		0				0

		186		7		986.627		3451.57		0		0				0

		187		7		998.623		3270.75		0		0				0

		188		7		-998.623		3270.75		0		0				0

		189		7		-986.627		3451.57		0		0				0

		190		7		-971.502		3679.55		0		0				0

		191		7		-952.058		3972.63		0		0				0

		192		8		-926.467		4358.38		0		0				0

		193		8		-891.799		4880.94		0		0				0

		194		9		-843.058		5615.62		0		0				0

		195		11		-771.004		6701.7		0		0				0

		196		13		-656.487		8427.83		0		0				0

		197		48		-1211.52		40000		0		0				1

		198		48		-0.000197004		40000		0		0				1

		199		48		-0.000250266		40000		0		0				1

		200		48		1269.81		40000		0		0				1

		201		12		673.47		8171.84		0		0				0

		202		10		786.923		6461.75		0		0				0

		203		9		857.753		5394.11		0		0				0

		204		8		905.452		4675.13		0		0				0

		205		7		939.289		4165.11		0		0				0

		206		7		964.23		3789.16		0		0				0

		207		7		983.164		3503.77		0		0				0

		208		6		997.89		3281.81		0		0				0

		209		6		1000		3075.22		0		0				0

		210		6		-1000		3075.22		0		0				0

		211		6		-997.89		3281.81		0		0				0

		212		7		-983.164		3503.77		0		0				0

		213		7		-964.23		3789.16		0		0				0

		214		7		-939.289		4165.11		0		0				0

		215		8		-905.452		4675.13		0		0				0

		216		9		-857.753		5394.11		0		0				0

		217		10		-786.923		6461.75		0		0				0

		218		12		-673.47		8171.84		0		0				0

		219		48		-1269.81		40000		0		0				1

		220		48		-0.000250266		40000		0		0				1

																				%%

																62				28.2

																158				71.8





180 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		------------------		-------------------		---------------		-----------------		----------------

		1		49		-7.91E-05		40000		0		0				1

		2		49		351.618		40000		0		0				1

		3		50		690.279		40000		0		0				1

		4		50		1003.18		40000		0		0				1

		5		50		1277.12		40000		0		0				1

		6		50		1488.69		40000		0		0				1

		7		14		596.016		9339.31		0		0				0

		8		14		628.376		8851.55		0		0				0

		9		13		655.003		8450.2		0		0				0

		10		13		676.334		8128.67		0		0				0

		11		12		693.607		7868.31		0		0				0

		12		12		-693.607		7868.31		0		0				0

		13		13		-676.335		8128.66		0		0				0

		14		13		-655.003		8450.2		0		0				0

		15		14		-628.376		8851.55		0		0				0

		16		14		-596.016		9339.31		0		0				0

		17		50		-1488.7		40000		0		0				1

		18		50		-1277.12		40000		0		0				1

		19		50		-1003.18		40000		0		0				1

		20		50		-690.28		40000		0		0				1

		21		49		-351.618		40000		0		0				1

		22		49		-7.91E-05		40000		0		0				1

		23		49		-0.000354273		40000		0		0				1

		24		49		492.82		40000		0		0				1

		25		49		964.81		40000		0		0				1

		26		49		1368.48		40000		0		0				1

		27		14		596.3		9335.04		0		0				0

		28		13		652.048		8494.75		0		0				0

		29		12		695.294		7842.89		0		0				0

		30		12		728.48		7342.67		0		0				0

		31		11		754.349		6952.74		0		0				0

		32		11		774.817		6644.22		0		0				0

		33		11		791.231		6396.81		0		0				0

		34		11		-791.231		6396.81		0		0				0

		35		11		-774.817		6644.22		0		0				0

		36		11		-754.349		6952.74		0		0				0

		37		12		-728.48		7342.67		0		0				0

		38		12		-695.294		7842.89		0		0				0

		39		13		-652.048		8494.75		0		0				0

		40		14		-596.3		9335.04		0		0				0

		41		49		-1368.48		40000		0		0				1

		42		49		-964.811		40000		0		0				1

		43		49		-492.82		40000		0		0				1

		44		49		-0.000354273		40000		0		0				1

		45		49		4.46E-05		40000		0		0				1

		46		49		600.766		40000		0		0				1

		47		49		1179.07		40000		0		0				1

		48		49		1607.96		40000		0		0				1

		49		13		652.179		8492.77		0		0				0

		50		12		709.209		7633.14		0		0				0

		51		11		751.527		6995.27		0		0				0

		52		11		783.637		6511.28		0		0				0

		53		10		808.481		6136.79		0		0				0

		54		10		828.033		5842.09		0		0				0

		55		10		843.65		5606.69		0		0				0

		56		10		-843.65		5606.69		0		0				0

		57		10		-828.033		5842.09		0		0				0

		58		10		-808.481		6136.79		0		0				0

		59		11		-783.637		6511.28		0		0				0

		60		11		-751.527		6995.27		0		0				0

		61		12		-709.209		7633.14		0		0				0

		62		13		-652.179		8492.77		0		0				0

		63		49		-1607.96		40000		0		0				1

		64		49		-1179.07		40000		0		0				1

		65		49		-600.766		40000		0		0				1

		66		49		4.46E-05		40000		0		0				1

		67		49		-0.000129221		40000		0		0				1

		68		49		691.742		40000		0		0				1

		69		49		1380.23		40000		0		0				1

		70		14		613.28		9079.09		0		0				0

		71		12		692		7892.53		0		0				0

		72		11		748.172		7045.85		0		0				0

		73		10		789.425		6424.03		0		0				0

		74		10		820.52		5955.34		0		0				0

		75		9		844.467		5594.38		0		0				0

		76		9		863.249		5311.27		0		0				0

		77		9		878.217		5085.65		0		0				0

		78		9		-878.217		5085.65		0		0				0

		79		9		-863.249		5311.27		0		0				0

		80		9		-844.467		5594.38		0		0				0

		81		10		-820.52		5955.34		0		0				0

		82		10		-789.425		6424.03		0		0				0

		83		11		-748.172		7045.85		0		0				0

		84		12		-692		7892.53		0		0				0

		85		14		-613.28		9079.09		0		0				0

		86		49		-1380.23		40000		0		0				1

		87		49		-691.742		40000		0		0				1

		88		49		-0.000129221		40000		0		0				1

		89		49		1.48E-06		40000		0		0				1

		90		49		771.206		40000		0		0				1

		91		49		1602.86		40000		0		0				1

		92		13		643.632		8621.59		0		0				0

		93		12		721.993		7440.45		0		0				0

		94		11		777.184		6608.54		0		0				0

		95		10		817.439		6001.77		0		0				0

		96		9		847.638		5546.57		0		0				0

		97		9		870.82		5197.16		0		0				0

		98		9		888.961		4923.72		0		0				0

		99		8		903.396		4706.13		0		0				0

		100		8		-903.396		4706.13		0		0				0

		101		9		-888.961		4923.72		0		0				0

		102		9		-870.82		5197.16		0		0				0

		103		9		-847.638		5546.57		0		0				0

		104		10		-817.439		6001.77		0		0				0

		105		11		-777.184		6608.54		0		0				0

		106		12		-721.993		7440.45		0		0				0

		107		13		-643.632		8621.59		0		0				0

		108		49		-1602.86		40000		0		0				1

		109		49		-771.206		40000		0		0				1

		110		49		1.48E-06		40000		0		0				1

		111		49		-0.00038803		40000		0		0				1

		112		49		900.66		40000		0		0				1

		113		49		2164.86		40000		0		0				1

		114		12		688.333		7947.81		0		0				0

		115		11		765.238		6788.6		0		0				0

		116		10		818.575		5984.64		0		0				0

		117		9		857.11		5403.8		0		0				0

		118		9		885.835		4970.82		0		0				0

		119		8		907.792		4639.87		0		0				0

		120		8		924.926		4381.61		0		0				0

		121		8		938.535		4176.47		0		0				0

		122		8		-938.535		4176.47		0		0				0

		123		8		-924.926		4381.61		0		0				0

		124		8		-907.792		4639.87		0		0				0

		125		9		-885.835		4970.82		0		0				0

		126		9		-857.11		5403.8		0		0				0

		127		10		-818.575		5984.64		0		0				0

		128		11		-765.238		6788.6		0		0				0

		129		12		-688.333		7947.81		0		0				0

		130		49		-2164.86		40000		0		0				1

		131		49		-900.661		40000		0		0				1

		132		49		-0.00038803		40000		0		0				1

		133		48		-6.25E-05		40000		0		0				1

		134		48		992.487		40000		0		0				1

		135		14		604.486		9211.65		0		0				0

		136		11		720.363		7465.01		0		0				0

		137		10		795.697		6329.5		0		0				0

		138		9		847.425		5549.79		0		0				0

		139		9		884.56		4990.04		0		0				0

		140		8		912.129		4574.49		0		0				0

		141		8		933.145		4257.71		0		0				0

		142		7		949.521		4010.87		0		0				0

		143		7		962.518		3814.97		0		0				0

		144		7		-962.518		3814.97		0		0				0

		145		7		-949.521		4010.87		0		0				0

		146		8		-933.145		4257.71		0		0				0

		147		8		-912.129		4574.49		0		0				0

		148		9		-884.561		4990.04		0		0				0

		149		9		-847.425		5549.79		0		0				0

		150		10		-795.697		6329.5		0		0				0

		151		11		-720.363		7465.01		0		0				0

		152		14		-604.486		9211.65		0		0				0

		153		48		-992.487		40000		0		0				1

		154		48		-6.25E-05		40000		0		0				1

		155		48		-6.14E-05		40000		0		0				1

		156		48		1057.44		40000		0		0				1

		157		13		629.213		8838.94		0		0				0

		158		11		744.877		7095.51		0		0				0

		159		10		818.743		5982.12		0		0				0

		160		9		869.092		5223.2		0		0				0

		161		8		905.072		4680.87		0		0				0

		162		8		931.704		4279.43		0		0				0

		163		7		951.975		3973.9		0		0				0

		164		7		967.75		3736.11		0		0				0

		165		7		980.268		3547.43		0		0				0

		166		7		-980.268		3547.43		0		0				0

		167		7		-967.75		3736.11		0		0				0

		168		7		-951.975		3973.9		0		0				0

		169		8		-931.704		4279.43		0		0				0

		170		8		-905.072		4680.87		0		0				0

		171		9		-869.092		5223.2		0		0				0

		172		10		-818.743		5982.12		0		0				0

		173		11		-744.877		7095.51		0		0				0

		174		13		-629.213		8838.94		0		0				0

		175		48		-1057.44		40000		0		0				1

		176		48		-6.14E-05		40000		0		0				1

		177		48		0.000141725		40000		0		0				1

		178		48		1115.84		40000		0		0				1

		179		13		649.58		8531.95		0		0				0

		180		11		764.473		6800.14		0		0				0

		181		9		837.005		5706.85		0		0				0

		182		8		886.164		4965.87		0		0				0

		183		8		921.169		4438.23		0		0				0

		184		7		947.029		4048.44		0		0				0

		185		7		966.683		3752.19		0		0				0

		186		7		981.976		3521.68		0		0				0

		187		6		994.106		3338.84		0		0				0

		188		6		-994.106		3338.84		0		0				0

		189		7		-981.976		3521.68		0		0				0

		190		7		-966.683		3752.19		0		0				0

		191		7		-947.029		4048.44		0		0				0

		192		8		-921.169		4438.23		0		0				0

		193		8		-886.164		4965.87		0		0				0

		194		9		-837.005		5706.85		0		0				0

		195		11		-764.473		6800.14		0		0				0

		196		13		-649.58		8531.95		0		0				0

		197		48		-1115.84		40000		0		0				1

		198		48		0.000141725		40000		0		0				1

		199		48		-0.000254275		40000		0		0				1

		200		48		1180.3		40000		0		0				1

		201		12		666.734		8273.37		0		0				0

		202		10		780.624		6556.69		0		0				0

		203		9		851.953		5481.54		0		0				0

		204		8		900.074		4756.21		0		0				0

		205		7		934.247		4241.11		0		0				0

		206		7		959.451		3861.2		0		0				0

		207		7		978.592		3572.69		0		0				0

		208		6		993.482		3348.25		0		0				0

		209		6		1000		3152.91		0		0				0

		210		6		-1000		3152.91		0		0				0

		211		6		-993.482		3348.25		0		0				0

		212		7		-978.592		3572.69		0		0				0

		213		7		-959.451		3861.21		0		0				0

		214		7		-934.247		4241.11		0		0				0

		215		8		-900.074		4756.21		0		0				0

		216		9		-851.953		5481.54		0		0				0

		217		10		-780.624		6556.69		0		0				0

		218		12		-666.734		8273.37		0		0				0

		219		48		-1180.3		40000		0		0				1

		220		48		-0.000254275		40000		0		0				1

																				%%

																62				28.2

																158				71.8





190 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		49		-0.000129697		40000		0		0				1

		2		49		354.912		40000		0		0				1

		3		49		704.112		40000		0		0				1

		4		49		1047.09		40000		0		0				1

		5		49		1398.23		40000		0		0				1

		6		49		1814.88		40000		0		0				1

		7		14		592.395		9393.89		0		0				0

		8		14		620.548		8969.55		0		0				0

		9		13		647.155		8568.49		0		0				0

		10		13		668.472		8247.18		0		0				0

		11		12		685.775		7986.37		0		0				0

		12		12		-685.775		7986.37		0		0				0

		13		13		-668.472		8247.18		0		0				0

		14		13		-647.155		8568.49		0		0				0

		15		14		-620.547		8969.57		0		0				0

		16		14		-592.395		9393.89		0		0				0

		17		49		-1814.88		40000		0		0				1

		18		49		-1398.23		40000		0		0				1

		19		49		-1047.09		40000		0		0				1

		20		49		-704.113		40000		0		0				1

		21		49		-354.912		40000		0		0				1

		22		49		-0.000129697		40000		0		0				1

		23		49		-0.000406166		40000		0		0				1

		24		49		493.437		40000		0		0				1

		25		49		985.856		40000		0		0				1

		26		49		1512.96		40000		0		0				1

		27		14		591.833		9402.37		0		0				0

		28		13		644.426		8609.62		0		0				0

		29		12		687.682		7957.62		0		0				0

		30		12		720.971		7455.85		0		0				0

		31		11		746.951		7064.25		0		0				0

		32		11		767.523		6754.17		0		0				0

		33		10		784.03		6505.36		0		0				0

		34		10		-784.03		6505.36		0		0				0

		35		11		-767.523		6754.17		0		0				0

		36		11		-746.951		7064.25		0		0				0

		37		12		-720.971		7455.85		0		0				0

		38		12		-687.682		7957.62		0		0				0

		39		13		-644.38		8610.33		0		0				0

		40		14		-591.833		9402.37		0		0				0

		41		49		-1512.97		40000		0		0				1

		42		49		-985.856		40000		0		0				1

		43		49		-493.437		40000		0		0				1

		44		49		-0.000406166		40000		0		0				1

		45		49		-0.000125971		40000		0		0				1

		46		49		593.236		40000		0		0				1

		47		49		1180.96		40000		0		0				1

		48		49		1691.46		40000		0		0				1

		49		13		644.614		8606.79		0		0				0

		50		12		701.788		7744.99		0		0				0

		51		11		744.277		7104.56		0		0				0

		52		11		776.557		6618		0		0				0

		53		10		801.556		6241.17		0		0				0

		54		10		821.243		5944.43		0		0				0

		55		9		836.976		5707.28		0		0				0

		56		9		-836.976		5707.28		0		0				0

		57		10		-821.243		5944.43		0		0				0

		58		10		-801.556		6241.17		0		0				0

		59		11		-776.557		6618		0		0				0

		60		11		-744.277		7104.56		0		0				0

		61		12		-701.788		7744.99		0		0				0

		62		13		-644.614		8606.79		0		0				0

		63		49		-1691.46		40000		0		0				1

		64		49		-1180.96		40000		0		0				1

		65		49		-593.236		40000		0		0				1

		66		49		-0.000125971		40000		0		0				1

		67		49		-0.000234981		40000		0		0				1

		68		49		671.735		40000		0		0				1

		69		49		1307.44		40000		0		0				1

		70		14		606.209		9185.68		0		0				0

		71		12		684.667		8003.06		0		0				0

		72		11		741.038		7153.39		0		0				0

		73		10		782.513		6528.23		0		0				0

		74		10		813.812		6056.44		0		0				0

		75		9		837.937		5692.8		0		0				0

		76		9		856.871		5407.41		0		0				0

		77		9		871.965		5179.89		0		0				0

		78		9		-871.965		5179.89		0		0				0

		79		9		-856.871		5407.41		0		0				0

		80		9		-837.937		5692.8		0		0				0

		81		10		-813.812		6056.44		0		0				0

		82		10		-782.513		6528.22		0		0				0

		83		11		-741.038		7153.39		0		0				0

		84		12		-684.667		8003.06		0		0				0

		85		14		-606.209		9185.68		0		0				0

		86		49		-1307.44		40000		0		0				1

		87		49		-671.735		40000		0		0				1

		88		49		-0.000234981		40000		0		0				1

		89		48		6.07E-05		40000		0		0				1

		90		48		736.693		40000		0		0				1

		91		49		1385.2		40000		0		0				1

		92		13		636.29		8732.26		0		0				0

		93		12		714.864		7547.9		0		0				0

		94		11		770.315		6712.08		0		0				0

		95		10		810.825		6101.47		0		0				0

		96		9		841.25		5642.87		0		0				0

		97		9		864.622		5290.57		0		0				0

		98		9		882.923		5014.73		0		0				0

		99		8		897.49		4795.15		0		0				0

		100		8		-897.49		4795.15		0		0				0

		101		9		-882.923		5014.73		0		0				0

		102		9		-864.622		5290.57		0		0				0

		103		9		-841.25		5642.87		0		0				0

		104		10		-810.825		6101.47		0		0				0

		105		11		-770.315		6712.08		0		0				0

		106		12		-714.864		7547.9		0		0				0

		107		13		-636.29		8732.26		0		0				0

		108		49		-1385.21		40000		0		0				1

		109		48		-736.693		40000		0		0				1

		110		48		6.07E-05		40000		0		0				1

		111		48		-0.000379413		40000		0		0				1

		112		48		844.006		40000		0		0				1

		113		48		1524.38		40000		0		0				1

		114		12		681.306		8053.72		0		0				0

		115		11		758.502		6890.14		0		0				0

		116		10		812.166		6081.26		0		0				0

		117		9		850.993		5496.01		0		0				0

		118		8		879.965		5059.31		0		0				0

		119		8		902.124		4725.3		0		0				0

		120		8		919.423		4464.54		0		0				0

		121		8		933.169		4257.36		0		0				0

		122		8		-933.169		4257.36		0		0				0

		123		8		-919.423		4464.54		0		0				0

		124		8		-902.124		4725.3		0		0				0

		125		8		-879.965		5059.31		0		0				0

		126		9		-850.993		5496.01		0		0				0

		127		10		-812.166		6081.26		0		0				0

		128		11		-758.502		6890.14		0		0				0

		129		12		-681.306		8053.72		0		0				0

		130		48		-1524.38		40000		0		0				1

		131		48		-844.006		40000		0		0				1

		132		48		-0.000379413		40000		0		0				1

		133		48		-2.22E-05		40000		0		0				1

		134		48		937.611		40000		0		0				1

		135		13		597.588		9315.62		0		0				0

		136		11		713.603		7566.92		0		0				0

		137		10		789.305		6425.84		0		0				0

		138		9		841.396		5640.67		0		0				0

		139		8		878.841		5076.25		0		0				0

		140		8		906.664		4656.87		0		0				0

		141		8		927.886		4336.98		0		0				0

		142		7		944.428		4087.64		0		0				0

		143		7		957.559		3889.72		0		0				0

		144		7		-957.559		3889.72		0		0				0

		145		7		-944.428		4087.64		0		0				0

		146		8		-927.886		4336.98		0		0				0

		147		8		-906.664		4656.87		0		0				0

		148		8		-878.841		5076.25		0		0				0

		149		9		-841.396		5640.67		0		0				0

		150		10		-789.305		6425.84		0		0				0

		151		11		-713.603		7566.92		0		0				0

		152		13		-597.588		9315.62		0		0				0

		153		48		-937.611		40000		0		0				1

		154		48		-2.22E-05		40000		0		0				1

		155		48		-0.000151741		40000		0		0				1

		156		48		1022.85		40000		0		0				1

		157		13		622.331		8942.67		0		0				0

		158		11		738.371		7193.59		0		0				0

		159		10		812.652		6073.92		0		0				0

		160		9		863.384		5309.24		0		0				0

		161		8		899.681		4762.13		0		0				0

		162		8		926.57		4356.83		0		0				0

		163		7		947.044		4048.21		0		0				0

		164		7		962.984		3807.95		0		0				0

		165		7		975.633		3617.29		0		0				0

		166		7		-975.633		3617.29		0		0				0

		167		7		-962.984		3807.95		0		0				0

		168		7		-947.044		4048.21		0		0				0

		169		8		-926.57		4356.83		0		0				0

		170		8		-899.681		4762.13		0		0				0

		171		9		-863.384		5309.24		0		0				0

		172		10		-812.652		6073.92		0		0				0

		173		11		-738.371		7193.59		0		0				0

		174		13		-622.331		8942.67		0		0				0

		175		48		-1022.85		40000		0		0				1

		176		48		-0.000151741		40000		0		0				1

		177		47		0.000294545		40000		0		0				1

		178		47		1099.01		40000		0		0				1

		179		12		642.978		8631.46		0		0				0

		180		10		758.2		6894.7		0		0				0

		181		9		831.179		5794.68		0		0				0

		182		8		880.73		5047.77		0		0				0

		183		8		916.056		4515.3		0		0				0

		184		7		942.17		4121.68		0		0				0

		185		7		962.027		3822.38		0		0				0

		186		7		977.48		3589.45		0		0				0

		187		6		989.739		3404.66		0		0				0

		188		6		-989.739		3404.66		0		0				0

		189		7		-977.48		3589.45		0		0				0

		190		7		-962.027		3822.38		0		0				0

		191		7		-942.17		4121.68		0		0				0

		192		8		-916.056		4515.3		0		0				0

		193		8		-880.73		5047.77		0		0				0

		194		9		-831.179		5794.67		0		0				0

		195		10		-758.2		6894.7		0		0				0

		196		12		-642.978		8631.46		0		0				0

		197		47		-1099.01		40000		0		0				1

		198		47		0.000294545		40000		0		0				1

		199		47		-0.000238779		40000		0		0				1

		200		47		1171.76		40000		0		0				1

		201		12		660.29		8370.51		0		0				0

		202		10		774.565		6648.01		0		0				0

		203		9		846.36		5565.83		0		0				0

		204		8		894.88		4834.49		0		0				0

		205		7		929.373		4314.57		0		0				0

		206		7		954.827		3930.9		0		0				0

		207		7		974.168		3639.38		0		0				0

		208		6		989.215		3412.57		0		0				0

		209		6		1000		3228.77		0		0				0

		210		6		-1000		3228.77		0		0				0

		211		6		-989.215		3412.57		0		0				0

		212		7		-974.168		3639.38		0		0				0

		213		7		-954.827		3930.9		0		0				0

		214		7		-929.373		4314.57		0		0				0

		215		8		-894.88		4834.49		0		0				0

		216		9		-846.361		5565.83		0		0				0

		217		10		-774.565		6648.01		0		0				0

		218		12		-660.29		8370.51		0		0				0

		219		47		-1171.76		40000		0		0				1

		220		47		-0.000238779		40000		0		0				1

																				%%

																62				28.2

																158				71.8





200 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		49		-0.000215533		40000		0		0				1

		2		49		338.43		40000		0		0				1

		3		49		664.512		40000		0		0				1

		4		49		964.676		40000		0		0				1

		5		49		1219.98		40000		0		0				1

		6		49		1401.61		40000		0		0				1

		7		49		1487		40000		0		0				1

		8		13		613.353		9077.99		0		0				0

		9		13		639.615		8682.15		0		0				0

		10		12		660.993		8359.91		0		0				0

		11		12		678.325		8098.66		0		0				0

		12		12		-678.325		8098.66		0		0				0

		13		12		-660.993		8359.91		0		0				0

		14		13		-639.615		8682.15		0		0				0

		15		13		-613.353		9077.99		0		0				0

		16		49		-1487		40000		0		0				1

		17		49		-1401.61		40000		0		0				1

		18		49		-1219.98		40000		0		0				1

		19		49		-964.676		40000		0		0				1

		20		49		-664.513		40000		0		0				1

		21		49		-338.43		40000		0		0				1

		22		49		-0.000215533		40000		0		0				1

		23		48		-0.000493408		40000		0		0				1

		24		48		467.305		40000		0		0				1

		25		48		910.664		40000		0		0				1

		26		49		1295.89		40000		0		0				1

		27		48		1533.47		40000		0		0				1

		28		13		637.056		8720.72		0		0				0

		29		12		680.431		8066.92		0		0				0

		30		11		713.806		7563.85		0		0				0

		31		11		739.885		7170.76		0		0				0

		32		11		760.549		6859.28		0		0				0

		33		10		777.14		6609.2		0		0				0

		34		10		-777.141		6609.2		0		0				0

		35		11		-760.549		6859.28		0		0				0

		36		11		-739.884		7170.77		0		0				0

		37		11		-713.806		7563.85		0		0				0

		38		12		-680.431		8066.92		0		0				0

		39		13		-637.056		8720.72		0		0				0

		40		48		-1533.47		40000		0		0				1

		41		49		-1295.89		40000		0		0				1

		42		48		-910.664		40000		0		0				1

		43		48		-467.305		40000		0		0				1

		44		48		-0.000493408		40000		0		0				1

		45		48		-0.000298699		40000		0		0				1

		46		48		562.819		40000		0		0				1

		47		48		1092.96		40000		0		0				1

		48		48		1576.96		40000		0		0				1

		49		13		637.483		8714.29		0		0				0

		50		12		694.708		7851.72		0		0				0

		51		11		737.345		7209.04		0		0				0

		52		10		769.779		6720.16		0		0				0

		53		10		794.921		6341.19		0		0				0

		54		10		814.733		6042.56		0		0				0

		55		9		830.574		5803.79		0		0				0

		56		9		-830.574		5803.79		0		0				0

		57		10		-814.733		6042.56		0		0				0

		58		10		-794.921		6341.19		0		0				0

		59		10		-769.779		6720.16		0		0				0

		60		11		-737.345		7209.04		0		0				0

		61		12		-694.708		7851.71		0		0				0

		62		13		-637.483		8714.29		0		0				0

		63		48		-1576.96		40000		0		0				1

		64		48		-1092.96		40000		0		0				1

		65		48		-562.819		40000		0		0				1

		66		48		-0.000298699		40000		0		0				1

		67		48		-0.000127784		40000		0		0				1

		68		48		642.012		40000		0		0				1

		69		48		1251.13		40000		0		0				1

		70		13		599.563		9285.86		0		0				0

		71		12		677.669		8108.55		0		0				0

		72		11		734.212		7256.27		0		0				0

		73		10		775.888		6628.08		0		0				0

		74		10		807.377		6153.45		0		0				0

		75		9		831.668		5787.3		0		0				0

		76		9		850.742		5499.79		0		0				0

		77		9		865.956		5270.47		0		0				0

		78		9		-865.956		5270.47		0		0				0

		79		9		-850.742		5499.79		0		0				0

		80		9		-831.668		5787.3		0		0				0

		81		10		-807.377		6153.45		0		0				0

		82		10		-775.888		6628.08		0		0				0

		83		11		-734.211		7256.28		0		0				0

		84		12		-677.662		8108.66		0		0				0

		85		13		-599.563		9285.86		0		0				0

		86		48		-1251.13		40000		0		0				1

		87		48		-642.012		40000		0		0				1

		88		48		-0.000127784		40000		0		0				1

		89		48		-0.000161324		40000		0		0				1

		90		48		711.298		40000		0		0				1

		91		48		1391.75		40000		0		0				1

		92		13		629.4		8836.12		0		0				0

		93		11		708.045		7650.68		0		0				0

		94		10		763.73		6811.34		0		0				0

		95		10		804.475		6197.18		0		0				0

		96		9		835.11		5735.41		0		0				0

		97		9		858.662		5380.42		0		0				0

		98		8		877.113		5102.3		0		0				0

		99		8		891.805		4880.85		0		0				0

		100		8		-891.805		4880.84		0		0				0

		101		8		-877.113		5102.3		0		0				0

		102		9		-858.662		5380.42		0		0				0

		103		9		-835.11		5735.41		0		0				0

		104		10		-804.475		6197.18		0		0				0

		105		10		-763.73		6811.34		0		0				0

		106		11		-708.045		7650.68		0		0				0

		107		13		-629.4		8836.12		0		0				0

		108		48		-1391.75		40000		0		0				1

		109		48		-711.298		40000		0		0				1

		110		48		-0.000161324		40000		0		0				1

		111		48		-0.000423617		40000		0		0				1

		112		48		831.744		40000		0		0				1

		113		48		1521		40000		0		0				1

		114		12		674.572		8155.23		0		0				0

		115		11		752.04		6987.55		0		0				0

		116		10		806.004		6174.14		0		0				0

		117		9		845.104		5584.77		0		0				0

		118		8		874.308		5144.57		0		0				0

		119		8		896.66		4807.66		0		0				0

		120		8		914.117		4544.53		0		0				0

		121		8		927.991		4335.4		0		0				0

		122		8		-927.991		4335.4		0		0				0

		123		8		-914.117		4544.53		0		0				0

		124		8		-896.66		4807.66		0		0				0

		125		8		-874.308		5144.57		0		0				0

		126		9		-845.104		5584.77		0		0				0

		127		10		-806.004		6174.14		0		0				0

		128		11		-752.04		6987.55		0		0				0

		129		12		-674.572		8155.23		0		0				0

		130		48		-1521		40000		0		0				1

		131		48		-831.745		40000		0		0				1

		132		48		-0.000423617		40000		0		0				1

		133		47		-0.000169721		40000		0		0				1

		134		47		939.74		40000		0		0				1

		135		13		592.397		9393.87		0		0				0

		136		11		707.121		7664.61		0		0				0

		137		10		783.159		6518.48		0		0				0

		138		9		835.587		5728.22		0		0				0

		139		8		873.325		5159.39		0		0				0

		140		8		901.389		4736.39		0		0				0

		141		8		922.807		4413.54		0		0				0

		142		7		939.508		4161.81		0		0				0

		143		7		952.767		3961.96		0		0				0

		144		7		-952.767		3961.96		0		0				0

		145		7		-939.508		4161.81		0		0				0

		146		8		-922.807		4413.54		0		0				0

		147		8		-901.389		4736.39		0		0				0

		148		8		-873.325		5159.39		0		0				0

		149		9		-835.588		5728.22		0		0				0

		150		10		-783.159		6518.48		0		0				0

		151		11		-707.121		7664.61		0		0				0

		152		13		-592.397		9393.87		0		0				0

		153		47		-939.741		40000		0		0				1

		154		47		-0.000169721		40000		0		0				1

		155		47		-0.000182534		40000		0		0				1

		156		47		1046.06		40000		0		0				1

		157		13		615.976		9038.45		0		0				0

		158		11		732.121		7287.79		0		0				0

		159		9		806.785		6162.36		0		0				0

		160		9		857.875		5392.27		0		0				0

		161		8		894.472		4840.63		0		0				0

		162		8		921.606		4431.65		0		0				0

		163		7		942.276		4120.08		0		0				0

		164		7		958.373		3877.45		0		0				0

		165		7		971.148		3684.9		0		0				0

		166		7		-971.148		3684.9		0		0				0

		167		7		-958.373		3877.45		0		0				0

		168		7		-942.276		4120.08		0		0				0

		169		8		-921.606		4431.65		0		0				0

		170		8		-894.472		4840.63		0		0				0

		171		9		-857.875		5392.27		0		0				0

		172		9		-806.785		6162.36		0		0				0

		173		11		-732.121		7287.79		0		0				0

		174		13		-615.976		9038.45		0		0				0

		175		47		-1046.06		40000		0		0				1

		176		47		-0.000182534		40000		0		0				1

		177		47		-0.000197399		40000		0		0				1

		178		47		1154.38		40000		0		0				1

		179		12		636.686		8726.3		0		0				0

		180		10		752.164		6985.67		0		0				0

		181		9		825.558		5879.4		0		0				0

		182		8		875.481		5126.9		0		0				0

		183		8		911.111		4589.85		0		0				0

		184		7		937.468		4192.56		0		0				0

		185		7		957.518		3890.34		0		0				0

		186		7		973.124		3655.11		0		0				0

		187		6		985.508		3468.45		0		0				0

		188		6		-985.508		3468.45		0		0				0

		189		7		-973.124		3655.11		0		0				0

		190		7		-957.518		3890.34		0		0				0

		191		7		-937.468		4192.56		0		0				0

		192		8		-911.111		4589.85		0		0				0

		193		8		-875.481		5126.9		0		0				0

		194		9		-825.558		5879.4		0		0				0

		195		10		-752.164		6985.67		0		0				0

		196		12		-636.686		8726.3		0		0				0

		197		47		-1154.41		40000		0		0				1

		198		47		-0.000197399		40000		0		0				1

		199		47		-0.000355386		40000		0		0				1

		200		47		1255.48		40000		0		0				1

		201		12		654.076		8464.18		0		0				0

		202		10		768.728		6736		0		0				0

		203		9		840.959		5647.25		0		0				0

		204		8		889.856		4910.21		0		0				0

		205		7		924.654		4385.7		0		0				0

		206		7		950.349		3998.4		0		0				0

		207		7		969.879		3704.02		0		0				0

		208		6		985.076		3474.95		0		0				0

		209		6		997.134		3293.19		0		0				0

		210		6		-997.134		3293.19		0		0				0

		211		6		-985.076		3474.95		0		0				0

		212		7		-969.879		3704.02		0		0				0

		213		7		-950.349		3998.4		0		0				0

		214		7		-924.654		4385.7		0		0				0

		215		8		-889.856		4910.21		0		0				0

		216		9		-840.959		5647.25		0		0				0

		217		10		-768.728		6736		0		0				0

		218		12		-654.076		8464.18		0		0				0

		219		47		-1255.48		40000		0		0				1

		220		47		-0.000355386		40000		0		0				1

																				%%

																66				30.0

																154				70.0






_1148644530.xls
Chart1

		0		0

		10		10

		20		20

		30		30

		40		40

		50		50

		60		60

		70		70

		80		80

		90		90

		100		100

		110		110

		120		120

		130		130

		140		140

		150		150

		160		160

		170		170

		180		180

		190		190

		200		200



Anode signal

Cupola signal

Cupola bias, Volt

Signal distribution, %%

28.1818181818

71.8181818182

42.7272727273

57.2727272727

50.9090909091

49.0909090909

56.3636363636

43.6363636364

63.6363636364

36.3636363636

67.2727272727

32.7272727273

71.8181818182

28.1818181818

74.5454545455

25.4545454545

78.1818181818

21.8181818182

79.0909090909

20.9090909091

80

20

82.7272727273

17.2727272727

86.3636363636

13.6363636364

87.2727272727

12.7272727273

89.0909090909

10.9090909091

90.4545454545

9.5454545455

90.9090909091

9.0909090909

91.8181818182

8.1818181818

92.7272727273

7.2727272727

94.5454545455

5.4545454545

94.5454545455

5.4545454545



Summary

		

		Bias Potential		Anode		Cupola

				%		%

		0		28.2		71.8

		10		42.7		57.3

		20		50.9		49.1

		30		56.4		43.6

		40		63.6		36.4

		50		67.3		32.7

		60		71.8		28.2

		70		74.5		25.5

		80		78.2		21.8

		90		79.1		20.9

		100		80.0		20.0

		110		82.7		17.3

		120		86.4		13.6

		130		87.3		12.7

		140		89.1		10.9

		150		90.5		9.5

		160		90.9		9.1

		170		91.8		8.2

		180		92.7		7.3

		190		94.5		5.5

		200		94.5		5.5





Summary

		



Anode signal

Cupola signal

Cupola bias, Volt

Signal distribution, %%



0 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		-----------		---------------------		----------------

		1		82		0.000142919		40000		0		0				1

		2		82		-129.929		40000		0		0				1

		3		82		-367.596		40000		0		0				1

		4		83		165.81		40000		0		0				1

		5		18		347.324		937.745		0		0				0

		6		18		427.204		904.155		0		0				0

		7		18		497.159		867.659		0		0				0

		8		18		558.04		829.814		0		0				0

		9		18		610.729		791.839		0		0				0

		10		18		656.159		754.622		0		0				0

		11		18		695.259		718.76		0		0				0

		12		18		-695.259		718.76		0		0				0

		13		18		-656.159		754.622		0		0				0

		14		18		-610.729		791.839		0		0				0

		15		18		-558.04		829.814		0		0				0

		16		18		-497.159		867.659		0		0				0

		17		18		-427.204		904.155		0		0				0

		18		18		-347.324		937.745		0		0				0

		19		83		-165.81		40000		0		0				1

		20		82		367.596		40000		0		0				1

		21		82		129.929		40000		0		0				1

		22		82		0.000142919		40000		0		0				1

		23		76		1.58E-05		40000		0		0				1

		24		76		130.118		40000		0		0				1

		25		77		171.194		40000		0		0				1

		26		13		270.809		962.633		0		0				0

		27		13		358.326		933.597		0		0				0

		28		13		436.017		899.938		0		0				0

		29		13		504.514		863.404		0		0				0

		30		13		564.352		825.535		0		0				0

		31		13		616.258		787.544		0		0				0

		32		13		661.079		750.316		0		0				0

		33		13		699.691		714.446		0		0				0

		34		13		-699.691		714.446		0		0				0

		35		13		-661.079		750.316		0		0				0

		36		13		-616.258		787.544		0		0				0

		37		13		-564.352		825.535		0		0				0

		38		13		-504.514		863.404		0		0				0

		39		13		-436.017		899.938		0		0				0

		40		13		-358.326		933.597		0		0				0

		41		13		-270.809		962.633		0		0				0

		42		77		-171.194		40000		0		0				1

		43		76		-130.118		40000		0		0				1

		44		76		1.58E-05		40000		0		0				1

		45		73		3.30E-05		40000		0		0				1

		46		74		311.796		40000		0		0				1

		47		74		547.26		40000		0		0				1

		48		11		275.833		961.206		0		0				0

		49		11		361.993		932.181		0		0				0

		50		11		438.951		898.511		0		0				0

		51		11		506.96		861.969		0		0				0

		52		11		566.45		824.096		0		0				0

		53		11		618.096		786.103		0		0				0

		54		11		662.714		748.873		0		0				0

		55		11		701.163		713.001		0		0				0

		56		11		-701.163		713.001		0		0				0

		57		11		-662.714		748.873		0		0				0

		58		11		-618.096		786.103		0		0				0

		59		11		-566.45		824.096		0		0				0

		60		11		-506.96		861.969		0		0				0

		61		11		-438.951		898.511		0		0				0

		62		11		-361.993		932.181		0		0				0

		63		11		-275.833		961.206		0		0				0

		64		74		-547.26		40000		0		0				1

		65		74		-311.796		40000		0		0				1

		66		73		3.30E-05		40000		0		0				1

		67		71		-1.04E-05		40000		0		0				1

		68		72		451.434		40000		0		0				1

		69		72		831.508		40000		0		0				1

		70		10		278.366		960.475		0		0				0

		71		10		363.826		931.467		0		0				0

		72		10		440.418		897.793		0		0				0

		73		10		508.183		861.249		0		0				0

		74		10		567.499		823.374		0		0				0

		75		10		619.014		785.38		0		0				0

		76		10		663.53		748.149		0		0				0

		77		10		701.898		712.277		0		0				0

		78		10		-701.898		712.277		0		0				0

		79		10		-663.53		748.149		0		0				0

		80		10		-619.014		785.38		0		0				0

		81		10		-567.499		823.374		0		0				0

		82		10		-508.183		861.249		0		0				0

		83		10		-440.418		897.793		0		0				0

		84		10		-363.826		931.467		0		0				0

		85		10		-278.366		960.475		0		0				0

		86		72		-831.508		40000		0		0				1

		87		72		-451.434		40000		0		0				1

		88		71		-1.04E-05		40000		0		0				1

		89		70		-0.000104516		40000		0		0				1

		90		70		564.856		40000		0		0				1

		91		70		1058.45		40000		0		0				1

		92		9		279.911		960.026		0		0				0

		93		9		364.93		931.035		0		0				0

		94		9		441.298		897.36		0		0				0

		95		9		508.917		860.816		0		0				0

		96		9		568.128		822.94		0		0				0

		97		9		619.564		784.946		0		0				0

		98		9		664.02		747.715		0		0				0

		99		9		702.339		711.842		0		0				0

		100		9		-702.339		711.842		0		0				0

		101		9		-664.02		747.715		0		0				0

		102		9		-619.564		784.946		0		0				0

		103		9		-568.128		822.94		0		0				0

		104		9		-508.917		860.816		0		0				0

		105		9		-441.298		897.36		0		0				0

		106		9		-364.93		931.035		0		0				0

		107		9		-279.911		960.026		0		0				0

		108		70		-1058.45		40000		0		0				1

		109		70		-564.856		40000		0		0				1

		110		70		-0.000104516		40000		0		0				1

		111		68		-3.35E-05		40000		0		0				1

		112		68		744.354		40000		0		0				1

		113		68		1423.71		40000		0		0				1

		114		8		281.5		959.561		0		0				0

		115		8		366.174		930.546		0		0				0

		116		8		442.299		896.868		0		0				0

		117		8		509.753		860.321		0		0				0

		118		8		568.845		822.445		0		0				0

		119		8		620.193		784.449		0		0				0

		120		8		664.579		747.218		0		0				0

		121		8		702.843		711.345		0		0				0

		122		8		-702.843		711.345		0		0				0

		123		8		-664.579		747.218		0		0				0

		124		8		-620.193		784.449		0		0				0

		125		8		-568.845		822.445		0		0				0

		126		8		-509.753		860.321		0		0				0

		127		8		-442.299		896.868		0		0				0

		128		8		-366.174		930.546		0		0				0

		129		8		-281.5		959.561		0		0				0

		130		68		-1423.71		40000		0		0				1

		131		68		-744.354		40000		0		0				1

		132		68		-3.35E-05		40000		0		0				1

		133		66		-5.08E-05		40000		0		0				1

		134		66		883.64		40000		0		0				1

		135		66		1711.58		40000		0		0				1

		136		7		282.378		959.303		0		0				0

		137		7		366.878		930.269		0		0				0

		138		7		442.857		896.592		0		0				0

		139		7		510.217		860.045		0		0				0

		140		7		569.244		822.169		0		0				0

		141		7		620.542		784.173		0		0				0

		142		7		664.889		746.942		0		0				0

		143		7		703.122		711.069		0		0				0

		144		7		-703.122		711.069		0		0				0

		145		7		-664.889		746.942		0		0				0

		146		7		-620.542		784.173		0		0				0

		147		7		-569.244		822.169		0		0				0

		148		7		-510.217		860.045		0		0				0

		149		7		-442.857		896.592		0		0				0

		150		7		-366.878		930.269		0		0				0

		151		7		-282.378		959.303		0		0				0

		152		66		-1711.58		40000		0		0				1

		153		66		-883.64		40000		0		0				1

		154		66		-5.08E-05		40000		0		0				1

		155		65		6.75E-06		40000		0		0				1

		156		65		1002.4		40000		0		0				1

		157		65		1982.94		40000		0		0				1

		158		7		282.904		959.148		0		0				0

		159		7		367.191		930.146		0		0				0

		160		7		443.215		896.415		0		0				0

		161		7		510.515		859.869		0		0				0

		162		7		569.499		821.992		0		0				0

		163		7		620.765		783.997		0		0				0

		164		7		665.087		746.766		0		0				0

		165		7		703.3		710.893		0		0				0

		166		7		-703.3		710.893		0		0				0

		167		7		-665.087		746.766		0		0				0

		168		7		-620.765		783.997		0		0				0

		169		7		-569.499		821.992		0		0				0

		170		7		-510.515		859.869		0		0				0

		171		7		-443.215		896.415		0		0				0

		172		7		-367.191		930.146		0		0				0

		173		7		-282.904		959.148		0		0				0

		174		65		-1982.94		40000		0		0				1

		175		65		-1002.4		40000		0		0				1

		176		65		6.75E-06		40000		0		0				1

		177		63		6.05E-05		40000		0		0				1

		178		63		1097.63		40000		0		0				1

		179		64		2161.37		40000		0		0				1

		180		6		283.243		959.048		0		0				0

		181		6		367.621		929.976		0		0				0

		182		6		443.455		896.296		0		0				0

		183		6		510.717		859.749		0		0				0

		184		6		569.672		821.872		0		0				0

		185		6		620.917		783.876		0		0				0

		186		6		665.223		746.645		0		0				0

		187		6		703.423		710.772		0		0				0

		188		6		-703.423		710.772		0		0				0

		189		6		-665.223		746.645		0		0				0

		190		6		-620.917		783.876		0		0				0

		191		6		-569.672		821.872		0		0				0

		192		6		-510.717		859.749		0		0				0

		193		6		-443.455		896.296		0		0				0

		194		6		-367.621		929.976		0		0				0

		195		6		-283.243		959.048		0		0				0

		196		64		-2161.37		40000		0		0				1

		197		63		-1097.63		40000		0		0				1

		198		63		6.05E-05		40000		0		0				1

		199		62		4.06E-05		40000		0		0				1

		200		62		1183.38		40000		0		0				1

		201		62		2299.83		40000		0		0				1

		202		6		283.743		958.9		0		0				0

		203		6		367.845		929.887		0		0				0

		204		6		443.635		896.208		0		0				0

		205		6		510.867		859.66		0		0				0

		206		6		569.802		821.782		0		0				0

		207		6		621.03		783.787		0		0				0

		208		6		665.324		746.555		0		0				0

		209		6		703.513		710.682		0		0				0

		210		6		-703.513		710.682		0		0				0

		211		6		-665.324		746.555		0		0				0

		212		6		-621.03		783.787		0		0				0

		213		6		-569.802		821.782		0		0				0

		214		6		-510.867		859.66		0		0				0

		215		6		-443.635		896.208		0		0				0

		216		6		-367.845		929.887		0		0				0

		217		6		-283.743		958.9		0		0				0

		218		62		-2299.83		40000		0		0				1

		219		62		-1183.38		40000		0		0				1

		220		62		4.06E-05		40000		0		0				1

																				%%

										Electrons Reaching Anode						62				28.2

										Electrons falling on Cupola						158				71.8





10 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		------------------		---------------		-----------------		----------------

		1		70		0.000180446		40000		0		0				1

		2		70		422.056		40000		0		0				1

		3		70		832.811		40000		0		0				1

		4		70		1222.34		40000		0		0				1

		5		70		1582.53		40000		0		0				1

		6		70		1906.64		40000		0		0				1

		7		70		2189.2		40000		0		0				1

		8		9		267.076		963.676		0		0				0

		9		9		292.862		956.155		0		0				0

		10		9		315.291		948.995		0		0				0

		11		9		335.001		942.218		0		0				0

		12		9		-335.001		942.218		0		0				0

		13		9		-315.291		948.995		0		0				0

		14		9		-292.862		956.155		0		0				0

		15		9		-267.076		963.676		0		0				0

		16		70		-2189.2		40000		0		0				1

		17		70		-1906.64		40000		0		0				1

		18		70		-1582.53		40000		0		0				1

		19		70		-1222.34		40000		0		0				1

		20		70		-832.811		40000		0		0				1

		21		70		-422.056		40000		0		0				1

		22		70		0.000180446		40000		0		0				1

		23		69		0.000131577		40000		0		0				1

		24		69		571.313		40000		0		0				1

		25		69		1127.28		40000		0		0				1

		26		69		1653.09		40000		0		0				1

		27		69		2131.49		40000		0		0				1

		28		70		2671.59		40000		0		0				1

		29		8		300.109		953.905		0		0				0

		30		8		336.3		941.755		0		0				0

		31		8		368.338		929.692		0		0				0

		32		8		396.581		918		0		0				0

		33		8		421.432		906.86		0		0				0

		34		8		-421.432		906.86		0		0				0

		35		8		-396.581		918		0		0				0

		36		8		-368.338		929.692		0		0				0

		37		8		-336.3		941.755		0		0				0

		38		8		-300.109		953.905		0		0				0

		39		70		-2671.59		40000		0		0				1

		40		69		-2131.49		40000		0		0				1

		41		69		-1653.09		40000		0		0				1

		42		69		-1127.28		40000		0		0				1

		43		69		-571.313		40000		0		0				1

		44		69		0.000131577		40000		0		0				1

		45		68		-6.75E-05		40000		0		0				1

		46		68		679.414		40000		0		0				1

		47		68		1342.23		40000		0		0				1

		48		68		1974.49		40000		0		0				1

		49		68		2576.27		40000		0		0				1

		50		7		290.906		956.752		0		0				0

		51		8		336.309		941.752		0		0				0

		52		8		376.733		926.322		0		0				0

		53		8		412.527		910.945		0		0				0

		54		8		444.105		895.975		0		0				0

		55		8		471.842		881.683		0		0				0

		56		8		-471.842		881.683		0		0				0

		57		8		-444.105		895.975		0		0				0

		58		8		-412.547		910.936		0		0				0

		59		8		-376.733		926.322		0		0				0

		60		8		-336.309		941.752		0		0				0

		61		7		-290.906		956.752		0		0				0

		62		68		-2576.27		40000		0		0				1

		63		68		-1974.49		40000		0		0				1

		64		68		-1342.23		40000		0		0				1

		65		68		-679.414		40000		0		0				1

		66		68		-6.75E-05		40000		0		0				1

		67		67		-0.000117749		40000		0		0				1

		68		67		765.728		40000		0		0				1

		69		67		1511.8		40000		0		0				1

		70		67		2227.25		40000		0		0				1

		71		67		2997.51		40000		0		0				1

		72		7		312.464		949.93		0		0				0

		73		7		361.115		932.521		0		0				0

		74		7		404.424		914.572		0		0				0

		75		7		442.707		896.666		0		0				0

		76		7		476.409		879.224		0		0				0

		77		7		506.022		862.521		0		0				0

		78		7		-506.022		862.521		0		0				0

		79		7		-476.409		879.224		0		0				0

		80		7		-442.707		896.666		0		0				0

		81		7		-404.424		914.572		0		0				0

		82		7		-361.115		932.521		0		0				0

		83		7		-312.464		949.93		0		0				0

		84		67		-2997.51		40000		0		0				1

		85		67		-2227.25		40000		0		0				1

		86		67		-1511.8		40000		0		0				1

		87		67		-765.728		40000		0		0				1

		88		67		-0.000117749		40000		0		0				1

		89		66		-3.42E-06		40000		0		0				1

		90		66		839.498		40000		0		0				1

		91		66		1656.32		40000		0		0				1

		92		66		2422.37		40000		0		0				1

		93		7		271.237		962.512		0		0				0

		94		7		328.327		944.564		0		0				0

		95		7		379.335		925.259		0		0				0

		96		7		424.702		905.333		0		0				0

		97		7		464.767		885.433		0		0				0

		98		7		500.003		866.024		0		0				0

		99		7		530.935		847.413		0		0				0

		100		7		-530.935		847.413		0		0				0

		101		7		-500.003		866.024		0		0				0

		102		7		-464.767		885.433		0		0				0

		103		7		-424.702		905.333		0		0				0

		104		7		-379.335		925.259		0		0				0

		105		7		-328.327		944.564		0		0				0

		106		7		-271.237		962.512		0		0				0

		107		66		-2422.37		40000		0		0				1

		108		66		-1656.32		40000		0		0				1

		109		66		-839.498		40000		0		0				1

		110		66		-3.42E-06		40000		0		0				1

		111		65		6.87E-05		40000		0		0				1

		112		65		963.74		40000		0		0				1

		113		65		1906.33		40000		0		0				1

		114		65		2810.46		40000		0		0				1

		115		6		290.041		957.014		0		0				0

		116		6		350.493		936.565		0		0				0

		117		6		404.74		914.432		0		0				0

		118		6		452.91		891.556		0		0				0

		119		6		495.378		868.677		0		0				0

		120		6		532.66		846.329		0		0				0

		121		6		565.328		824.866		0		0				0

		122		6		-565.328		824.866		0		0				0

		123		6		-532.66		846.329		0		0				0

		124		6		-495.378		868.677		0		0				0

		125		6		-452.91		891.556		0		0				0

		126		6		-404.74		914.432		0		0				0

		127		6		-350.493		936.565		0		0				0

		128		6		-290.041		957.014		0		0				0

		129		65		-2810.46		40000		0		0				1

		130		65		-1906.33		40000		0		0				1

		131		65		-963.74		40000		0		0				1

		132		65		6.87E-05		40000		0		0				1

		133		64		0.000105767		40000		0		0				1

		134		64		1064.35		40000		0		0				1

		135		64		2098.45		40000		0		0				1

		136		64		2997.4		40000		0		0				1

		137		6		302.47		953.159		0		0				0

		138		6		365.35		930.87		0		0				0

		139		6		421.722		906.725		0		0				0

		140		6		471.714		881.751		0		0				0

		141		6		515.725		856.754		0		0				0

		142		6		554.333		832.295		0		0				0

		143		6		588.118		808.775		0		0				0

		144		6		-588.118		808.775		0		0				0

		145		6		-554.333		832.295		0		0				0

		146		6		-515.725		856.754		0		0				0

		147		6		-471.714		881.751		0		0				0

		148		6		-421.722		906.725		0		0				0

		149		6		-365.35		930.87		0		0				0

		150		6		-302.47		953.159		0		0				0

		151		64		-2997.4		40000		0		0				1

		152		64		-2098.45		40000		0		0				1

		153		64		-1064.35		40000		0		0				1

		154		64		0.000105767		40000		0		0				1

		155		63		4.02E-05		40000		0		0				1

		156		63		1153.41		40000		0		0				1

		157		63		2279.83		40000		0		0				1

		158		63		3385.75		40000		0		0				1

		159		6		311.281		950.318		0		0				0

		160		6		375.889		926.665		0		0				0

		161		6		433.976		900.925		0		0				0

		162		6		485.256		874.372		0		0				0

		163		6		530.347		847.78		0		0				0

		164		6		569.821		821.769		0		0				0

		165		6		604.306		796.752		0		0				0

		166		6		-604.306		796.752		0		0				0

		167		6		-569.821		821.769		0		0				0

		168		6		-530.347		847.78		0		0				0

		169		6		-485.256		874.372		0		0				0

		170		6		-433.976		900.925		0		0				0

		171		6		-375.889		926.665		0		0				0

		172		6		-311.281		950.318		0		0				0

		173		63		-3385.75		40000		0		0				1

		174		63		-2279.83		40000		0		0				1

		175		63		-1153.41		40000		0		0				1

		176		63		4.02E-05		40000		0		0				1

		177		62		0.000155958		40000		0		0				1

		178		62		1232.46		40000		0		0				1

		179		62		2437.64		40000		0		0				1

		180		62		3483.42		40000		0		0				1

		181		5		318.251		948.007		0		0				0

		182		5		384.209		923.246		0		0				0

		183		5		443.231		896.407		0		0				0

		184		5		495.469		868.626		0		0				0

		185		5		541.356		840.793		0		0				0

		186		5		581.484		813.558		0		0				0

		187		5		616.503		787.352		0		0				0

		188		5		-616.503		787.352		0		0				0

		189		5		-581.484		813.558		0		0				0

		190		5		-541.356		840.793		0		0				0

		191		5		-495.469		868.626		0		0				0

		192		5		-443.231		896.407		0		0				0

		193		5		-384.209		923.246		0		0				0

		194		5		-318.251		948.007		0		0				0

		195		62		-3483.42		40000		0		0				1

		196		62		-2437.64		40000		0		0				1

		197		62		-1232.46		40000		0		0				1

		198		62		0.000155958		40000		0		0				1

		199		61		-0.000157547		40000		0		0				1

		200		61		1302		40000		0		0				1

		201		61		2579		40000		0		0				1

		202		61		3839.17		40000		0		0				1

		203		5		323.614		946.189		0		0				0

		204		5		390.598		920.561		0		0				0

		205		5		450.501		892.776		0		0				0

		206		5		503.477		864.009		0		0				0

		207		5		549.973		835.183		0		0				0

		208		5		590.595		806.968		0		0				0

		209		5		626.009		779.815		0		0				0

		210		5		-626.009		779.815		0		0				0

		211		5		-590.595		806.968		0		0				0

		212		5		-549.973		835.183		0		0				0

		213		5		-503.477		864.009		0		0				0

		214		5		-450.501		892.776		0		0				0

		215		5		-390.598		920.561		0		0				0

		216		5		-323.614		946.189		0		0				0

		217		61		-3839.17		40000		0		0				1

		218		61		-2579		40000		0		0				1

		219		61		-1302		40000		0		0				1

		220		61		-0.000157547		40000		0		0				1

																				%%

																94				42.7

																126				57.3





20 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		67		0.000158044		40000		0		0				1

		2		67		434.079		40000		0		0				1

		3		67		856.716		40000		0		0				1

		4		67		1258.41		40000		0		0				1

		5		67		1632.72		40000		0		0				1

		6		67		1977.31		40000		0		0				1

		7		67		2294.03		40000		0		0				1

		8		67		2585.96		40000		0		0				1

		9		67		2841.43		40000		0		0				1

		10		67		3009.96		40000		0		0				1

		11		67		3094.13		40000		0		0				1

		12		67		-3094.13		40000		0		0				1

		13		67		-3009.96		40000		0		0				1

		14		67		-2841.43		40000		0		0				1

		15		67		-2585.96		40000		0		0				1

		16		67		-2294.03		40000		0		0				1

		17		67		-1977.31		40000		0		0				1

		18		67		-1632.72		40000		0		0				1

		19		67		-1258.41		40000		0		0				1

		20		67		-856.716		40000		0		0				1

		21		67		-434.079		40000		0		0				1

		22		67		0.000158044		40000		0		0				1

		23		66		0.00015304		40000		0		0				1

		24		66		587.969		40000		0		0				1

		25		66		1159.45		40000		0		0				1

		26		66		1699.77		40000		0		0				1

		27		66		2197.43		40000		0		0				1

		28		67		2641.33		40000		0		0				1

		29		67		3001.1		40000		0		0				1

		30		7		269.152         9		63.098		0		0				0

		31		7		294.458         9		55.664		0		0				0

		32		7		316.768         9		48.503		0		0				0

		33		7		336.39         9		41.723		0		0				0

		34		7		-336.39         9		41.723		0		0				0

		35		7		-316.768        9		48.503		0		0				0

		36		7		-294.458        9		55.664		0		0				0

		37		7		-269.152        9		63.098		0		0				0

		38		67		-3001.1		40000		0		0				1

		39		67		-2641.33		40000		0		0				1

		40		66		-2197.43		40000		0		0				1

		41		66		-1699.77		40000		0		0				1

		42		66		-1159.45		40000		0		0				1

		43		66		-587.969		40000		0		0				1

		44		66		0.00015304		40000		0		0				1

		45		65		0.000145476		40000		0		0				1

		46		65		700.749		40000		0		0				1

		47		66		1382.78		40000		0		0				1

		48		66		2027.74		40000		0		0				1

		49		66		2613.61		40000		0		0				1

		50		66		3097.39		40000		0		0				1

		51		6		275.903         9		61.185		0		0				0

		52		6		308.778         9		51.134		0		0				0

		53		6		337.916         9		41.176		0		0				0

		54		6		363.626         9		31.545		0		0				0

		55		6		386.257         9		22.391		0		0				0

		56		6		-386.257        9		22.391		0		0				0

		57		6		-363.626        9		31.545		0		0				0

		58		6		-337.916        9		41.176		0		0				0

		59		6		-308.778        9		51.134		0		0				0

		60		6		-275.903        9		61.185		0		0				0

		61		66		-3097.44		40000		0		0				1

		62		66		-2613.61		40000		0		0				1

		63		66		-2027.74		40000		0		0				1

		64		66		-1382.78		40000		0		0				1

		65		65		-700.749		40000		0		0				1

		66		65		0.000145476		40000		0		0				1

		67		65		0.000188678		40000		0		0				1

		68		65		791.119		40000		0		0				1

		69		65		1564.3		40000		0		0				1

		70		65		2308.16		40000		0		0				1

		71		65		3023.92		40000		0		0				1

		72		65		3544.98		40000		0		0				1

		73		6		301.576         9		53.442		0		0				0

		74		6		337.566         9		41.302		0		0				0

		75		6		369.489         9		29.235		0		0				0

		76		6		397.651         9		17.537		0		0				0

		77		6		422.438         9		6.392		0		0				0

		78		6		-422.438        9		6.392		0		0				0

		79		6		-397.651        9		17.537		0		0				0

		80		6		-369.489        9		29.235		0		0				0

		81		6		-337.566        9		41.302		0		0				0

		82		6		-301.576        9		53.442		0		0				0

		83		65		-3544.98		40000		0		0				1

		84		65		-3023.92		40000		0		0				1

		85		65		-2308.16		40000		0		0				1

		86		65		-1564.3		40000		0		0				1

		87		65		-791.119		40000		0		0				1

		88		65		0.000188678		40000		0		0				1

		89		64		0.000146157		40000		0		0				1

		90		64		865.088		40000		0		0				1

		91		64		1709.22		40000		0		0				1

		92		65		2517.05		40000		0		0				1

		93		65		3279.96		40000		0		0				1

		94		6		278.181         9		60.529		0		0				0

		95		6		321.299         9		46.978		0		0				0

		96		6		359.738         9		33.053		0		0				0

		97		6		393.771         9		19.209		0		0				0

		98		6		423.775         9		5.767		0		0				0

		99		6		450.267         8		92.894		0		0				0

		100		6		-450.267        8		92.894		0		0				0

		101		6		-423.775        9		5.767		0		0				0

		102		6		-393.771        9		19.209		0		0				0

		103		6		-359.738        9		33.053		0		0				0

		104		6		-321.299        9		46.978		0		0				0

		105		6		-278.181        9		60.529		0		0				0

		106		65		-3279.96		40000		0		0				1

		107		65		-2517.05		40000		0		0				1

		108		64		-1709.22		40000		0		0				1

		109		64		-865.088		40000		0		0				1

		110		64		0.000146157		40000		0		0				1

		111		63		0.00015277		40000		0		0				1

		112		63		987.064		40000		0		0				1

		113		63		1946.24		40000		0		0				1

		114		64		2853.83		40000		0		0				1

		115		64		3650.78		40000		0		0				1

		116		6		303.749         9		52.752		0		0				0

		117		6		350.758         9		36.466		0		0				0

		118		6		392.651         9		19.688		0		0				0

		119		6		429.719         9		2.962		0		0				0

		120		6		462.38         8		86.682		0		0				0

		121		6		491.093         8		71.107		0		0				0

		122		6		-491.093        8		71.107		0		0				0

		123		6		-462.38         8		86.682		0		0				0

		124		6		-429.719        9		2.962		0		0				0

		125		6		-392.651        9		19.688		0		0				0

		126		6		-350.758        9		36.466		0		0				0

		127		6		-303.749        9		52.752		0		0				0

		128		64		-3650.78		40000		0		0				1

		129		64		-2853.83		40000		0		0				1

		130		63		-1946.24		40000		0		0				1

		131		63		-987.064		40000		0		0				1

		132		63		0.00015277		40000		0		0				1

		133		62		9.41E-05		40000		0		0				1

		134		62		1091.47		40000		0		0				1

		135		62		2152.65		40000		0		0				1

		136		63		3143.39		40000		0		0				1

		137		5		266.782         9		63.757		0		0				0

		138		5		322.019         9		46.733		0		0				0

		139		5		371.75         9		28.333		0		0				0

		140		5		416.047         9		9.343		0		0				0

		141		5		455.209         8		90.385		0		0				0

		142		5		489.679         8		71.903		0		0				0

		143		5		519.961		854.19		0		0				0

		144		5		-519.961		854.19		0		0				0

		145		5		-489.679        8		71.903		0		0				0

		146		5		-455.209        8		90.385		0		0				0

		147		5		-416.047        9		9.343		0		0				0

		148		5		-371.75         9		28.333		0		0				0

		149		5		-322.019        9		46.733		0		0				0

		150		5		-266.782        9		63.757		0		0				0

		151		63		-3143.39		40000		0		0				1

		152		62		-2152.65		40000		0		0				1

		153		62		-1091.47		40000		0		0				1

		154		62		9.41E-05		40000		0		0				1

		155		61		8.77E-05		40000		0		0				1

		156		61		1182.55		40000		0		0				1

		157		62		2341.07		40000		0		0				1

		158		62		3445.24		40000		0		0				1

		159		5		278.402         9		60.465		0		0				0

		160		5		335.82         9		41.926		0		0				0

		161		5		387.633         9		21.814		0		0				0

		162		5		433.742         9		1.037		0		0				0

		163		5		474.46         8		80.277		0		0				0

		164		5		510.261         8		60.019		0		0				0

		165		5		541.678         8		40.586		0		0				0

		166		5		-541.678        8		40.586		0		0				0

		167		5		-510.261        8		60.019		0		0				0

		168		5		-474.46         8		80.277		0		0				0

		169		5		-433.742        9		1.037		0		0				0

		170		5		-387.633        9		21.814		0		0				0

		171		5		-335.82         9		41.926		0		0				0

		172		5		-278.402        9		60.465		0		0				0

		173		62		-3445.24		40000		0		0				1

		174		62		-2341.07		40000		0		0				1

		175		61		-1182.55		40000		0		0				1

		176		61		8.77E-05		40000		0		0				1

		177		61		0.000107051		40000		0		0				1

		178		61		1259.29		40000		0		0				1

		179		61		2497.86		40000		0		0				1

		180		61		3772.66		40000		0		0				1

		181		5		287.039         9		57.919		0		0				0

		182		5		346.58		938.02		0		0				0

		183		5		400.048         9		16.494		0		0				0

		184		5		447.609         8		94.229		0		0				0

		185		5		489.586         8		71.955		0		0				0

		186		5		526.409         8		50.231		0		0				0

		187		5		558.692         8		29.375		0		0				0

		188		5		-558.692        8		29.375		0		0				0

		189		5		-526.409        8		50.231		0		0				0

		190		5		-489.586        8		71.955		0		0				0

		191		5		-447.609        8		94.229		0		0				0

		192		5		-400.048        9		16.494		0		0				0

		193		5		-346.58		938.02		0		0				0

		194		5		-287.039        9		57.919		0		0				0

		195		61		-3772.66		40000		0		0				1

		196		61		-2497.86		40000		0		0				1

		197		61		-1259.29		40000		0		0				1

		198		61		0.000107051		40000		0		0				1

		199		60		0.000140066		40000		0		0				1

		200		60		1324.69		40000		0		0				1

		201		60		2615.4		40000		0		0				1

		202		61		3800.62		40000		0		0				1

		203		5		294.504		955.65		0		0				0

		204		5		355.453         9		34.694		0		0				0

		205		5		410.202         9		11.995		0		0				0

		206		5		458.817         8		88.531		0		0				0

		207		5		501.67         8		65.059		0		0				0

		208		5		539.28         8		42.126		0		0				0

		209		5		572.224         8		20.098		0		0				0

		210		5		-572.224        8		20.098		0		0				0

		211		5		-539.28         8		42.126		0		0				0

		212		5		-501.67         8		65.059		0		0				0

		213		5		-458.817        8		88.531		0		0				0

		214		5		-410.202        9		11.995		0		0				0

		215		5		-355.453        9		34.694		0		0				0

		216		5		-294.504		955.65		0		0				0

		217		61		-3800.62		40000		0		0				1

		218		60		-2615.4		40000		0		0				1

		219		60		-1324.69		40000		0		0				1

		220		60		0.000140066		40000		0		0				1

																				%%

																112				50.9

																108				49.1





30 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		65		-0.000118038		40000		0		0				1

		2		65		421.741		40000		0		0				1

		3		65		832.279		40000		0		0				1

		4		65		1221.83		40000		0		0				1

		5		65		1582.83		40000		0		0				1

		6		65		1909.92		40000		0		0				1

		7		65		2199.12		40000		0		0				1

		8		66		2447.35		40000		0		0				1

		9		66		2653.11		40000		0		0				1

		10		66		2818.12		40000		0		0				1

		11		66		2947.87		40000		0		0				1

		12		66		-2947.87		40000		0		0				1

		13		66		-2818.12		40000		0		0				1

		14		66		-2653.11		40000		0		0				1

		15		66		-2447.35		40000		0		0				1

		16		65		-2199.12		40000		0		0				1

		17		65		-1909.92		40000		0		0				1

		18		65		-1582.83		40000		0		0				1

		19		65		-1221.83		40000		0		0				1

		20		65		-832.279		40000		0		0				1

		21		65		-421.741		40000		0		0				1

		22		65		-0.000118038		40000		0		0				1

		23		65		0.000221881		40000		0		0				1

		24		65		573.89		40000		0		0				1

		25		65		1132.72		40000		0		0				1

		26		65		1663.94		40000		0		0				1

		27		65		2159.85		40000		0		0				1

		28		65		2620.29		40000		0		0				1

		29		65		3058.77		40000		0		0				1

		30		66		3514.31		40000		0		0				1

		31		65		3121.44		40000		0		0				1

		32		6		272.975         9		62.021		0		0				0

		33		6		289.738         9		57.106		0		0				0

		34		6		-289.738        9		57.106		0		0				0

		35		6		-272.975        9		62.021		0		0				0

		36		65		-3121.44		40000		0		0				1

		37		66		-3514.31		40000		0		0				1

		38		65		-3058.77		40000		0		0				1

		39		65		-2620.29		40000		0		0				1

		40		65		-2159.85		40000		0		0				1

		41		65		-1663.94		40000		0		0				1

		42		65		-1132.72		40000		0		0				1

		43		65		-573.89		40000		0		0				1

		44		65		0.000221881		40000		0		0				1

		45		64		3009.89		40000		0		0				1

		46		64		681.411		40000		0		0				1

		47		64		1342.6		40000		0		0				1

		48		64		1964.74		40000		0		0				1

		49		64		2528.36		40000		0		0				1

		50		65		3005.53		40000		0		0				1

		51		65		3367.27		40000		0		0				1

		52		6		270.025         9		62.853		0		0				0

		53		6		295.274         9		55.413		0		0				0

		54		6		317.54         9		48.245		0		0				0

		55		6		337.139         9		41.455		0		0				0

		56		6		-337.139        9		41.455		0		0				0

		57		6		-317.54         9		48.245		0		0				0

		58		6		-295.274        9		55.413		0		0				0

		59		6		-270.025        9		62.853		0		0				0

		60		65		-3367.27		40000		0		0				1

		61		65		-3005.53		40000		0		0				1

		62		64		-2528.36		40000		0		0				1

		63		64		-1964.74		40000		0		0				1

		64		64		-1342.6		40000		0		0				1

		65		64		-681.411		40000		0		0				1

		66		64		3009.89		40000		0		0				1

		67		63		0.000126995		40000		0		0				1

		68		63		770.792		40000		0		0				1

		69		64		1518.26		40000		0		0				1

		70		64		2223.63		40000		0		0				1

		71		64		2884.22		40000		0		0				1

		72		64		3541.74		40000		0		0				1

		73		6		265.87         9		64.009		0		0				0

		74		6		297.54         9		54.709		0		0				0

		75		6		325.588         9		45.512		0		0				0

		76		6		350.319		936.63		0		0				0

		77		6		372.084         9		28.199		0		0				0

		78		6		-372.084        9		28.199		0		0				0

		79		6		-350.319		936.63		0		0				0

		80		6		-325.588        9		45.512		0		0				0

		81		6		-297.54         9		54.709		0		0				0

		82		6		-265.87         9		64.009		0		0				0

		83		64		-3541.74		40000		0		0				1

		84		64		-2884.22		40000		0		0				1

		85		64		-2223.63		40000		0		0				1

		86		64		-1518.26		40000		0		0				1

		87		63		-770.792		40000		0		0				1

		88		63		0.000126995		40000		0		0				1

		89		63		7.02E-05		40000		0		0				1

		90		63		849.466		40000		0		0				1

		91		63		1672.84		40000		0		0				1

		92		63		2444.64		40000		0		0				1

		93		63		3142.57		40000		0		0				1

		94		64		3670.73		40000		0		0				1

		95		5		285.803         9		58.288		0		0				0

		96		5		319.769         9		47.495		0		0				0

		97		5		349.874         9		36.797		0		0				0

		98		5		376.44         9		26.441		0		0				0

		99		5		399.833         9		16.588		0		0				0

		100		5		-399.833        9		16.588		0		0				0

		101		5		-376.44         9		26.441		0		0				0

		102		5		-349.874        9		36.797		0		0				0

		103		5		-319.769        9		47.495		0		0				0

		104		5		-285.803        9		58.288		0		0				0

		105		64		-3670.73		40000		0		0				1

		106		63		-3142.57		40000		0		0				1

		107		63		-2444.64		40000		0		0				1

		108		63		-1672.84		40000		0		0				1

		109		63		-849.498		40000		0		0				1

		110		63		7.02E-05		40000		0		0				1

		111		62		0.000146897		40000		0		0				1

		112		62		983.744		40000		0		0				1

		113		62		1946.1		40000		0		0				1

		114		62		2865.06		40000		0		0				1

		115		62		3665.78		40000		0		0				1

		116		5		274.158         9		61.685		0		0				0

		117		5		315.904         9		48.791		0		0				0

		118		5		353.441         9		35.457		0		0				0

		119		5		386.764         9		22.179		0		0				0

		120		5		416.171         9		9.286		0		0				0

		121		5		442.059         8		96.986		0		0				0

		122		5		-442.059        8		96.986		0		0				0

		123		5		-416.171        9		9.286		0		0				0

		124		5		-386.764        9		22.179		0		0				0

		125		5		-353.441        9		35.457		0		0				0

		126		5		-315.904        9		48.791		0		0				0

		127		5		-274.158        9		61.685		0		0				0

		128		62		-3665.78		40000		0		0				1

		129		62		-2865.06		40000		0		0				1

		130		62		-1946.1		40000		0		0				1

		131		62		-983.744		40000		0		0				1

		132		62		0.000146897		40000		0		0				1

		133		61		0.000180192		40000		0		0				1

		134		61		1087.14		40000		0		0				1

		135		61		2156.25		40000		0		0				1

		136		62		3222.12		40000		0		0				1

		137		62		4694.73		40000		0		0				1

		138		5		292.423         9		56.289		0		0				0

		139		5		337.67         9		41.265		0		0				0

		140		5		377.99		925.81		0		0				0

		141		5		413.781         9		10.377		0		0				0

		142		5		445.361         8		95.351		0		0				0

		143		5		473.042		881.04		0		0				0

		144		5		-473.042		881.04		0		0				0

		145		5		-445.361        8		95.351		0		0				0

		146		5		-413.781        9		10.377		0		0				0

		147		5		-377.99		925.81		0		0				0

		148		5		-337.67         9		41.265		0		0				0

		149		5		-292.423        9		56.289		0		0				0

		150		62		-4694.73		40000		0		0				1

		151		62		-3222.12		40000		0		0				1

		152		61		-2156.25		40000		0		0				1

		153		61		-1087.14		40000		0		0				1

		154		61		0.000180192		40000		0		0				1

		155		61		0.000147159		40000		0		0				1

		156		61		1168.26		40000		0		0				1

		157		61		2302.89		40000		0		0				1

		158		61		3349.71		40000		0		0				1

		159		62		4142.1		40000		0		0				1

		160		5		307.933         9		51.408		0		0				0

		161		5		355.205         9		34.788		0		0				0

		162		5		397.34         9		17.671		0		0				0

		163		5		434.777         9		0.538		0		0				0

		164		5		467.753         8		83.859		0		0				0

		165		5		496.745         8		67.896		0		0				0

		166		5		-496.745        8		67.896		0		0				0

		167		5		-467.753        8		83.859		0		0				0

		168		5		-434.777        9		0.538		0		0				0

		169		5		-397.34         9		17.671		0		0				0

		170		5		-355.205        9		34.788		0		0				0

		171		5		-307.933        9		51.408		0		0				0

		172		62		-4142.1		40000		0		0				1

		173		61		-3349.71		40000		0		0				1

		174		61		-2302.89		40000		0		0				1

		175		61		-1168.26		40000		0		0				1

		176		61		0.000147159		40000		0		0				1

		177		60		7.50E-05		40000		0		0				1

		178		60		1240.95		40000		0		0				1

		179		60		2441.07		40000		0		0				1

		180		60		3536.73		40000		0		0				1

		181		5		265.119         9		64.216		0		0				0

		182		5		319.853         9		47.467		0		0				0

		183		5		369.18         9		29.358		0		0				0

		184		5		413.139         9		10.668		0		0				0

		185		5		451.874         8		92.082		0		0				0

		186		5		485.915         8		74.006		0		0				0

		187		5		515.953         8		56.617		0		0				0

		188		5		-515.953        8		56.617		0		0				0

		189		5		-485.915        8		74.006		0		0				0

		190		5		-451.8         8		92.119		0		0				0

		191		5		-413.139        9		10.668		0		0				0

		192		5		-369.18         9		29.358		0		0				0

		193		5		-319.853        9		47.467		0		0				0

		194		5		-265.119        9		64.216		0		0				0

		195		60		-3536.73		40000		0		0				1

		196		60		-2441.07		40000		0		0				1

		197		60		-1240.95		40000		0		0				1

		198		60		7.50E-05		40000		0		0				1

		199		59		-4.52E-05		40000		0		0				1

		200		59		1309.61		40000		0		0				1

		201		59		2582.55		40000		0		0				1

		202		60		3804.62		40000		0		0				1

		203		5		273.155		961.97		0		0				0

		204		5		329.684         9		44.091		0		0				0

		205		5		380.544         9		24.763		0		0				0

		206		5		425.81         9		4.813		0		0				0

		207		5		465.814         8		84.883		0		0				0

		208		5		501.017         8		65.437		0		0				0

		209		5		531.734         8		46.911		0		0				0

		210		5		-531.734        8		46.911		0		0				0

		211		5		-501.017        8		65.437		0		0				0

		212		5		-465.814        8		84.883		0		0				0

		213		5		-425.81         9		4.813		0		0				0

		214		5		-380.544        9		24.763		0		0				0

		215		5		-329.684        9		44.091		0		0				0

		216		5		-273.155		961.97		0		0				0

		217		60		-3804.62		40000		0		0				1

		218		59		-2582.55		40000		0		0				1

		219		59		-1309.61		40000		0		0				1

		220		59		-4.52E-05		40000		0		0				1

																				%%

																124				56.4

																96				43.6





40 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		64		0.000170355		40000		0		0				1

		2		64		398.31		40000		0		0				1

		3		64		784.989		40000		0		0				1

		4		64		1150.01		40000		0		0				1

		5		64		1486.02		40000		0		0				1

		6		64		1788.58		40000		0		0				1

		7		64		2055.75		40000		0		0				1

		8		64		2287.34		40000		0		0				1

		9		64		2484.43		40000		0		0				1

		10		65		2649.22		40000		0		0				1

		11		65		2785.02		40000		0		0				1

		12		65		-2785.02		40000		0		0				1

		13		65		-2649.22		40000		0		0				1

		14		64		-2484.43		40000		0		0				1

		15		64		-2287.34		40000		0		0				1

		16		64		-2055.75		40000		0		0				1

		17		64		-1788.58		40000		0		0				1

		18		64		-1486.02		40000		0		0				1

		19		64		-1150.01		40000		0		0				1

		20		64		-784.988		40000		0		0				1

		21		64		-398.309		40000		0		0				1

		22		64		0.000170355		40000		0		0				1

		23		63		0.000216813		40000		0		0				1

		24		63		550.37		40000		0		0				1

		25		63		1083.93		40000		0		0				1

		26		64		1586.08		40000		0		0				1

		27		64		2046.62		40000		0		0				1

		28		64		2461.99		40000		0		0				1

		29		64		2837.5		40000		0		0				1

		30		64		3190.9		40000		0		0				1

		31		64		3571.57		40000		0		0				1

		32		64		4199.74		40000		0		0				1

		33		64		3369.71		40000		0		0				1

		34		64		-3369.71		40000		0		0				1

		35		64		-4199.73		40000		0		0				1

		36		64		-3571.57		40000		0		0				1

		37		64		-3190.9		40000		0		0				1

		38		64		-2837.5		40000		0		0				1

		39		64		-2461.99		40000		0		0				1

		40		64		-2046.62		40000		0		0				1

		41		64		-1586.08		40000		0		0				1

		42		63		-1083.93		40000		0		0				1

		43		63		-550.37		40000		0		0				1

		44		63		0.000216813		40000		0		0				1

		45		63		0.000145076		40000		0		0				1

		46		63		666.496		40000		0		0				1

		47		63		1313.72		40000		0		0				1

		48		63		1922.48		40000		0		0				1

		49		63		2472.85		40000		0		0				1

		50		63		2942.65		40000		0		0				1

		51		63		3314.48		40000		0		0				1

		52		64		3588.46		40000		0		0				1

		53		5		266.649         9		63.794		0		0				0

		54		5		286.454         9		58.094		0		0				0

		55		5		303.874         9		52.712		0		0				0

		56		5		-303.874        9		52.712		0		0				0

		57		5		-286.454        9		58.094		0		0				0

		58		5		-266.649        9		63.794		0		0				0

		59		64		-3588.46		40000		0		0				1

		60		63		-3314.48		40000		0		0				1

		61		63		-2942.65		40000		0		0				1

		62		63		-2472.85		40000		0		0				1

		63		63		-1922.48		40000		0		0				1

		64		63		-1313.72		40000		0		0				1

		65		63		-666.496		40000		0		0				1

		66		63		0.000145076		40000		0		0				1

		67		62		0.00012727		40000		0		0				1

		68		63		761.013		40000		0		0				1

		69		63		1504.92		40000		0		0				1

		70		63		2217.23		40000		0		0				1

		71		63		2882.76		40000		0		0				1

		72		63		3460.05		40000		0		0				1

		73		63		3849.8		40000		0		0				1

		74		5		271.148         9		62.538		0		0				0

		75		5		295.912         9		55.215		0		0				0

		76		5		317.958         9		48.105		0		0				0

		77		5		337.44         9		41.347		0		0				0

		78		5		-337.44         9		41.347		0		0				0

		79		5		-317.958        9		48.105		0		0				0

		80		5		-295.912        9		55.215		0		0				0

		81		5		-271.148        9		62.538		0		0				0

		82		63		-3849.8		40000		0		0				1

		83		63		-3460.05		40000		0		0				1

		84		63		-2882.76		40000		0		0				1

		85		63		-2217.23		40000		0		0				1

		86		63		-1504.92		40000		0		0				1

		87		63		-761.013		40000		0		0				1

		88		62		0.00012727		40000		0		0				1

		89		62		7.85E-05		40000		0		0				1

		90		62		837.26		40000		0		0				1

		91		62		1658.86		40000		0		0				1

		92		62		2459.56		40000		0		0				1

		93		62		3266.08		40000		0		0				1

		94		62		4202.11		40000		0		0				1

		95		61		4422.98		40000		0		0				1

		96		5		291.697         9		56.511		0		0				0

		97		5		319.01         9		47.751		0		0				0

		98		5		343.151		939.28		0		0				0

		99		5		364.408         9		31.239		0		0				0

		100		5		-364.408        9		31.239		0		0				0

		101		5		-343.151		939.28		0		0				0

		102		5		-319.01         9		47.751		0		0				0

		103		5		-291.697        9		56.511		0		0				0

		104		61		-4422.98		40000		0		0				1

		105		62		-4202.11		40000		0		0				1

		106		62		-3266.08		40000		0		0				1

		107		62		-2459.56		40000		0		0				1

		108		62		-1658.86		40000		0		0				1

		109		62		-837.259		40000		0		0				1

		110		62		7.85E-05		40000		0		0				1

		111		61		0.00018682		40000		0		0				1

		112		61		953.206		40000		0		0				1

		113		61		1879.13		40000		0		0				1

		114		62		2748.24		40000		0		0				1

		115		62		3507.77		40000		0		0				1

		116		63		3990.46		40000		0		0				1

		117		5		291.057         9		56.706		0		0				0

		118		5		325.062         9		45.693		0		0				0

		119		5		355.275         9		34.762		0		0				0

		120		5		382.268         9		24.051		0		0				0

		121		5		406.028         9		13.861		0		0				0

		122		5		-406.028        9		13.861		0		0				0

		123		5		-382.268        9		24.051		0		0				0

		124		5		-355.275        9		34.762		0		0				0

		125		5		-325.062        9		45.693		0		0				0

		126		5		-291.057        9		56.706		0		0				0

		127		63		-3990.46		40000		0		0				1

		128		62		-3507.77		40000		0		0				1

		129		62		-2748.24		40000		0		0				1

		130		61		-1879.13		40000		0		0				1

		131		61		-953.205		40000		0		0				1

		132		61		0.00018682		40000		0		0				1

		133		61		0.000147069		40000		0		0				1

		134		61		1049.18		40000		0		0				1

		135		61		2062.38		40000		0		0				1

		136		61		2994.07		40000		0		0				1

		137		61		3768.23		40000		0		0				1

		138		5		271.403         9		62.466		0		0				0

		139		5		313.102		949.72		0		0				0

		140		5		350.356         9		36.617		0		0				0

		141		5		383.397         9		23.584		0		0				0

		142		5		412.564         9		10.929		0		0				0

		143		5		437.922         8		99.013		0		0				0

		144		5		-437.922        8		99.013		0		0				0

		145		5		-412.564        9		10.929		0		0				0

		146		5		-383.397        9		23.584		0		0				0

		147		5		-350.356        9		36.617		0		0				0

		148		5		-313.102		949.72		0		0				0

		149		5		-271.403        9		62.466		0		0				0

		150		61		-3768.23		40000		0		0				1

		151		61		-2994.07		40000		0		0				1

		152		61		-2062.38		40000		0		0				1

		153		61		-1049.18		40000		0		0				1

		154		61		0.000147069		40000		0		0				1

		155		60		9.33E-05		40000		0		0				1

		156		60		1137.17		40000		0		0				1

		157		60		2240.5		40000		0		0				1

		158		60		3292.16		40000		0		0				1

		159		60		4256.46		40000		0		0				1

		160		5		286.331         9		58.131		0		0				0

		161		5		330.528         9		43.796		0		0				0

		162		5		369.952         9		29.051		0		0				0

		163		5		404.869         9		14.375		0		0				0

		164		5		435.668         9		0.108		0		0				0

		165		5		462.781         8		86.473		0		0				0

		166		5		-462.781        8		86.473		0		0				0

		167		5		-435.668        9		0.108		0		0				0

		168		5		-404.869        9		14.375		0		0				0

		169		5		-369.952        9		29.051		0		0				0

		170		5		-330.528        9		43.796		0		0				0

		171		5		-286.331        9		58.131		0		0				0

		172		60		-4256.46		40000		0		0				1

		173		60		-3292.16		40000		0		0				1

		174		60		-2240.5		40000		0		0				1

		175		60		-1137.17		40000		0		0				1

		176		60		9.33E-05		40000		0		0				1

		177		59		0.000124461		40000		0		0				1

		178		59		1217.7		40000		0		0				1

		179		59		2398.81		40000		0		0				1

		180		60		3536.11		40000		0		0				1

		181		60		4434.31		40000		0		0				1

		182		4		298.922         9		54.278		0		0				0

		183		4		344.99         9		38.606		0		0				0

		184		4		386.058         9		22.474		0		0				0

		185		4		422.419         9		6.401		0		0				0

		186		4		454.475         8		90.759		0		0				0

		187		5		482.695         8		75.788		0		0				0

		188		5		-482.695        8		75.788		0		0				0

		189		4		-454.475        8		90.759		0		0				0

		190		4		-422.419        9		6.401		0		0				0

		191		4		-386.058        9		22.474		0		0				0

		192		4		-344.99         9		38.606		0		0				0

		193		4		-298.922        9		54.278		0		0				0

		194		60		-4434.31		40000		0		0				1

		195		60		-3536.11		40000		0		0				1

		196		59		-2398.81		40000		0		0				1

		197		59		-1217.7		40000		0		0				1

		198		59		0.000124461		40000		0		0				1

		199		58		-0.000136824		40000		0		0				1

		200		59		1293.08		40000		0		0				1

		201		59		2539.83		40000		0		0				1

		202		59		3681.37		40000		0		0				1

		203		59		4546.54		40000		0		0				1

		204		4		309.65         9		50.851		0		0				0

		205		4		357.24         9		34.013		0		0				0

		206		4		399.689         9		16.651		0		0				0

		207		4		437.269         8		99.331		0		0				0

		208		4		470.388		882.46		0		0				0

		209		4		499.512         8		66.307		0		0				0

		210		4		-499.512        8		66.307		0		0				0

		211		4		-470.388		882.46		0		0				0

		212		4		-437.269        8		99.331		0		0				0

		213		4		-399.689        9		16.651		0		0				0

		214		4		-357.24         9		34.013		0		0				0

		215		4		-309.65         9		50.851		0		0				0

		216		59		-4546.54		40000		0		0				1

		217		59		-3681.37		40000		0		0				1

		218		59		-2539.83		40000		0		0				1

		219		59		-1293.08		40000		0		0				1

		220		58		-0.000136824		40000		0		0				1

																				%%

																140				63.6

																80				36.4





50 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		63		0.000195125		40000		0		0				1

		2		63		389.598		40000		0		0				1

		3		63		767.91		40000		0		0				1

		4		63		1124.95		40000		0		0				1

		5		63		1452.98		40000		0		0				1

		6		63		1746.98		40000		0		0				1

		7		63		2004.72		40000		0		0				1

		8		63		2226.47		40000		0		0				1

		9		63		2414.62		40000		0		0				1

		10		63		2572.93		40000		0		0				1

		11		63		2705.79		40000		0		0				1

		12		63		-2705.79		40000		0		0				1

		13		63		-2572.93		40000		0		0				1

		14		63		-2414.62		40000		0		0				1

		15		63		-2226.47		40000		0		0				1

		16		63		-2004.72		40000		0		0				1

		17		63		-1746.98		40000		0		0				1

		18		63		-1452.98		40000		0		0				1

		19		63		-1124.95		40000		0		0				1

		20		63		-767.91		40000		0		0				1

		21		63		-389.598		40000		0		0				1

		22		63		0.000195125		40000		0		0				1

		23		63		0.000174848		40000		0		0				1

		24		63		542.287		40000		0		0				1

		25		63		1071.84		40000		0		0				1

		26		63		1578.1		40000		0		0				1

		27		63		2053.87		40000		0		0				1

		28		63		2494.83		40000		0		0				1

		29		63		2896.54		40000		0		0				1

		30		63		3248.79		40000		0		0				1

		31		63		3531.27		40000		0		0				1

		32		62		3722.8		40000		0		0				1

		33		62		3828.3		40000		0		0				1

		34		62		-3828.3		40000		0		0				1

		35		62		-3722.8		40000		0		0				1

		36		63		-3531.27		40000		0		0				1

		37		63		-3248.79		40000		0		0				1

		38		63		-2896.54		40000		0		0				1

		39		63		-2494.83		40000		0		0				1

		40		63		-2053.87		40000		0		0				1

		41		63		-1578.1		40000		0		0				1

		42		63		-1071.84		40000		0		0				1

		43		63		-542.287		40000		0		0				1

		44		63		0.000174848		40000		0		0				1

		45		62		0.000173054		40000		0		0				1

		46		62		645.921		40000		0		0				1

		47		62		1276.59		40000		0		0				1

		48		62		1880.57		40000		0		0				1

		49		63		2454.11		40000		0		0				1

		50		63		3006.29		40000		0		0				1

		51		63		3570.28		40000		0		0				1

		52		63		4250.19		40000		0		0				1

		53		64		5484.49		40000		0		0				1

		54		63		3424.55		40000		0		0				1

		55		5		278.67         9		60.387		0		0				0

		56		5		-278.67         9		60.387		0		0				0

		57		63		-3424.55		40000		0		0				1

		58		64		-5484.49		40000		0		0				1

		59		63		-4250.19		40000		0		0				1

		60		63		-3570.28		40000		0		0				1

		61		63		-3006.29		40000		0		0				1

		62		63		-2454.11		40000		0		0				1

		63		62		-1880.57		40000		0		0				1

		64		62		-1276.59		40000		0		0				1

		65		62		-645.921		40000		0		0				1

		66		62		0.000173054		40000		0		0				1

		67		62		-9.24E-05		40000		0		0				1

		68		62		725.586		40000		0		0				1

		69		62		1430.01		40000		0		0				1

		70		62		2093.68		40000		0		0				1

		71		62		2699.76		40000		0		0				1

		72		63		3230.03		40000		0		0				1

		73		63		3647.01		40000		0		0				1

		74		64		3854.15		40000		0		0				1

		75		5		274.207         9		61.671		0		0				0

		76		5		294.181		955.75		0		0				0

		77		5		311.439         9		50.266		0		0				0

		78		5		-311.439        9		50.266		0		0				0

		79		5		-294.181		955.75		0		0				0

		80		5		-274.207        9		61.671		0		0				0

		81		64		-3854.15		40000		0		0				1

		82		63		-3647.01		40000		0		0				1

		83		63		-3230.03		40000		0		0				1

		84		62		-2699.76		40000		0		0				1

		85		62		-2093.68		40000		0		0				1

		86		62		-1430.01		40000		0		0				1

		87		62		-725.586		40000		0		0				1

		88		62		-9.24E-05		40000		0		0				1

		89		61		0.000220684		40000		0		0				1

		90		61		794.229		40000		0		0				1

		91		61		1562.02		40000		0		0				1

		92		62		2275.96		40000		0		0				1

		93		62		2904.87		40000		0		0				1

		94		62		3414.74		40000		0		0				1

		95		63		3788.3		40000		0		0				1

		96		5		271.226         9		62.516		0		0				0

		97		5		296.422         9		55.057		0		0				0

		98		5		318.643         9		47.875		0		0				0

		99		5		338.224         9		41.066		0		0				0

		100		5		-338.224        9		41.066		0		0				0

		101		5		-318.643        9		47.875		0		0				0

		102		5		-296.422        9		55.057		0		0				0

		103		5		-271.226        9		62.516		0		0				0

		104		63		-3788.3		40000		0		0				1

		105		62		-3414.74		40000		0		0				1

		106		62		-2904.87		40000		0		0				1

		107		62		-2275.96		40000		0		0				1

		108		61		-1562.02		40000		0		0				1

		109		61		-794.229		40000		0		0				1

		110		61		0.000220684		40000		0		0				1

		111		61		0.000204379		40000		0		0				1

		112		61		915.591		40000		0		0				1

		113		61		1802.44		40000		0		0				1

		114		61		2636.81		40000		0		0				1

		115		61		3396.51		40000		0		0				1

		116		61		3999.65		40000		0		0				1

		117		5		270.971         9		62.587		0		0				0

		118		5		303.087         9		52.963		0		0				0

		119		5		331.588         9		43.424		0		0				0

		120		5		356.744         9		34.202		0		0				0

		121		5		378.897         9		25.439		0		0				0

		122		5		-378.897        9		25.439		0		0				0

		123		5		-356.744        9		34.202		0		0				0

		124		5		-331.588        9		43.424		0		0				0

		125		5		-303.087        9		52.963		0		0				0

		126		5		-270.971        9		62.587		0		0				0

		127		61		-3999.65		40000		0		0				1

		128		61		-3396.51		40000		0		0				1

		129		61		-2636.81		40000		0		0				1

		130		61		-1802.44		40000		0		0				1

		131		61		-915.591		40000		0		0				1

		132		61		0.000204379		40000		0		0				1

		133		60		0.000140004		40000		0		0				1

		134		60		1021.07		40000		0		0				1

		135		60		2010.92		40000		0		0				1

		136		60		2954.32		40000		0		0				1

		137		60		3909.31		40000		0		0				1

		138		61		3988.77		40000		0		0				1

		139		4		292.91		956.14		0		0				0

		140		4		327.75         9		44.764		0		0				0

		141		4		358.607         9		33.489		0		0				0

		142		5		385.869         9		22.553		0		0				0

		143		5		409.887         9		12.136		0		0				0

		144		5		-409.887        9		12.136		0		0				0

		145		5		-385.869        9		22.553		0		0				0

		146		4		-358.607        9		33.489		0		0				0

		147		4		-327.75         9		44.764		0		0				0

		148		4		-292.91		956.14		0		0				0

		149		61		-3988.77		40000		0		0				1

		150		60		-3909.31		40000		0		0				1

		151		60		-2954.32		40000		0		0				1

		152		60		-2010.92		40000		0		0				1

		153		60		-1021.07		40000		0		0				1

		154		60		0.000140004		40000		0		0				1

		155		59		0.000207283		40000		0		0				1

		156		59		1115.96		40000		0		0				1

		157		59		2191.14		40000		0		0				1

		158		60		3183.51		40000		0		0				1

		159		60		4021.38		40000		0		0				1

		160		4		269.71         9		62.942		0		0				0

		161		4		311.051         9		50.393		0		0				0

		162		4		347.937         9		37.518		0		0				0

		163		4		380.639         9		24.724		0		0				0

		164		4		409.504         9		12.308		0		0				0

		165		4		434.926         9		0.466		0		0				0

		166		4		-434.926        9		0.466		0		0				0

		167		4		-409.504        9		12.308		0		0				0

		168		4		-380.639        9		24.724		0		0				0

		169		4		-347.937        9		37.518		0		0				0

		170		4		-311.051        9		50.393		0		0				0

		171		4		-269.71         9		62.942		0		0				0

		172		60		-4021.38		40000		0		0				1

		173		60		-3183.51		40000		0		0				1

		174		59		-2191.14		40000		0		0				1

		175		59		-1115.96		40000		0		0				1

		176		59		0.000207283		40000		0		0				1

		177		59		0.000160992		40000		0		0				1

		178		59		1205.89		40000		0		0				1

		179		59		2365.67		40000		0		0				1

		180		59		3412.54		40000		0		0				1

		181		59		4255.73		40000		0		0				1

		182		4		282.857         9		59.162		0		0				0

		183		4		325.99         9		45.373		0		0				0

		184		4		364.575         9		31.174		0		0				0

		185		4		398.817         9		17.031		0		0				0

		186		4		429.046         9		3.283		0		0				0

		187		4		455.666         8		90.151		0		0				0

		188		4		-455.666        8		90.151		0		0				0

		189		4		-429.046        9		3.283		0		0				0

		190		4		-398.817        9		17.031		0		0				0

		191		4		-364.575        9		31.174		0		0				0

		192		4		-325.99         9		45.373		0		0				0

		193		4		-282.857        9		59.162		0		0				0

		194		59		-4255.73		40000		0		0				1

		195		59		-3412.54		40000		0		0				1

		196		59		-2365.67		40000		0		0				1

		197		59		-1205.89		40000		0		0				1

		198		59		0.000160992		40000		0		0				1

		199		58		0.000139431		40000		0		0				1

		200		58		1291.26		40000		0		0				1

		201		58		2543.86		40000		0		0				1

		202		58		3688.66		40000		0		0				1

		203		58		4587.96		40000		0		0				1

		204		4		293.86         9		55.848		0		0				0

		205		4		338.512         9		40.962		0		0				0

		206		4		378.586         9		25.566		0		0				0

		207		4		414.158         9		10.205		0		0				0

		208		4		445.548         8		95.258		0		0				0

		209		4		473.177         8		80.968		0		0				0

		210		4		-473.177        8		80.968		0		0				0

		211		4		-445.548        8		95.258		0		0				0

		212		4		-414.158        9		10.205		0		0				0

		213		4		-378.586        9		25.566		0		0				0

		214		4		-338.512        9		40.962		0		0				0

		215		4		-293.86         9		55.848		0		0				0

		216		58		-4587.96		40000		0		0				1

		217		58		-3688.66		40000		0		0				1

		218		58		-2543.86		40000		0		0				1

		219		58		-1291.26		40000		0		0				1

		220		58		0.000139431		40000		0		0				1

																				%%

																148				67.3

																72				32.7





60 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		62		0.00017284		40000		0		0				1

		2		62		375.834		40000		0		0				1

		3		62		741.983		40000		0		0				1

		4		63		1090.23		40000		0		0				1

		5		63		1414.81		40000		0		0				1

		6		63		1712.76		40000		0		0				1

		7		63		1983.53		40000		0		0				1

		8		63		2228.33		40000		0		0				1

		9		63		2449.31		40000		0		0				1

		10		63		2648.88		40000		0		0				1

		11		63		2829.23		40000		0		0				1

		12		63		-2829.23		40000		0		0				1

		13		63		-2648.88		40000		0		0				1

		14		63		-2449.31		40000		0		0				1

		15		63		-2228.33		40000		0		0				1

		16		63		-1983.53		40000		0		0				1

		17		63		-1712.76		40000		0		0				1

		18		63		-1414.81		40000		0		0				1

		19		63		-1090.23		40000		0		0				1

		20		62		-741.983		40000		0		0				1

		21		62		-375.834		40000		0		0				1

		22		62		0.00017284		40000		0		0				1

		23		62		0.000222919		40000		0		0				1

		24		62		508.367		40000		0		0				1

		25		62		1001.52		40000		0		0				1

		26		62		1466.33		40000		0		0				1

		27		62		1892.82		40000		0		0				1

		28		62		2275.08		40000		0		0				1

		29		63		2611.16		40000		0		0				1

		30		63		2902.27		40000		0		0				1

		31		63		3151.04		40000		0		0				1

		32		63		3360.53		40000		0		0				1

		33		64		3531.42		40000		0		0				1

		34		64		-3531.42		40000		0		0				1

		35		63		-3360.53		40000		0		0				1

		36		63		-3151.04		40000		0		0				1

		37		63		-2902.27		40000		0		0				1

		38		63		-2611.16		40000		0		0				1

		39		62		-2275.08		40000		0		0				1

		40		62		-1892.82		40000		0		0				1

		41		62		-1466.33		40000		0		0				1

		42		62		-1001.52		40000		0		0				1

		43		62		-508.367		40000		0		0				1

		44		62		0.000222919		40000		0		0				1

		45		61		0.000123152		40000		0		0				1

		46		62		607.358		40000		0		0				1

		47		62		1195.21		40000		0		0				1

		48		62		1745.65		40000		0		0				1

		49		62		2243.2		40000		0		0				1

		50		62		2675.15		40000		0		0				1

		51		62		3032.9		40000		0		0				1

		52		62		3316.27		40000		0		0				1

		53		63		3533.01		40000		0		0				1

		54		63		3695.54		40000		0		0				1

		55		63		3816.16		40000		0		0				1

		56		63		-3816.16		40000		0		0				1

		57		63		-3695.54		40000		0		0				1

		58		63		-3533.01		40000		0		0				1

		59		62		-3316.27		40000		0		0				1

		60		62		-3032.9		40000		0		0				1

		61		62		-2675.15		40000		0		0				1

		62		62		-2243.2		40000		0		0				1

		63		62		-1745.65		40000		0		0				1

		64		62		-1195.21		40000		0		0				1

		65		62		-607.358		40000		0		0				1

		66		61		0.000123152		40000		0		0				1

		67		61		0.000278651		40000		0		0				1

		68		61		691.436		40000		0		0				1

		69		61		1361.42		40000		0		0				1

		70		61		1991.39		40000		0		0				1

		71		61		2566.42		40000		0		0				1

		72		61		3071.2		40000		0		0				1

		73		61		3482.23		40000		0		0				1

		74		62		3778.74		40000		0		0				1

		75		62		3983.87		40000		0		0				1

		76		5		274.513         9		61.583		0		0				0

		77		5		291.265         9		56.642		0		0				0

		78		5		-291.265        9		56.642		0		0				0

		79		5		-274.513        9		61.583		0		0				0

		80		62		-3983.87		40000		0		0				1

		81		62		-3778.74		40000		0		0				1

		82		61		-3482.23		40000		0		0				1

		83		61		-3071.2		40000		0		0				1

		84		61		-2566.42		40000		0		0				1

		85		61		-1991.39		40000		0		0				1

		86		61		-1361.42		40000		0		0				1

		87		61		-691.436		40000		0		0				1

		88		61		0.000278651		40000		0		0				1

		89		61		0.000160691		40000		0		0				1

		90		61		765.347		40000		0		0				1

		91		61		1507.85		40000		0		0				1

		92		61		2210.79		40000		0		0				1

		93		61		2869.62		40000		0		0				1

		94		61		3494.18		40000		0		0				1

		95		61		4063.66		40000		0		0				1

		96		60		4304.54		40000		0		0				1

		97		5		276.971         9		60.878		0		0				0

		98		5		297.819         9		54.622		0		0				0

		99		5		316.165         9		48.704		0		0				0

		100		5		-316.165        9		48.704		0		0				0

		101		5		-297.819        9		54.622		0		0				0

		102		5		-276.971        9		60.878		0		0				0

		103		60		-4304.54		40000		0		0				1

		104		61		-4063.66		40000		0		0				1

		105		61		-3494.18		40000		0		0				1

		106		61		-2869.62		40000		0		0				1

		107		61		-2210.79		40000		0		0				1

		108		61		-1507.85		40000		0		0				1

		109		61		-765.347		40000		0		0				1

		110		61		0.000160691		40000		0		0				1

		111		60		0.000228992		40000		0		0				1

		112		60		892.955		40000		0		0				1

		113		60		1755.46		40000		0		0				1

		114		60		2562.71		40000		0		0				1

		115		60		3301.76		40000		0		0				1

		116		61		3955.96		40000		0		0				1

		117		62		4176.86		40000		0		0				1

		118		4		285.231         9		58.459		0		0				0

		119		4		311.896         9		50.116		0		0				0

		120		4		335.402         9		42.075		0		0				0

		121		4		356.087         9		34.453		0		0				0

		122		4		-356.087        9		34.453		0		0				0

		123		4		-335.402        9		42.075		0		0				0

		124		4		-311.896        9		50.116		0		0				0

		125		4		-285.231        9		58.459		0		0				0

		126		62		-4176.86		40000		0		0				1

		127		61		-3955.96		40000		0		0				1

		128		60		-3301.76		40000		0		0				1

		129		60		-2562.71		40000		0		0				1

		130		60		-1755.49		40000		0		0				1

		131		60		-892.955		40000		0		0				1

		132		60		0.000228992		40000		0		0				1

		133		59		0.000140241		40000		0		0				1

		134		59		1006.92		40000		0		0				1

		135		59		1975.3		40000		0		0				1

		136		59		2860.91		40000		0		0				1

		137		60		3610.47		40000		0		0				1

		138		60		4178.92		40000		0		0				1

		139		4		277.601         9		60.696		0		0				0

		140		4		310.213         9		50.667		0		0				0

		141		4		339.144         9		40.734		0		0				0

		142		4		364.702         9		31.124		0		0				0

		143		4		387.225         9		21.985		0		0				0

		144		4		-387.225        9		21.985		0		0				0

		145		4		-364.702        9		31.124		0		0				0

		146		4		-339.144        9		40.734		0		0				0

		147		4		-310.213        9		50.667		0		0				0

		148		4		-277.601        9		60.696		0		0				0

		149		60		-4178.92		40000		0		0				1

		150		60		-3610.47		40000		0		0				1

		151		59		-2860.91		40000		0		0				1

		152		59		-1975.3		40000		0		0				1

		153		59		-1006.92		40000		0		0				1

		154		59		0.000140241		40000		0		0				1

		155		59		0.000228447		40000		0		0				1

		156		59		1113.19		40000		0		0				1

		157		59		2190.55		40000		0		0				1

		158		59		3181.43		40000		0		0				1

		159		59		4003.33		40000		0		0				1

		160		59		4604.93		40000		0		0				1

		161		4		295.661         9		55.293		0		0				0

		162		4		330.248         9		43.894		0		0				0

		163		4		361.057         9		32.544		0		0				0

		164		4		388.315         9		21.527		0		0				0

		165		4		412.348         9		11.026		0		0				0

		166		4		-412.348        9		11.026		0		0				0

		167		4		-388.315        9		21.527		0		0				0

		168		4		-361.057        9		32.544		0		0				0

		169		4		-330.248        9		43.894		0		0				0

		170		4		-295.661        9		55.293		0		0				0

		171		59		-4604.93		40000		0		0				1

		172		59		-4003.33		40000		0		0				1

		173		59		-3181.43		40000		0		0				1

		174		59		-2190.55		40000		0		0				1

		175		59		-1113.19		40000		0		0				1

		176		59		0.000228447		40000		0		0				1

		177		58		0.000202763		40000		0		0				1

		178		58		1208.26		40000		0		0				1

		179		58		2395.63		40000		0		0				1

		180		58		3536.03		40000		0		0				1

		181		58		4513.56		40000		0		0				1

		182		4		270.387         9		62.752		0		0				0

		183		4		310.35         9		50.623		0		0				0

		184		4		346.757         9		37.955		0		0				0

		185		4		379.238         9		25.299		0		0				0

		186		4		407.957         9		13.001		0		0				0

		187		4		433.262         9		1.268		0		0				0

		188		4		-433.262        9		1.268		0		0				0

		189		4		-407.957        9		13.001		0		0				0

		190		4		-379.238        9		25.299		0		0				0

		191		4		-346.757        9		37.955		0		0				0

		192		4		-310.35         9		50.623		0		0				0

		193		4		-270.387        9		62.752		0		0				0

		194		58		-4513.56		40000		0		0				1

		195		58		-3536.03		40000		0		0				1

		196		58		-2395.63		40000		0		0				1

		197		58		-1208.26		40000		0		0				1

		198		58		0.000202763		40000		0		0				1

		199		58		0.000132275		40000		0		0				1

		200		58		1287.43		40000		0		0				1

		201		58		2570.3		40000		0		0				1

		202		58		3885.15		40000		0		0				1

		203		57		5311.1		40000		0		0				1

		204		4		280.367         9		59.893		0		0				0

		205		4		322.5		946.57		0		0				0

		206		4		360.755         9		32.661		0		0				0

		207		4		394.723		918.8		0		0				0

		208		4		424.693         9		5.338		0		0				0

		209		4		451.071         8		92.488		0		0				0

		210		4		-451.071        8		92.488		0		0				0

		211		4		-424.693        9		5.338		0		0				0

		212		4		-394.723		918.8		0		0				0

		213		4		-360.755        9		32.661		0		0				0

		214		4		-322.5		946.57		0		0				0

		215		4		-280.367        9		59.893		0		0				0

		216		57		-5311.1		40000		0		0				1

		217		58		-3885.15		40000		0		0				1

		218		58		-2570.3		40000		0		0				1

		219		58		-1287.43		40000		0		0				1

		220		58		0.000132275		40000		0		0				1

																				%%

																158				71.8

																62				28.2





70 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		62		0.000155228		40000		0		0				1

		2		62		349.401		40000		0		0				1

		3		62		688.35		40000		0		0				1

		4		62		1007.8		40000		0		0				1

		5		62		1301.08		40000		0		0				1

		6		62		1564.32		40000		0		0				1

		7		62		1796.14		40000		0		0				1

		8		62		1997.67		40000		0		0				1

		9		62		2171.11		40000		0		0				1

		10		62		2319.59		40000		0		0				1

		11		62		2446.52		40000		0		0				1

		12		62		-2446.52		40000		0		0				1

		13		62		-2319.59		40000		0		0				1

		14		62		-2171.11		40000		0		0				1

		15		62		-1997.67		40000		0		0				1

		16		62		-1796.14		40000		0		0				1

		17		62		-1564.32		40000		0		0				1

		18		62		-1301.08		40000		0		0				1

		19		62		-1007.8		40000		0		0				1

		20		62		-688.35		40000		0		0				1

		21		62		-349.401		40000		0		0				1

		22		62		0.000155228		40000		0		0				1

		23		61		0.000234536		40000		0		0				1

		24		61		484.469		40000		0		0				1

		25		61		954.532		40000		0		0				1

		26		61		1397.81		40000		0		0				1

		27		61		1805.19		40000		0		0				1

		28		61		2170.86		40000		0		0				1

		29		62		2491.51		40000		0		0				1

		30		62		2765.36		40000		0		0				1

		31		62		2992.07		40000		0		0				1

		32		62		3173.57		40000		0		0				1

		33		62		3314.93		40000		0		0				1

		34		62		-3314.93		40000		0		0				1

		35		62		-3173.57		40000		0		0				1

		36		62		-2992.07		40000		0		0				1

		37		62		-2765.36		40000		0		0				1

		38		62		-2491.51		40000		0		0				1

		39		61		-2170.86		40000		0		0				1

		40		61		-1805.19		40000		0		0				1

		41		61		-1397.81		40000		0		0				1

		42		61		-954.532		40000		0		0				1

		43		61		-484.469		40000		0		0				1

		44		61		0.000234536		40000		0		0				1

		45		61		0.000116572		40000		0		0				1

		46		61		586.783		40000		0		0				1

		47		61		1156.24		40000		0		0				1

		48		61		1694.5		40000		0		0				1

		49		61		2193.52		40000		0		0				1

		50		61		2652.77		40000		0		0				1

		51		61		3077.54		40000		0		0				1

		52		61		3474.84		40000		0		0				1

		53		61		3837.15		40000		0		0				1

		54		61		4106.59		40000		0		0				1

		55		61		4201.15		40000		0		0				1

		56		61		-4201.15		40000		0		0				1

		57		61		-4106.59		40000		0		0				1

		58		61		-3837.15		40000		0		0				1

		59		61		-3474.84		40000		0		0				1

		60		61		-3077.54		40000		0		0				1

		61		61		-2652.77		40000		0		0				1

		62		61		-2193.52		40000		0		0				1

		63		61		-1694.5		40000		0		0				1

		64		61		-1156.24		40000		0		0				1

		65		61		-586.783		40000		0		0				1

		66		61		0.000116572		40000		0		0				1

		67		60		0.000169786		40000		0		0				1

		68		60		672.305		40000		0		0				1

		69		60		1323.17		40000		0		0				1

		70		61		1935.29		40000		0		0				1

		71		61		2499.23		40000		0		0				1

		72		61		3016.53		40000		0		0				1

		73		61		3502.83		40000		0		0				1

		74		62		4001.6		40000		0		0				1

		75		62		4662.16		40000		0		0				1

		76		61		3711.99		40000		0		0				1

		77		4		273.85         9		61.772		0		0				0

		78		4		-273.85         9		61.772		0		0				0

		79		61		-3711.99		40000		0		0				1

		80		62		-4662.16		40000		0		0				1

		81		62		-4001.6		40000		0		0				1

		82		61		-3502.83		40000		0		0				1

		83		61		-3016.53		40000		0		0				1

		84		61		-2499.23		40000		0		0				1

		85		61		-1935.29		40000		0		0				1

		86		60		-1323.17		40000		0		0				1

		87		60		-672.305		40000		0		0				1

		88		60		0.000169786		40000		0		0				1

		89		60		0.000245418		40000		0		0				1

		90		60		748.753		40000		0		0				1

		91		60		1471.2		40000		0		0				1

		92		60		2143.84		40000		0		0				1

		93		60		2748.67		40000		0		0				1

		94		61		3273.45		40000		0		0				1

		95		61		3702.4		40000		0		0				1

		96		61		3987.41		40000		0		0				1

		97		62		4073.22		40000		0		0				1

		98		4		281.524         9		59.554		0		0				0

		99		4		298.648         9		54.363		0		0				0

		100		4		-298.648        9		54.363		0		0				0

		101		4		-281.524        9		59.554		0		0				0

		102		62		-4073.22		40000		0		0				1

		103		61		-3987.41		40000		0		0				1

		104		61		-3702.4		40000		0		0				1

		105		61		-3273.45		40000		0		0				1

		106		60		-2748.67		40000		0		0				1

		107		60		-2143.84		40000		0		0				1

		108		60		-1471.2		40000		0		0				1

		109		60		-748.753		40000		0		0				1

		110		60		0.000245418		40000		0		0				1

		111		59		0.000168293		40000		0		0				1

		112		59		887.482		40000		0		0				1

		113		59		1743.55		40000		0		0				1

		114		59		2533.58		40000		0		0				1

		115		59		3218.41		40000		0		0				1

		116		60		3765.72		40000		0		0				1

		117		60		4176.32		40000		0		0				1

		118		4		271.616         9		62.406		0		0				0

		119		4		296.555         9		55.016		0		0				0

		120		4		318.581         9		47.896		0		0				0

		121		4		337.996         9		41.147		0		0				0

		122		4		-337.996        9		41.147		0		0				0

		123		4		-318.581        9		47.896		0		0				0

		124		4		-296.555        9		55.016		0		0				0

		125		4		-271.616        9		62.406		0		0				0

		126		60		-4176.32		40000		0		0				1

		127		60		-3765.72		40000		0		0				1

		128		59		-3218.41		40000		0		0				1

		129		59		-2533.58		40000		0		0				1

		130		59		-1743.55		40000		0		0				1

		131		59		-887.482		40000		0		0				1

		132		59		0.000168293		40000		0		0				1

		133		59		-0.000152184		40000		0		0				1

		134		59		1010.09		40000		0		0				1

		135		59		1995.95		40000		0		0				1

		136		59		2928.91		40000		0		0				1

		137		59		3756.64		40000		0		0				1

		138		58		4378.86		40000		0		0				1

		139		4		265.773         9		64.036		0		0				0

		140		4		295.981         9		55.194		0		0				0

		141		4		323.079         9		46.372		0		0				0

		142		4		347.173         9		37.801		0		0				0

		143		4		368.467         9		29.641		0		0				0

		144		4		-368.467        9		29.641		0		0				0

		145		4		-347.173        9		37.801		0		0				0

		146		4		-323.079        9		46.372		0		0				0

		147		4		-295.981        9		55.194		0		0				0

		148		4		-265.773        9		64.036		0		0				0

		149		58		-4378.86		40000		0		0				1

		150		59		-3756.64		40000		0		0				1

		151		59		-2928.91		40000		0		0				1

		152		59		-1995.95		40000		0		0				1

		153		59		-1010.09		40000		0		0				1

		154		59		-0.000152184		40000		0		0				1

		155		58		0.000218983		40000		0		0				1

		156		58		1111.25		40000		0		0				1

		157		58		2211.86		40000		0		0				1

		158		58		3306.23		40000		0		0				1

		159		58		4410.99		40000		0		0				1

		160		57		5239.21		40000		0		0				1

		161		4		282.838         9		59.168		0		0				0

		162		4		315.109         9		49.055		0		0				0

		163		4		344.374         9		38.832		0		0				0

		164		4		370.343         9		28.895		0		0				0

		165		4		393.249         9		19.432		0		0				0

		166		4		-393.249        9		19.432		0		0				0

		167		4		-370.343        9		28.895		0		0				0

		168		4		-344.374        9		38.832		0		0				0

		169		4		-315.109        9		49.055		0		0				0

		170		4		-282.838        9		59.168		0		0				0

		171		57		-5239.21		40000		0		0				1

		172		58		-4410.99		40000		0		0				1

		173		58		-3306.23		40000		0		0				1

		174		58		-2211.86		40000		0		0				1

		175		58		-1111.25		40000		0		0				1

		176		58		0.000218983		40000		0		0				1

		177		58		0.000229263		40000		0		0				1

		178		58		1189.02		40000		0		0				1

		179		58		2372.25		40000		0		0				1

		180		58		3582.15		40000		0		0				1

		181		59		5030.02		40000		0		0				1

		182		56		5175.96		40000		0		0				1

		183		4		295.665         9		55.292		0		0				0

		184		4		330.856         9		43.681		0		0				0

		185		4		362.054         9		32.157		0		0				0

		186		4		389.865         9		20.872		0		0				0

		187		4		414.107         9		10.228		0		0				0

		188		4		-414.107        9		10.228		0		0				0

		189		4		-389.865        9		20.872		0		0				0

		190		4		-362.054        9		32.157		0		0				0

		191		4		-330.856        9		43.681		0		0				0

		192		4		-295.665        9		55.292		0		0				0

		193		56		-5175.96		40000		0		0				1

		194		59		-5030.02		40000		0		0				1

		195		58		-3582.15		40000		0		0				1

		196		58		-2372.25		40000		0		0				1

		197		58		-1189.02		40000		0		0				1

		198		58		0.000229263		40000		0		0				1

		199		57		0.00018622		40000		0		0				1

		200		57		1247.86		40000		0		0				1

		201		58		2479.54		40000		0		0				1

		202		58		3699.95		40000		0		0				1

		203		59		4933.8		40000		0		0				1

		204		4		268.905         9		63.167		0		0				0

		205		4		308.158         9		51.335		0		0				0

		206		4		344.941         9		38.624		0		0				0

		207		4		377.492         9		26.013		0		0				0

		208		4		406.188		913.79		0		0				0

		209		4		431.434         9		2.144		0		0				0

		210		4		-431.434        9		2.144		0		0				0

		211		4		-406.188		913.79		0		0				0

		212		4		-377.492        9		26.013		0		0				0

		213		4		-344.941        9		38.624		0		0				0

		214		4		-308.158        9		51.335		0		0				0

		215		4		-268.905        9		63.167		0		0				0

		216		59		-4933.8		40000		0		0				1

		217		58		-3699.95		40000		0		0				1

		218		58		-2479.54		40000		0		0				1

		219		57		-1247.86		40000		0		0				1

		220		57		0.00018622		40000		0		0				1

																				%%

																164				74.5

																56				25.5





80 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		61		0.000230997		40000		0		0				1

		2		61		336.855		40000		0		0				1

		3		61		663.991		40000		0		0				1

		4		61		973.104		40000		0		0				1

		5		61		1258.24		40000		0		0				1

		6		61		1516.06		40000		0		0				1

		7		61		1745.5		40000		0		0				1

		8		61		1947.2		40000		0		0				1

		9		61		2122.8		40000		0		0				1

		10		61		2274.49		40000		0		0				1

		11		61		2404.65		40000		0		0				1

		12		61		-2404.65		40000		0		0				1

		13		61		-2274.49		40000		0		0				1

		14		61		-2122.8		40000		0		0				1

		15		61		-1947.2		40000		0		0				1

		16		61		-1745.5		40000		0		0				1

		17		61		-1516.06		40000		0		0				1

		18		61		-1258.24		40000		0		0				1

		19		61		-973.104		40000		0		0				1

		20		61		-663.991		40000		0		0				1

		21		61		-336.855		40000		0		0				1

		22		61		0.000230997		40000		0		0				1

		23		61		0.00016472		40000		0		0				1

		24		61		471.087		40000		0		0				1

		25		61		927.741		40000		0		0				1

		26		61		1357.89		40000		0		0				1

		27		61		1753.49		40000		0		0				1

		28		61		2111.15		40000		0		0				1

		29		61		2431.67		40000		0		0				1

		30		61		2718.8		40000		0		0				1

		31		61		2978.86		40000		0		0				1

		32		61		3218.77		40000		0		0				1

		33		61		3446.14		40000		0		0				1

		34		61		-3446.14		40000		0		0				1

		35		61		-3218.77		40000		0		0				1

		36		61		-2978.86		40000		0		0				1

		37		61		-2718.8		40000		0		0				1

		38		61		-2431.67		40000		0		0				1

		39		61		-2111.15		40000		0		0				1

		40		61		-1753.49		40000		0		0				1

		41		61		-1357.89		40000		0		0				1

		42		61		-927.741		40000		0		0				1

		43		61		-471.087		40000		0		0				1

		44		61		0.00016472		40000		0		0				1

		45		60		0.000122476		40000		0		0				1

		46		60		575.395		40000		0		0				1

		47		60		1131.36		40000		0		0				1

		48		60		1650.94		40000		0		0				1

		49		60		2121.7		40000		0		0				1

		50		60		2536.94		40000		0		0				1

		51		61		2895.47		40000		0		0				1

		52		61		3199.89		40000		0		0				1

		53		61		3453.47		40000		0		0				1

		54		61		3657.78		40000		0		0				1

		55		62		3808.09		40000		0		0				1

		56		62		-3808.09		40000		0		0				1

		57		61		-3657.78		40000		0		0				1

		58		61		-3453.47		40000		0		0				1

		59		61		-3199.89		40000		0		0				1

		60		61		-2895.47		40000		0		0				1

		61		60		-2536.94		40000		0		0				1

		62		60		-2121.7		40000		0		0				1

		63		60		-1650.94		40000		0		0				1

		64		60		-1131.36		40000		0		0				1

		65		60		-575.395		40000		0		0				1

		66		60		0.000122476		40000		0		0				1

		67		60		0.000196866		40000		0		0				1

		68		60		666.751		40000		0		0				1

		69		60		1310.37		40000		0		0				1

		70		60		1909.1		40000		0		0				1

		71		60		2444.19		40000		0		0				1

		72		60		2902.35		40000		0		0				1

		73		60		3278.02		40000		0		0				1

		74		60		3574.81		40000		0		0				1

		75		60		3801.31		40000		0		0				1

		76		61		3968.69		40000		0		0				1

		77		61		4089.68		40000		0		0				1

		78		61		-4089.68		40000		0		0				1

		79		61		-3968.69		40000		0		0				1

		80		60		-3801.31		40000		0		0				1

		81		60		-3574.81		40000		0		0				1

		82		60		-3278.01		40000		0		0				1

		83		60		-2902.35		40000		0		0				1

		84		60		-2444.19		40000		0		0				1

		85		60		-1909.1		40000		0		0				1

		86		60		-1310.37		40000		0		0				1

		87		60		-666.751		40000		0		0				1

		88		60		0.000196866		40000		0		0				1

		89		60		9.23E-05		40000		0		0				1

		90		60		749.558		40000		0		0				1

		91		60		1475.12		40000		0		0				1

		92		60		2152.42		40000		0		0				1

		93		60		2756.66		40000		0		0				1

		94		60		3265.15		40000		0		0				1

		95		60		3666.1		40000		0		0				1

		96		60		3967.84		40000		0		0				1

		97		60		4194.22		40000		0		0				1

		98		4		268.995         9		63.142		0		0				0

		99		4		284.98         9		58.533		0		0				0

		100		4		-284.98         9		58.533		0		0				0

		101		4		-268.995        9		63.142		0		0				0

		102		60		-4194.22		40000		0		0				1

		103		60		-3967.84		40000		0		0				1

		104		60		-3666.1		40000		0		0				1

		105		60		-3265.15		40000		0		0				1

		106		60		-2756.66		40000		0		0				1

		107		60		-2152.42		40000		0		0				1

		108		60		-1475.12		40000		0		0				1

		109		60		-749.558		40000		0		0				1

		110		60		9.23E-05		40000		0		0				1

		111		59		0.000238213		40000		0		0				1

		112		59		890.819		40000		0		0				1

		113		59		1764.53		40000		0		0				1

		114		59		2605.77		40000		0		0				1

		115		59		3395.67		40000		0		0				1

		116		59		4084.9		40000		0		0				1

		117		58		4562.28		40000		0		0				1

		118		58		4790.88		40000		0		0				1

		119		4		284.157         9		58.778		0		0				0

		120		4		304.598         9		52.481		0		0				0

		121		4		322.814         9		46.462		0		0				0

		122		4		-322.814        9		46.462		0		0				0

		123		4		-304.598        9		52.481		0		0				0

		124		4		-284.157        9		58.778		0		0				0

		125		58		-4790.88		40000		0		0				1

		126		58		-4562.28		40000		0		0				1

		127		59		-4084.9		40000		0		0				1

		128		59		-3395.67		40000		0		0				1

		129		59		-2605.77		40000		0		0				1

		130		59		-1764.53		40000		0		0				1

		131		59		-890.819		40000		0		0				1

		132		59		0.000238213		40000		0		0				1

		133		59		0.00019261		40000		0		0				1

		134		59		997.74		40000		0		0				1

		135		59		1986.01		40000		0		0				1

		136		59		2970.49		40000		0		0				1

		137		59		3996.28		40000		0		0				1

		138		58		5185.75		40000		0		0				1

		139		56		6247.74		40000		0		0				1

		140		4		282.917         9		59.145		0		0				0

		141		4		308.705         9		51.158		0		0				0

		142		4		331.787         9		43.354		0		0				0

		143		4		352.175         9		35.934		0		0				0

		144		4		-352.175        9		35.934		0		0				0

		145		4		-331.787        9		43.354		0		0				0

		146		4		-308.705        9		51.158		0		0				0

		147		4		-282.917        9		59.145		0		0				0

		148		56		-6247.74		40000		0		0				1

		149		58		-5185.75		40000		0		0				1

		150		59		-3996.28		40000		0		0				1

		151		59		-2970.49		40000		0		0				1

		152		59		-1986.01		40000		0		0				1

		153		59		-997.74		40000		0		0				1

		154		59		0.00019261		40000		0		0				1

		155		58		0.000262118		40000		0		0				1

		156		58		1075.27		40000		0		0				1

		157		58		2136.58		40000		0		0				1

		158		59		3186.87		40000		0		0				1

		159		59		4278.47		40000		0		0				1

		160		61		5655.33		40000		0		0				1

		161		4		270.457         9		62.732		0		0				0

		162		4		300.587         9		53.755		0		0				0

		163		4		328.939         9		44.351		0		0				0

		164		4		354.254         9		35.149		0		0				0

		165		4		376.263         9		26.513		0		0				0

		166		4		-376.263        9		26.513		0		0				0

		167		4		-354.254        9		35.149		0		0				0

		168		4		-328.939        9		44.351		0		0				0

		169		4		-300.587        9		53.755		0		0				0

		170		4		-270.457        9		62.732		0		0				0

		171		61		-5655.33		40000		0		0				1

		172		59		-4278.47		40000		0		0				1

		173		59		-3186.87		40000		0		0				1

		174		58		-2136.58		40000		0		0				1

		175		58		-1075.27		40000		0		0				1

		176		58		0.000262118		40000		0		0				1

		177		58		0.000191718		40000		0		0				1

		178		58		1134.97		40000		0		0				1

		179		58		2240.8		40000		0		0				1

		180		58		3288.2		40000		0		0				1

		181		59		4214.66		40000		0		0				1

		182		61		4671.66		40000		0		0				1

		183		4		286.941         9		57.948		0		0				0

		184		4		318.772         9		47.831		0		0				0

		185		4		348.677         9		37.243		0		0				0

		186		4		373.648		927.57		0		0				0

		187		4		396.847         9		17.885		0		0				0

		188		4		-396.847        9		17.885		0		0				0

		189		4		-373.648		927.57		0		0				0

		190		4		-347.977        9		37.503		0		0				0

		191		4		-318.772        9		47.831		0		0				0

		192		4		-286.941        9		57.948		0		0				0

		193		61		-4671.66		40000		0		0				1

		194		59		-4214.66		40000		0		0				1

		195		58		-3288.2		40000		0		0				1

		196		58		-2240.8		40000		0		0				1

		197		58		-1134.97		40000		0		0				1

		198		58		0.000191718		40000		0		0				1

		199		57		-0.000142712		40000		0		0				1

		200		57		1187.2		40000		0		0				1

		201		57		2328.59		40000		0		0				1

		202		58		3366.31		40000		0		0				1

		203		59		4199.26		40000		0		0				1

		204		60		4695.07		40000		0		0				1

		205		4		299.197         9		54.191		0		0				0

		206		4		334.295         9		42.469		0		0				0

		207		4		365.601         9		30.772		0		0				0

		208		4		391.967         9		19.979		0		0				0

		209		4		414.9          9		9.867		0		0				0

		210		4		-414.9         9		9.867		0		0				0

		211		4		-391.967        9		19.979		0		0				0

		212		4		-365.601        9		30.772		0		0				0

		213		4		-334.295        9		42.469		0		0				0

		214		4		-299.197        9		54.191		0		0				0

		215		60		-4695.07		40000		0		0				1

		216		59		-4199.26		40000		0		0				1

		217		58		-3366.31		40000		0		0				1

		218		57		-2328.59		40000		0		0				1

		219		57		-1187.2		40000		0		0				1

		220		57		-0.000142712		40000		0		0				1

																				%%

																172				78.2

																48				21.8





90 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		60		0.000124716		40000		0		0				1

		2		60		328.996		40000		0		0				1

		3		60		647.937		40000		0		0				1

		4		60		948.257		40000		0		0				1

		5		60		1223.91		40000		0		0				1

		6		61		1471.71		40000		0		0				1

		7		61		1691.02		40000		0		0				1

		8		61		1883.13		40000		0		0				1

		9		61		2050.45		40000		0		0				1

		10		61		2195.92		40000		0		0				1

		11		61		2322.53		40000		0		0				1

		12		61		-2322.53		40000		0		0				1

		13		61		-2195.92		40000		0		0				1

		14		61		-2050.45		40000		0		0				1

		15		61		-1883.13		40000		0		0				1

		16		61		-1691.02		40000		0		0				1

		17		61		-1471.71		40000		0		0				1

		18		60		-1223.91		40000		0		0				1

		19		60		-948.257		40000		0		0				1

		20		60		-647.937		40000		0		0				1

		21		60		-328.996		40000		0		0				1

		22		60		0.000124716		40000		0		0				1

		23		60		0.000261359		40000		0		0				1

		24		60		467.794		40000		0		0				1

		25		60		920.286		40000		0		0				1

		26		60		1344.02		40000		0		0				1

		27		60		1728.85		40000		0		0				1

		28		60		2068.86		40000		0		0				1

		29		60		2362.33		40000		0		0				1

		30		60		2611.16		40000		0		0				1

		31		60		2819.67		40000		0		0				1

		32		60		2993.29		40000		0		0				1

		33		60		3137.55		40000		0		0				1

		34		60		-3137.55		40000		0		0				1

		35		60		-2993.29		40000		0		0				1

		36		60		-2819.67		40000		0		0				1

		37		60		-2611.16		40000		0		0				1

		38		60		-2362.33		40000		0		0				1

		39		60		-2068.86		40000		0		0				1

		40		60		-1728.85		40000		0		0				1

		41		60		-1344.02		40000		0		0				1

		42		60		-920.286		40000		0		0				1

		43		60		-467.794		40000		0		0				1

		44		60		0.000261359		40000		0		0				1

		45		60		0.000209604		40000		0		0				1

		46		60		578.766		40000		0		0				1

		47		60		1139.75		40000		0		0				1

		48		60		1666.4		40000		0		0				1

		49		60		2144.52		40000		0		0				1

		50		60		2563.21		40000		0		0				1

		51		60		2916.29		40000		0		0				1

		52		60		3203.58		40000		0		0				1

		53		60		3430.85		40000		0		0				1

		54		60		3607.88		40000		0		0				1

		55		60		3745.35		40000		0		0				1

		56		60		-3745.35		40000		0		0				1

		57		60		-3607.88		40000		0		0				1

		58		60		-3430.85		40000		0		0				1

		59		60		-3203.58		40000		0		0				1

		60		60		-2916.29		40000		0		0				1

		61		60		-2563.21		40000		0		0				1

		62		60		-2144.52		40000		0		0				1

		63		60		-1666.4		40000		0		0				1

		64		60		-1139.75		40000		0		0				1

		65		60		-578.766		40000		0		0				1

		66		60		0.000209604		40000		0		0				1

		67		60		0.000256562		40000		0		0				1

		68		60		672.246		40000		0		0				1

		69		60		1327.52		40000		0		0				1

		70		60		1949.78		40000		0		0				1

		71		59		2523.75		40000		0		0				1

		72		59		3033.49		40000		0		0				1

		73		59		3462.03		40000		0		0				1

		74		59		3796.43		40000		0		0				1

		75		59		4036.7		40000		0		0				1

		76		59		4202.4		40000		0		0				1

		77		59		4320.36		40000		0		0				1

		78		59		-4320.36		40000		0		0				1

		79		59		-4202.4		40000		0		0				1

		80		59		-4036.7		40000		0		0				1

		81		59		-3796.43		40000		0		0				1

		82		59		-3462.03		40000		0		0				1

		83		59		-3033.49		40000		0		0				1

		84		59		-2523.75		40000		0		0				1

		85		60		-1949.78		40000		0		0				1

		86		60		-1327.52		40000		0		0				1

		87		60		-672.246		40000		0		0				1

		88		60		0.000256562		40000		0		0				1

		89		59		0.000212081		40000		0		0				1

		90		59		750.551		40000		0		0				1

		91		59		1486.88		40000		0		0				1

		92		59		2197.31		40000		0		0				1

		93		59		2872.74		40000		0		0				1

		94		59		3501.19		40000		0		0				1

		95		59		4053.6		40000		0		0				1

		96		58		4468.03		40000		0		0				1

		97		58		4688.42		40000		0		0				1

		98		58		4767.73		40000		0		0				1

		99		4		274.227         9		61.665		0		0				0

		100		4		-274.227        9		61.665		0		0				0

		101		58		-4767.73		40000		0		0				1

		102		58		-4688.42		40000		0		0				1

		103		58		-4468.03		40000		0		0				1

		104		59		-4053.6		40000		0		0				1

		105		59		-3501.19		40000		0		0				1

		106		59		-2872.74		40000		0		0				1

		107		59		-2197.31		40000		0		0				1

		108		59		-1486.88		40000		0		0				1

		109		59		-750.551		40000		0		0				1

		110		59		0.000212081		40000		0		0				1

		111		59		0.000188142		40000		0		0				1

		112		59		868.473		40000		0		0				1

		113		59		1724.96		40000		0		0				1

		114		59		2566.67		40000		0		0				1

		115		59		3411.14		40000		0		0				1

		116		59		4319.62		40000		0		0				1

		117		59		5450.73		40000		0		0				1

		118		4		266.923         9		63.718		0		0				0

		119		4		271.378         9		62.473		0		0				0

		120		4		290.752         9		56.799		0		0				0

		121		4		308.729		951.15		0		0				0

		122		4		-308.729		951.15		0		0				0

		123		4		-290.752        9		56.799		0		0				0

		124		4		-271.378        9		62.473		0		0				0

		125		4		-266.923        9		63.718		0		0				0

		126		59		-5450.73		40000		0		0				1

		127		59		-4319.62		40000		0		0				1

		128		59		-3411.14		40000		0		0				1

		129		59		-2566.67		40000		0		0				1

		130		59		-1724.96		40000		0		0				1

		131		59		-868.473		40000		0		0				1

		132		59		0.000188142		40000		0		0				1

		133		58		0.000177294		40000		0		0				1

		134		59		950.672		40000		0		0				1

		135		59		1880.78		40000		0		0				1

		136		59		2775.83		40000		0		0				1

		137		60		3631.36		40000		0		0				1

		138		61		4437.32		40000		0		0				1

		139		63		4966.79		40000		0		0				1

		140		4		272.353         9		62.197		0		0				0

		141		4		296.824         9		54.932		0		0				0

		142		4		317.377		948.3		0		0				0

		143		4		336.994         9		41.507		0		0				0

		144		4		-336.994        9		41.507		0		0				0

		145		4		-317.377		948.3		0		0				0

		146		4		-296.824        9		54.932		0		0				0

		147		4		-272.353        9		62.197		0		0				0

		148		63		-4966.79		40000		0		0				1

		149		61		-4437.32		40000		0		0				1

		150		60		-3631.36		40000		0		0				1

		151		59		-2775.83		40000		0		0				1

		152		59		-1880.78		40000		0		0				1

		153		59		-950.672		40000		0		0				1

		154		58		0.000177294		40000		0		0				1

		155		58		0.000329879		40000		0		0				1

		156		58		1016.2		40000		0		0				1

		157		58		1997.22		40000		0		0				1

		158		59		2906.35		40000		0		0				1

		159		59		3692.28		40000		0		0				1

		160		60		4250.74		40000		0		0				1

		161		62		4437.01		40000		0		0				1

		162		4		292.888         9		56.147		0		0				0

		163		4		319.64         9		47.539		0		0				0

		164		4		342.031         9		39.689		0		0				0

		165		4		363.124         9		31.741		0		0				0

		166		4		-363.124        9		31.741		0		0				0

		167		4		-342.031        9		39.689		0		0				0

		168		4		-319.64         9		47.539		0		0				0

		169		4		-292.888        9		56.147		0		0				0

		170		62		-4437.01		40000		0		0				1

		171		60		-4250.74		40000		0		0				1

		172		59		-3692.28		40000		0		0				1

		173		59		-2906.35		40000		0		0				1

		174		58		-1997.22		40000		0		0				1

		175		58		-1016.2		40000		0		0				1

		176		58		0.000329879		40000		0		0				1

		177		57		0.000319197		40000		0		0				1

		178		57		1076.81		40000		0		0				1

		179		58		2106.12		40000		0		0				1

		180		58		3034.93		40000		0		0				1

		181		59		3801.35		40000		0		0				1

		182		59		4351.85		40000		0		0				1

		183		4		276.281         9		61.077		0		0				0

		184		4		308.623         9		51.185		0		0				0

		185		4		337.34         9		41.383		0		0				0

		186		4		362.762         9		31.882		0		0				0

		187		4		385.203         9		22.832		0		0				0

		188		4		-385.203        9		22.832		0		0				0

		189		4		-362.762        9		31.882		0		0				0

		190		4		-337.34         9		41.383		0		0				0

		191		4		-308.623        9		51.185		0		0				0

		192		4		-276.281        9		61.077		0		0				0

		193		59		-4351.85		40000		0		0				1

		194		59		-3801.35		40000		0		0				1

		195		58		-3034.93		40000		0		0				1

		196		58		-2106.12		40000		0		0				1

		197		57		-1076.81		40000		0		0				1

		198		57		0.000319197		40000		0		0				1

		199		57		0.000168184		40000		0		0				1

		200		57		1137.69		40000		0		0				1

		201		57		2220.67		40000		0		0				1

		202		57		3187.83		40000		0		0				1

		203		58		3982.29		40000		0		0				1

		204		58		4587.03		40000		0		0				1

		205		4		288.3		957.54		0		0				0

		206		4		322.345         9		46.622		0		0				0

		207		4		352.548         9		35.794		0		0				0

		208		4		379.246         9		25.296		0		0				0

		209		4		402.793         9		15.291		0		0				0

		210		4		-402.793        9		15.291		0		0				0

		211		4		-379.246        9		25.296		0		0				0

		212		4		-352.548        9		35.794		0		0				0

		213		4		-322.345        9		46.622		0		0				0

		214		4		-288.3		957.54		0		0				0

		215		58		-4587.03		40000		0		0				1

		216		58		-3982.29		40000		0		0				1

		217		57		-3187.83		40000		0		0				1

		218		57		-2220.67		40000		0		0				1

		219		57		-1137.69		40000		0		0				1

		220		57		0.000168184		40000		0		0				1

																				%%

																174				79.1

																46				20.9





100 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		60		0.000200255		40000		0		0				1

		2		60		330.115		40000		0		0				1

		3		60		650.001		40000		0		0				1

		4		60		950.785		40000		0		0				1

		5		60		1225.94		40000		0		0				1

		6		60		1471.69		40000		0		0				1

		7		60		1686.91		40000		0		0				1

		8		60		1872.55		40000		0		0				1

		9		60		2030.97		40000		0		0				1

		10		60		2165.27		40000		0		0				1

		11		60		2278.79		40000		0		0				1

		12		60		-2278.79		40000		0		0				1

		13		60		-2165.27		40000		0		0				1

		14		60		-2030.97		40000		0		0				1

		15		60		-1872.55		40000		0		0				1

		16		60		-1686.91		40000		0		0				1

		17		60		-1471.69		40000		0		0				1

		18		60		-1225.94		40000		0		0				1

		19		60		-950.785		40000		0		0				1

		20		60		-650.001		40000		0		0				1

		21		60		-330.115		40000		0		0				1

		22		60		0.000200255		40000		0		0				1

		23		60		0.000372984		40000		0		0				1

		24		60		473.485		40000		0		0				1

		25		60		933.884		40000		0		0				1

		26		60		1369.41		40000		0		0				1

		27		60		1770.53		40000		0		0				1

		28		60		2130.35		40000		0		0				1

		29		60		2444.81		40000		0		0				1

		30		60		2712.65		40000		0		0				1

		31		59		2935.28		40000		0		0				1

		32		59		3116.42		40000		0		0				1

		33		59		3261.44		40000		0		0				1

		34		59		-3261.44		40000		0		0				1

		35		59		-3116.42		40000		0		0				1

		36		59		-2935.28		40000		0		0				1

		37		60		-2712.65		40000		0		0				1

		38		60		-2444.81		40000		0		0				1

		39		60		-2130.35		40000		0		0				1

		40		60		-1770.53		40000		0		0				1

		41		60		-1369.41		40000		0		0				1

		42		60		-933.884		40000		0		0				1

		43		60		-473.485		40000		0		0				1

		44		60		0.000372984		40000		0		0				1

		45		60		0.000156044		40000		0		0				1

		46		60		579.412		40000		0		0				1

		47		60		1146.13		40000		0		0				1

		48		60		1689.11		40000		0		0				1

		49		60		2199.69		40000		0		0				1

		50		60		2671.07		40000		0		0				1

		51		59		3096.64		40000		0		0				1

		52		59		3468.58		40000		0		0				1

		53		59		3778.25		40000		0		0				1

		54		59		4017.89		40000		0		0				1

		55		58		4187.48		40000		0		0				1

		56		58		-4187.48		40000		0		0				1

		57		59		-4017.89		40000		0		0				1

		58		59		-3778.25		40000		0		0				1

		59		59		-3468.58		40000		0		0				1

		60		59		-3096.64		40000		0		0				1

		61		60		-2671.07		40000		0		0				1

		62		60		-2199.69		40000		0		0				1

		63		60		-1689.11		40000		0		0				1

		64		60		-1146.13		40000		0		0				1

		65		60		-579.412		40000		0		0				1

		66		60		0.000156044		40000		0		0				1

		67		59		0.000211949		40000		0		0				1

		68		60		660.289		40000		0		0				1

		69		60		1308.64		40000		0		0				1

		70		60		1936.27		40000		0		0				1

		71		60		2539.61		40000		0		0				1

		72		60		3120.58		40000		0		0				1

		73		60		3683.72		40000		0		0				1

		74		59		4228.47		40000		0		0				1

		75		59		4735.76		40000		0		0				1

		76		58		5146.63		40000		0		0				1

		77		57		5344.43		40000		0		0				1

		78		57		-5344.43		40000		0		0				1

		79		58		-5146.63		40000		0		0				1

		80		59		-4735.76		40000		0		0				1

		81		59		-4228.47		40000		0		0				1

		82		60		-3683.72		40000		0		0				1

		83		60		-3120.58		40000		0		0				1

		84		60		-2539.61		40000		0		0				1

		85		60		-1936.27		40000		0		0				1

		86		60		-1308.64		40000		0		0				1

		87		60		-660.289		40000		0		0				1

		88		59		0.000211949		40000		0		0				1

		89		59		0.000191269		40000		0		0				1

		90		59		723.301		40000		0		0				1

		91		59		1433.51		40000		0		0				1

		92		60		2122.12		40000		0		0				1

		93		60		2789.42		40000		0		0				1

		94		60		3448.84		40000		0		0				1

		95		60		4132.69		40000		0		0				1

		96		60		4902.23		40000		0		0				1

		97		59		5835.1		40000		0		0				1

		98		4		266.443         9		63.851		0		0				0

		99		57		4555.6		40000		0		0				1

		100		57		-4555.6		40000		0		0				1

		101		4		-266.443        9		63.851		0		0				0

		102		59		-5835.1		40000		0		0				1

		103		60		-4902.23		40000		0		0				1

		104		60		-4132.69		40000		0		0				1

		105		60		-3448.84		40000		0		0				1

		106		60		-2789.42		40000		0		0				1

		107		60		-2122.12		40000		0		0				1

		108		59		-1433.51		40000		0		0				1

		109		59		-723.301		40000		0		0				1

		110		59		0.000191269		40000		0		0				1

		111		59		0.000138938		40000		0		0				1

		112		59		817.78		40000		0		0				1

		113		59		1613.15		40000		0		0				1

		114		59		2367.35		40000		0		0				1

		115		60		3066.09		40000		0		0				1

		116		60		3694.86		40000		0		0				1

		117		61		4217.3		40000		0		0				1

		118		63		4469.02		40000		0		0				1

		119		4		266.721         9		63.774		0		0				0

		120		4		281.016         9		59.703		0		0				0

		121		4		297.632         9		54.681		0		0				0

		122		4		-297.632        9		54.681		0		0				0

		123		4		-281.016        9		59.703		0		0				0

		124		4		-266.721        9		63.774		0		0				0

		125		63		-4469.02		40000		0		0				1

		126		61		-4217.3		40000		0		0				1

		127		60		-3694.86		40000		0		0				1

		128		60		-3066.09		40000		0		0				1

		129		59		-2367.35		40000		0		0				1

		130		59		-1613.15		40000		0		0				1

		131		59		-817.78		40000		0		0				1

		132		59		0.000138938		40000		0		0				1

		133		58		0.000131573		40000		0		0				1

		134		58		893.94		40000		0		0				1

		135		58		1753.33		40000		0		0				1

		136		59		2543.94		40000		0		0				1

		137		59		3230.19		40000		0		0				1

		138		60		3771.12		40000		0		0				1

		139		61		4117.5		40000		0		0				1

		140		4		264.405         9		64.412		0		0				0

		141		4		288.387         9		57.514		0		0				0

		142		4		309.502         9		50.899		0		0				0

		143		4		328.191         9		44.611		0		0				0

		144		4		-328.191        9		44.611		0		0				0

		145		4		-309.502        9		50.899		0		0				0

		146		4		-288.387        9		57.514		0		0				0

		147		4		-264.405        9		64.412		0		0				0

		148		61		-4117.5		40000		0		0				1

		149		60		-3771.12		40000		0		0				1

		150		59		-3230.19		40000		0		0				1

		151		59		-2543.94		40000		0		0				1

		152		58		-1753.33		40000		0		0				1

		153		58		-893.94		40000		0		0				1

		154		58		0.000131573		40000		0		0				1

		155		58		0.000264189		40000		0		0				1

		156		58		965.92		40000		0		0				1

		157		58		1887.99		40000		0		0				1

		158		58		2721.36		40000		0		0				1

		159		58		3424.32		40000		0		0				1

		160		59		3968.36		40000		0		0				1

		161		60		4356.37		40000		0		0				1

		162		4		282.176         9		59.363		0		0				0

		163		4		308.36		951.27		0		0				0

		164		4		331.454         9		43.472		0		0				0

		165		4		351.808         9		36.072		0		0				0

		166		4		-351.808        9		36.072		0		0				0

		167		4		-331.454        9		43.472		0		0				0

		168		4		-308.36		951.27		0		0				0

		169		4		-282.176        9		59.363		0		0				0

		170		60		-4356.37		40000		0		0				1

		171		59		-3968.36		40000		0		0				1

		172		58		-3424.32		40000		0		0				1

		173		58		-2721.36		40000		0		0				1

		174		58		-1887.99		40000		0		0				1

		175		58		-965.92		40000		0		0				1

		176		58		0.000264189		40000		0		0				1

		177		57		0.000287172		40000		0		0				1

		178		57		1037.78		40000		0		0				1

		179		57		2027.24		40000		0		0				1

		180		57		2918.66		40000		0		0				1

		181		58		3666.78		40000		0		0				1

		182		58		4253.05		40000		0		0				1

		183		4		266.341         9		63.879		0		0				0

		184		4		297.775         9		54.636		0		0				0

		185		4		325.613         9		45.503		0		0				0

		186		4		350.189         9		36.679		0		0				0

		187		4		371.849         9		28.293		0		0				0

		188		4		-371.849        9		28.293		0		0				0

		189		4		-350.189        9		36.679		0		0				0

		190		4		-325.613        9		45.503		0		0				0

		191		4		-297.775        9		54.636		0		0				0

		192		4		-266.341        9		63.879		0		0				0

		193		58		-4253.05		40000		0		0				1

		194		58		-3666.78		40000		0		0				1

		195		57		-2918.66		40000		0		0				1

		196		57		-2027.24		40000		0		0				1

		197		57		-1037.78		40000		0		0				1

		198		57		0.000287172		40000		0		0				1

		199		56		0.000227842		40000		0		0				1

		200		56		1109.22		40000		0		0				1

		201		56		2171.28		40000		0		0				1

		202		57		3135.25		40000		0		0				1

		203		57		3946.91		40000		0		0				1

		204		57		4577.29		40000		0		0				1

		205		4		278.465         9		60.446		0		0				0

		206		4		311.41         9		50.276		0		0				0

		207		4		340.635         9		40.196		0		0				0

		208		4		366.454         9		30.436		0		0				0

		209		4		389.213         9		21.148		0		0				0

		210		4		-389.213        9		21.148		0		0				0

		211		4		-366.454        9		30.436		0		0				0

		212		4		-340.635        9		40.196		0		0				0

		213		4		-311.41         9		50.276		0		0				0

		214		4		-278.465        9		60.446		0		0				0

		215		57		-4577.29		40000		0		0				1

		216		57		-3946.91		40000		0		0				1

		217		57		-3135.25		40000		0		0				1

		218		56		-2171.28		40000		0		0				1

		219		56		-1109.22		40000		0		0				1

		220		56		0.000227842		40000		0		0				1

																				%%

																176				80.0

																44				20.0





110 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		60		0.000233043		40000		0		0				1

		2		60		334.312		40000		0		0				1

		3		60		659.47		40000		0		0				1

		4		60		967.425		40000		0		0				1

		5		60		1252.03		40000		0		0				1

		6		60		1509.37		40000		0		0				1

		7		60		1737.7		40000		0		0				1

		8		60		1937.08		40000		0		0				1

		9		60		2108.89		40000		0		0				1

		10		60		2255.43		40000		0		0				1

		11		60		2379.44		40000		0		0				1

		12		60		-2379.44		40000		0		0				1

		13		60		-2255.43		40000		0		0				1

		14		60		-2108.89		40000		0		0				1

		15		60		-1937.08		40000		0		0				1

		16		60		-1737.7		40000		0		0				1

		17		60		-1509.37		40000		0		0				1

		18		60		-1252.03		40000		0		0				1

		19		60		-967.425		40000		0		0				1

		20		60		-659.47		40000		0		0				1

		21		60		-334.312		40000		0		0				1

		22		60		0.000233043		40000		0		0				1

		23		60		0.000166556		40000		0		0				1

		24		60		466.662		40000		0		0				1

		25		60		922.867		40000		0		0				1

		26		60		1359.64		40000		0		0				1

		27		60		1770.42		40000		0		0				1

		28		60		2151.14		40000		0		0				1

		29		60		2499.71		40000		0		0				1

		30		60		2815.16		40000		0		0				1

		31		60		3096.97		40000		0		0				1

		32		59		3344.75		40000		0		0				1

		33		59		3558.3		40000		0		0				1

		34		59		-3558.3		40000		0		0				1

		35		59		-3344.75		40000		0		0				1

		36		60		-3096.97		40000		0		0				1

		37		60		-2815.16		40000		0		0				1

		38		60		-2499.71		40000		0		0				1

		39		60		-2151.14		40000		0		0				1

		40		60		-1770.42		40000		0		0				1

		41		60		-1359.64		40000		0		0				1

		42		60		-922.867		40000		0		0				1

		43		60		-466.662		40000		0		0				1

		44		60		0.000166556		40000		0		0				1

		45		60		0.000224996		40000		0		0				1

		46		60		555.426		40000		0		0				1

		47		60		1099.07		40000		0		0				1

		48		60		1621.53		40000		0		0				1

		49		60		2117.45		40000		0		0				1

		50		60		2585.83		40000		0		0				1

		51		60		3030.3		40000		0		0				1

		52		60		3456.71		40000		0		0				1

		53		60		3871.67		40000		0		0				1

		54		60		4279.67		40000		0		0				1

		55		60		4679.43		40000		0		0				1

		56		60		-4679.43		40000		0		0				1

		57		60		-4279.75		40000		0		0				1

		58		60		-3871.67		40000		0		0				1

		59		60		-3456.71		40000		0		0				1

		60		60		-3030.3		40000		0		0				1

		61		60		-2585.83		40000		0		0				1

		62		60		-2117.45		40000		0		0				1

		63		60		-1621.53		40000		0		0				1

		64		60		-1099.07		40000		0		0				1

		65		60		-555.426		40000		0		0				1

		66		60		0.000224996		40000		0		0				1

		67		59		0.000239		40000		0		0				1

		68		60		621.45		40000		0		0				1

		69		60		1227.86		40000		0		0				1

		70		60		1807.17		40000		0		0				1

		71		60		2352.39		40000		0		0				1

		72		60		2862.7		40000		0		0				1

		73		61		3343.95		40000		0		0				1

		74		61		3808.51		40000		0		0				1

		75		61		4277.79		40000		0		0				1

		76		62		4791.21		40000		0		0				1

		77		62		5419.84		40000		0		0				1

		78		62		-5420.03		40000		0		0				1

		79		62		-4791.21		40000		0		0				1

		80		61		-4277.79		40000		0		0				1

		81		61		-3808.51		40000		0		0				1

		82		61		-3343.95		40000		0		0				1

		83		60		-2862.7		40000		0		0				1

		84		60		-2352.43		40000		0		0				1

		85		60		-1807.17		40000		0		0				1

		86		60		-1227.86		40000		0		0				1

		87		60		-621.45		40000		0		0				1

		88		59		0.000239		40000		0		0				1

		89		59		0.000171653		40000		0		0				1

		90		59		675.333		40000		0		0				1

		91		59		1330.83		40000		0		0				1

		92		60		1949.54		40000		0		0				1

		93		60		2519.2		40000		0		0				1

		94		60		3031.99		40000		0		0				1

		95		61		3483.19		40000		0		0				1

		96		62		3864.26		40000		0		0				1

		97		63		4152.87		40000		0		0				1

		98		64		4273.62		40000		0		0				1

		99		67		3918.87		40000		0		0				1

		100		67		-3918.87		40000		0		0				1

		101		64		-4273.62		40000		0		0				1

		102		63		-4152.87		40000		0		0				1

		103		62		-3864.26		40000		0		0				1

		104		61		-3483.19		40000		0		0				1

		105		60		-3031.99		40000		0		0				1

		106		60		-2519.2		40000		0		0				1

		107		60		-1949.54		40000		0		0				1

		108		59		-1330.83		40000		0		0				1

		109		59		-675.333		40000		0		0				1

		110		59		0.000171653		40000		0		0				1

		111		59		10007.3         3		9998.7		0		0				0

		112		59		767.623		40000		0		0				1

		113		59		1505.02		40000		0		0				1

		114		59		2184.13		40000		0		0				1

		115		59		2780.49		40000		0		0				1

		116		60		3274.9		40000		0		0				1

		117		60		3651.29		40000		0		0				1

		118		61		3899.39		40000		0		0				1

		119		62		4022.56		40000		0		0				1

		120		4		273.277         9		61.935		0		0				0

		121		4		289.54         9		57.166		0		0				0

		122		4		-289.54         9		57.166		0		0				0

		123		4		-273.277        9		61.935		0		0				0

		124		62		-4022.56		40000		0		0				1

		125		61		-3899.39		40000		0		0				1

		126		60		-3651.29		40000		0		0				1

		127		60		-3274.9		40000		0		0				1

		128		59		-2780.49		40000		0		0				1

		129		59		-2184.13		40000		0		0				1

		130		59		-1505.02		40000		0		0				1

		131		59		-767.623		40000		0		0				1

		132		59		10007.3         3		9998.7		0		0				0

		133		58		0.000119293		40000		0		0				1

		134		58		853.352		40000		0		0				1

		135		58		1669.03		40000		0		0				1

		136		58		2410.72		40000		0		0				1

		137		59		3048.14		40000		0		0				1

		138		59		3562.63		40000		0		0				1

		139		59		3953.04		40000		0		0				1

		140		60		4235.65		40000		0		0				1

		141		4		277.905         9		60.608		0		0				0

		142		4		298.646         9		54.364		0		0				0

		143		4		316.86         9		48.472		0		0				0

		144		4		-316.86         9		48.472		0		0				0

		145		4		-298.646        9		54.364		0		0				0

		146		4		-277.905        9		60.608		0		0				0

		147		60		-4235.65		40000		0		0				1

		148		59		-3953.04		40000		0		0				1

		149		59		-3562.63		40000		0		0				1

		150		59		-3048.14		40000		0		0				1

		151		58		-2410.72		40000		0		0				1

		152		58		-1669.03		40000		0		0				1

		153		58		-853.352		40000		0		0				1

		154		58		0.000119293		40000		0		0				1

		155		57		0.000228484		40000		0		0				1

		156		57		936.982		40000		0		0				1

		157		57		1833.47		40000		0		0				1

		158		57		2649.34		40000		0		0				1

		159		58		3348.23		40000		0		0				1

		160		58		3909.28		40000		0		0				1

		161		58		4340.08		40000		0		0				1

		162		4		272.405         9		62.183		0		0				0

		163		4		297.804         9		54.627		0		0				0

		164		4		320.194         9		47.352		0		0				0

		165		4		339.902         9		40.461		0		0				0

		166		4		-339.902        9		40.461		0		0				0

		167		4		-320.194        9		47.352		0		0				0

		168		4		-297.804        9		54.627		0		0				0

		169		4		-272.405        9		62.183		0		0				0

		170		58		-4340.08		40000		0		0				1

		171		58		-3909.28		40000		0		0				1

		172		58		-3348.23		40000		0		0				1

		173		57		-2649.34		40000		0		0				1

		174		57		-1833.47		40000		0		0				1

		175		57		-936.982		40000		0		0				1

		176		57		0.000228484		40000		0		0				1

		177		57		0.000243491		40000		0		0				1

		178		57		1017.97		40000		0		0				1

		179		57		1997.51		40000		0		0				1

		180		57		2899.47		40000		0		0				1

		181		57		3679.6		40000		0		0				1

		182		57		4295.15		40000		0		0				1

		183		57		4752.17		40000		0		0				1

		184		4		287.945         9		57.647		0		0				0

		185		4		314.881         9		49.131		0		0				0

		186		4		338.671         9		40.905		0		0				0

		187		4		359.635         9		33.093		0		0				0

		188		4		-359.635        9		33.093		0		0				0

		189		4		-338.671        9		40.905		0		0				0

		190		4		-314.881        9		49.131		0		0				0

		191		4		-287.945        9		57.647		0		0				0

		192		57		-4752.17		40000		0		0				1

		193		57		-4295.15		40000		0		0				1

		194		57		-3679.6		40000		0		0				1

		195		57		-2899.47		40000		0		0				1

		196		57		-1997.51		40000		0		0				1

		197		57		-1017.97		40000		0		0				1

		198		57		0.000243491		40000		0		0				1

		199		56		9.11E-05		40000		0		0				1

		200		56		1094.28		40000		0		0				1

		201		56		2155.78		40000		0		0				1

		202		56		3154.68		40000		0		0				1

		203		56		4044.74		40000		0		0				1

		204		56		4718.62		40000		0		0				1

		205		4		269.592         9		62.975		0		0				0

		206		4		301.449         9		53.482		0		0				0

		207		4		329.758         9		44.066		0		0				0

		208		4		354.776         9		34.951		0		0				0

		209		4		376.825         9		26.285		0		0				0

		210		4		-376.825        9		26.285		0		0				0

		211		4		-354.776        9		34.951		0		0				0

		212		4		-329.758        9		44.066		0		0				0

		213		4		-301.449        9		53.482		0		0				0

		214		4		-269.592        9		62.975		0		0				0

		215		56		-4718.62		40000		0		0				1

		216		56		-4044.74		40000		0		0				1

		217		56		-3154.68		40000		0		0				1

		218		56		-2155.78		40000		0		0				1

		219		56		-1094.28		40000		0		0				1

		220		56		9.11E-05		40000		0		0				1

																				%%

																182				82.7

																38				17.3





120 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------------		-----------------		----------		-----------------		----------------

		1		60		14914.6		40000		0		0				1

		2		60		325.29		40000		0		0				1

		3		60		642.553		40000		0		0				1

		4		60		944.823		40000		0		0				1

		5		60		1226.91		40000		0		0				1

		6		60		1485.64		40000		0		0				1

		7		60		1719.66		40000		0		0				1

		8		60		1929.04		40000		0		0				1

		9		60		2114.86		40000		0		0				1

		10		60		2278.63		40000		0		0				1

		11		60		2422.13		40000		0		0				1

		12		60		-2422.13		40000		0		0				1

		13		60		-2278.63		40000		0		0				1

		14		60		-2114.9		40000		0		0				1

		15		60		-1929.04		40000		0		0				1

		16		60		-1719.66		40000		0		0				1

		17		60		-1485.64		40000		0		0				1

		18		60		-1226.91		40000		0		0				1

		19		60		-944.823		40000		0		0				1

		20		60		-642.553		40000		0		0				1

		21		60		-325.301		40000		0		0				1

		22		60		14914.6		40000		0		0				1

		23		60         0.		264525		40000		0		0				1

		24		60		438.358		40000		0		0				1

		25		60		865.99		40000		0		0				1

		26		60		1273.79		40000		0		0				1

		27		60		1655.81		40000		0		0				1

		28		60		2009.11		40000		0		0				1

		29		60		2333.6		40000		0		0				1

		30		60		2631.23		40000		0		0				1

		31		61		2904.98		40000		0		0				1

		32		61		3158.22		40000		0		0				1

		33		61		3393.99		40000		0		0				1

		34		61		-3393.99		40000		0		0				1

		35		61		-3158.22		40000		0		0				1

		36		61		-2904.98		40000		0		0				1

		37		60		-2631.23		40000		0		0				1

		38		60		-2333.6		40000		0		0				1

		39		60		-2009.11		40000		0		0				1

		40		60		-1655.81		40000		0		0				1

		41		60		-1273.79		40000		0		0				1

		42		60		-865.99		40000		0		0				1

		43		60		-438.358		40000		0		0				1

		44		60         0.		264525		40000		0		0				1

		45		60         0.		151705		40000		0		0				1

		46		60		515.215		40000		0		0				1

		47		60		1015.77		40000		0		0				1

		48		60		1489.3		40000		0		0				1

		49		60		1927.16		40000		0		0				1

		50		60		2325.48		40000		0		0				1

		51		61		2683.98		40000		0		0				1

		52		61		3005.46		40000		0		0				1

		53		61		3294.31		40000		0		0				1

		54		62		3556.92		40000		0		0				1

		55		62		3799.53		40000		0		0				1

		56		62		-3799.53		40000		0		0				1

		57		62		-3556.92		40000		0		0				1

		58		61		-3294.31		40000		0		0				1

		59		61		-3005.46		40000		0		0				1

		60		61		-2683.98		40000		0		0				1

		61		60		-2325.48		40000		0		0				1

		62		60		-1927.16		40000		0		0				1

		63		60		-1489.3		40000		0		0				1

		64		60		-1015.77		40000		0		0				1

		65		60		-515.215		40000		0		0				1

		66		60         0.		151705		40000		0		0				1

		67		59         0.		176971		40000		0		0				1

		68		59		577.426		40000		0		0				1

		69		59		1135.52		40000		0		0				1

		70		60		1657.53		40000		0		0				1

		71		60		2131.12		40000		0		0				1

		72		60		2548.73		40000		0		0				1

		73		61		2907.6		40000		0		0				1

		74		61		3206.96		40000		0		0				1

		75		62		3447.67		40000		0		0				1

		76		62		3628.97		40000		0		0				1

		77		63		3748.52		40000		0		0				1

		78		63		-3748.52		40000		0		0				1

		79		62		-3628.97		40000		0		0				1

		80		62		-3447.67		40000		0		0				1

		81		61		-3206.96		40000		0		0				1

		82		61		-2907.6		40000		0		0				1

		83		60		-2548.73		40000		0		0				1

		84		60		-2131.12		40000		0		0				1

		85		60		-1657.53		40000		0		0				1

		86		59		-1135.52		40000		0		0				1

		87		59		-577.426		40000		0		0				1

		88		59         0.		176971		40000		0		0				1

		89		59         0.		136128		40000		0		0				1

		90		59		633.089		40000		0		0				1

		91		59		1242.2		40000		0		0				1

		92		59		1806.29		40000		0		0				1

		93		59		2309.08		40000		0		0				1

		94		60		2740.22		40000		0		0				1

		95		60		3094.7		40000		0		0				1

		96		61		3372.2		40000		0		0				1

		97		61		3573.89		40000		0		0				1

		98		62		3703.6		40000		0		0				1

		99		63		3766.15		40000		0		0				1

		100		63		-3766.15		40000		0		0				1

		101		62		-3703.6		40000		0		0				1

		102		61		-3573.89		40000		0		0				1

		103		61		-3372.2		40000		0		0				1

		104		60		-3094.7		40000		0		0				1

		105		60		-2740.22		40000		0		0				1

		106		59		-2309.08		40000		0		0				1

		107		59		-1806.29		40000		0		0				1

		108		59		-1242.2		40000		0		0				1

		109		59		-633.089		40000		0		0				1

		110		59         0.		136128		40000		0		0				1

		111		58         0.		121656		40000		0		0				1

		112		58		735.984		40000		0		0				1

		113		58		1441.23		40000		0		0				1

		114		58		2087.59		40000		0		0				1

		115		59		2652.55		40000		0		0				1

		116		59		3122.82		40000		0		0				1

		117		59		3496.06		40000		0		0				1

		118		60		3780.53		40000		0		0				1

		119		60		3988.92		40000		0		0				1

		120		61		4136.59		40000		0		0				1

		121		4		279.196         9		60.234		0		0				0

		122		4		-279.196        9		60.234		0		0				0

		123		61		-4136.59		40000		0		0				1

		124		60		-3988.92		40000		0		0				1

		125		60		-3780.53		40000		0		0				1

		126		59		-3496.06		40000		0		0				1

		127		59		-3122.82		40000		0		0				1

		128		59		-2652.55		40000		0		0				1

		129		58		-2087.59		40000		0		0				1

		130		58		-1441.23		40000		0		0				1

		131		58		-735.984		40000		0		0				1

		132		58         0.		121656		40000		0		0				1

		133		57          0		0.00017908		40000		0		0				1

		134		57		832.935		40000		0		0				1

		135		57		1632.7		40000		0		0				1

		136		58		2367.64		40000		0		0				1

		137		58		3009.62		40000		0		0				1

		138		58		3538.91		40000		0		0				1

		139		58		3952.61		40000		0		0				1

		140		58		4268.02		40000		0		0				1

		141		4		268.533		963.27		0		0				0

		142		4		288.625         9		57.442		0		0				0

		143		4		306.282         9		51.941		0		0				0

		144		4		-306.282        9		51.941		0		0				0

		145		4		-288.625        9		57.442		0		0				0

		146		4		-268.533		963.27		0		0				0

		147		58		-4268.02		40000		0		0				1

		148		58		-3952.61		40000		0		0				1

		149		58		-3538.91		40000		0		0				1

		150		58		-3009.62		40000		0		0				1

		151		58		-2367.64		40000		0		0				1

		152		57		-1632.7		40000		0		0				1

		153		57		-832.935		40000		0		0				1

		154		57          0		0.00017908		40000		0		0				1

		155		57         0.		379989		40000		0		0				1

		156		57		923.737		40000		0		0				1

		157		57		1816.39		40000		0		0				1

		158		57		2648.27		40000		0		0				1

		159		57		3387.41		40000		0		0				1

		160		57		3993.72		40000		0		0				1

		161		57		4442.31		40000		0		0				1

		162		57		4767.49		40000		0		0				1

		163		4		288.236         9		57.559		0		0				0

		164		4		309.899         9		50.769		0		0				0

		165		4		328.98         9		44.337		0		0				0

		166		4		-328.98         9		44.337		0		0				0

		167		4		-309.899        9		50.769		0		0				0

		168		4		-288.236        9		57.559		0		0				0

		169		57		-4767.49		40000		0		0				1

		170		57		-4442.31		40000		0		0				1

		171		57		-3993.72		40000		0		0				1

		172		57		-3387.41		40000		0		0				1

		173		57		-2648.27		40000		0		0				1

		174		57		-1816.39		40000		0		0				1

		175		57		-923.737		40000		0		0				1

		176		57         0.		379989		40000		0		0				1

		177		56         0.		238816		40000		0		0				1

		178		56		1006.55		40000		0		0				1

		179		56		1986.99		40000		0		0				1

		180		56		2921.07		40000		0		0				1

		181		56		3787.02		40000		0		0				1

		182		56		4509.15		40000		0		0				1

		183		56		4951.45		40000		0		0				1

		184		4		278.989         9		60.294		0		0				0

		185		4		305.069		952.33		0		0				0

		186		4		328.131         9		44.632		0		0				0

		187		4		348.456         9		37.325		0		0				0

		188		4		-348.456        9		37.325		0		0				0

		189		4		-328.131        9		44.632		0		0				0

		190		4		-305.069		952.33		0		0				0

		191		4		-278.989        9		60.294		0		0				0

		192		56		-4951.45		40000		0		0				1

		193		56		-4509.15		40000		0		0				1

		194		56		-3787.02		40000		0		0				1

		195		56		-2921.07		40000		0		0				1

		196		56		-1986.99		40000		0		0				1

		197		56		-1006.55		40000		0		0				1

		198		56         0.		238816		40000		0		0				1

		199		56         0.		188636		40000		0		0				1

		200		56		1080.4		40000		0		0				1

		201		56		2139.38		40000		0		0				1

		202		56		3171.27		40000		0		0				1

		203		56		4199.05		40000		0		0				1

		204		55		5170.71		40000		0		0				1

		205		54		5427.47		40000		0		0				1

		206		4		292.281         9		56.332		0		0				0

		207		4		319.803         9		47.484		0		0				0

		208		4		344.106         9		38.931		0		0				0

		209		4		365.508         9		30.808		0		0				0

		210		4		-365.508        9		30.808		0		0				0

		211		4		-344.106        9		38.931		0		0				0

		212		4		-319.803        9		47.484		0		0				0

		213		4		-292.281        9		56.332		0		0				0

		214		54		-5427.47		40000		0		0				1

		215		55		-5170.71		40000		0		0				1

		216		56		-4199.05		40000		0		0				1

		217		56		-3171.27		40000		0		0				1

		218		56		-2139.38		40000		0		0				1

		219		56		-1080.4		40000		0		0				1

		220		56         0.		188636		40000		0		0				1

																				%%

																190				86.4

																30				13.6





130 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		60		0.000175744		40000		0		0				1

		2		60		301.801		40000		0		0				1

		3		60		595.716		40000		0		0				1

		4		60		874.992		40000		0		0				1

		5		60		1134.78		40000		0		0				1

		6		60		1372.37		40000		0		0				1

		7		60		1586.95		40000		0		0				1

		8		60		1779.15		40000		0		0				1

		9		61		1950.46		40000		0		0				1

		10		61		2102.81		40000		0		0				1

		11		61		2238.28		40000		0		0				1

		12		61		-2238.28		40000		0		0				1

		13		61		-2102.81		40000		0		0				1

		14		61		-1950.46		40000		0		0				1

		15		60		-1779.15		40000		0		0				1

		16		60		-1586.95		40000		0		0				1

		17		60		-1372.37		40000		0		0				1

		18		60		-1134.78		40000		0		0				1

		19		60		-874.992		40000		0		0				1

		20		60		-595.715		40000		0		0				1

		21		60		-301.801		40000		0		0				1

		22		60		0.000175744		40000		0		0				1

		23		60		0.000233889		40000		0		0				1

		24		60		404.202		40000		0		0				1

		25		60		796.31		40000		0		0				1

		26		60		1166.06		40000		0		0				1

		27		60		1506.13		40000		0		0				1

		28		60		1812.99		40000		0		0				1

		29		60		2085.95		40000		0		0				1

		30		61		2326.62		40000		0		0				1

		31		61		2538.08		40000		0		0				1

		32		61		2723.79		40000		0		0				1

		33		61		2887.9		40000		0		0				1

		34		61		-2887.9		40000		0		0				1

		35		61		-2723.79		40000		0		0				1

		36		61		-2538.08		40000		0		0				1

		37		61		-2326.62		40000		0		0				1

		38		60		-2085.95		40000		0		0				1

		39		60		-1812.99		40000		0		0				1

		40		60		-1506.13		40000		0		0				1

		41		60		-1166.06		40000		0		0				1

		42		60		-796.31		40000		0		0				1

		43		60		-404.202		40000		0		0				1

		44		60		0.000233889		40000		0		0				1

		45		59		0.000246283		40000		0		0				1

		46		59		480.667		40000		0		0				1

		47		59		944.814		40000		0		0				1

		48		60		1378.27		40000		0		0				1

		49		60		1770.87		40000		0		0				1

		50		60		2117.15		40000		0		0				1

		51		60		2415.61		40000		0		0				1

		52		61		2668.52		40000		0		0				1

		53		61		2879.43		40000		0		0				1

		54		61		3052.74		40000		0		0				1

		55		62		3193.5		40000		0		0				1

		56		62		-3193.5		40000		0		0				1

		57		61		-3052.74		40000		0		0				1

		58		61		-2879.43		40000		0		0				1

		59		61		-2668.52		40000		0		0				1

		60		60		-2415.61		40000		0		0				1

		61		60		-2117.15		40000		0		0				1

		62		60		-1770.87		40000		0		0				1

		63		60		-1378.27		40000		0		0				1

		64		59		-944.814		40000		0		0				1

		65		59		-480.667		40000		0		0				1

		66		59		0.000246283		40000		0		0				1

		67		59		0.000246056		40000		0		0				1

		68		59		547.25		40000		0		0				1

		69		59		1073.97		40000		0		0				1

		70		59		1562.17		40000		0		0				1

		71		59		1998.91		40000		0		0				1

		72		60		2376.79		40000		0		0				1

		73		60		2694.03		40000		0		0				1

		74		60		2953.04		40000		0		0				1

		75		60		3159.33		40000		0		0				1

		76		61		3318.99		40000		0		0				1

		77		61		3438.11		40000		0		0				1

		78		61		-3438.11		40000		0		0				1

		79		61		-3318.99		40000		0		0				1

		80		60		-3159.33		40000		0		0				1

		81		60		-2953.04		40000		0		0				1

		82		60		-2694.03		40000		0		0				1

		83		60		-2376.79		40000		0		0				1

		84		59		-1998.91		40000		0		0				1

		85		59		-1562.17		40000		0		0				1

		86		59		-1073.97		40000		0		0				1

		87		59		-547.25		40000		0		0				1

		88		59		0.000246056		40000		0		0				1

		89		58		0.000378418		40000		0		0				1

		90		58		608.735		40000		0		0				1

		91		58		1193.77		40000		0		0				1

		92		59		1734.09		40000		0		0				1

		93		59		2213.79		40000		0		0				1

		94		59		2623.64		40000		0		0				1

		95		59		2961.69		40000		0		0				1

		96		60		3232.22		40000		0		0				1

		97		60		3442.78		40000		0		0				1

		98		60		3602.08		40000		0		0				1

		99		61		3719.37		40000		0		0				1

		100		61		-3719.37		40000		0		0				1

		101		60		-3602.08		40000		0		0				1

		102		60		-3442.78		40000		0		0				1

		103		60		-3232.22		40000		0		0				1

		104		59		-2961.69		40000		0		0				1

		105		59		-2623.69		40000		0		0				1

		106		59		-2213.79		40000		0		0				1

		107		59		-1734.09		40000		0		0				1

		108		58		-1193.77		40000		0		0				1

		109		58		-608.735		40000		0		0				1

		110		58		0.000378418		40000		0		0				1

		111		58		0.000270141		40000		0		0				1

		112		58		722.175		40000		0		0				1

		113		58		1417.61		40000		0		0				1

		114		58		2061.67		40000		0		0				1

		115		58		2633.48		40000		0		0				1

		116		58		3117.87		40000		0		0				1

		117		58		3508.58		40000		0		0				1

		118		58		3812.24		40000		0		0				1

		119		58		4043.82		40000		0		0				1

		120		59		4220.01		40000		0		0				1

		121		4		269.945         9		62.876		0		0				0

		122		4		-269.945        9		62.876		0		0				0

		123		59		-4220.01		40000		0		0				1

		124		58		-4043.82		40000		0		0				1

		125		58		-3812.24		40000		0		0				1

		126		58		-3508.58		40000		0		0				1

		127		58		-3117.87		40000		0		0				1

		128		58		-2633.48		40000		0		0				1

		129		58		-2061.67		40000		0		0				1

		130		58		-1417.61		40000		0		0				1

		131		58		-722.175		40000		0		0				1

		132		58		0.000270141		40000		0		0				1

		133		57		0.000230875		40000		0		0				1

		134		57		824.323		40000		0		0				1

		135		57		1623.1		40000		0		0				1

		136		57		2373.5		40000		0		0				1

		137		57		3053.56		40000		0		0				1

		138		57		3636.61		40000		0		0				1

		139		57		4091.66		40000		0		0				1

		140		57		4415.39		40000		0		0				1

		141		57		4646.7		40000		0		0				1

		142		4		279.568         9		60.126		0		0				0

		143		4		296.641         9		54.989		0		0				0

		144		4		-296.641        9		54.989		0		0				0

		145		4		-279.568        9		60.126		0		0				0

		146		57		-4646.7		40000		0		0				1

		147		57		-4415.39		40000		0		0				1

		148		57		-4091.66		40000		0		0				1

		149		57		-3636.61		40000		0		0				1

		150		57		-3053.56		40000		0		0				1

		151		57		-2373.5		40000		0		0				1

		152		57		-1623.17		40000		0		0				1

		153		57		-824.323		40000		0		0				1

		154		57		0.000230875		40000		0		0				1

		155		56		0.000162103		40000		0		0				1

		156		56		914.439		40000		0		0				1

		157		56		1807.02		40000		0		0				1

		158		56		2662.38		40000		0		0				1

		159		56		3471.93		40000		0		0				1

		160		56		4208.37		40000		0		0				1

		161		56		4753.94		40000		0		0				1

		162		55		5013.01		40000		0		0				1

		163		4		279.463         9		60.156		0		0				0

		164		4		300.452         9		53.797		0		0				0

		165		4		318.956		947.77		0		0				0

		166		4		-318.956		947.77		0		0				0

		167		4		-300.452        9		53.797		0		0				0

		168		4		-279.463        9		60.156		0		0				0

		169		55		-5013.01		40000		0		0				1

		170		56		-4753.94		40000		0		0				1

		171		56		-4208.37		40000		0		0				1

		172		56		-3471.93		40000		0		0				1

		173		56		-2662.38		40000		0		0				1

		174		56		-1807.02		40000		0		0				1

		175		56		-914.439		40000		0		0				1

		176		56		0.000162103		40000		0		0				1

		177		56		0.000373856		40000		0		0				1

		178		56		992.739		40000		0		0				1

		179		56		1965.9		40000		0		0				1

		180		56		2913.09		40000		0		0				1

		181		56		3859.11		40000		0		0				1

		182		56		4871.39		40000		0		0				1

		183		54		5691.3		40000		0		0				1

		184		4		270.603         9		62.691		0		0				0

		185		4		296.016         9		55.183		0		0				0

		186		4		318.462         9		47.936		0		0				0

		187		4		338.216         9		41.068		0		0				0

		188		4		-338.216        9		41.068		0		0				0

		189		4		-318.462        9		47.936		0		0				0

		190		4		-296.016        9		55.183		0		0				0

		191		4		-270.603        9		62.691		0		0				0

		192		54		-5691.3		40000		0		0				1

		193		56		-4871.39		40000		0		0				1

		194		56		-3859.11		40000		0		0				1

		195		56		-2913.09		40000		0		0				1

		196		56		-1965.9		40000		0		0				1

		197		56		-992.739		40000		0		0				1

		198		56		0.000373856		40000		0		0				1

		199		55		0.000234203		40000		0		0				1

		200		55		1060.6		40000		0		0				1

		201		55		2099.88		40000		0		0				1

		202		56		3113.72		40000		0		0				1

		203		56		4147.89		40000		0		0				1

		204		56		5468.69		40000		0		0				1

		205		56		4325.78		40000		0		0				1

		206		3		284.108         9		58.792		0		0				0

		207		3		310.84         9		50.462		0		0				0

		208		4		334.415         9		42.426		0		0				0

		209		4		355.198         9		34.791		0		0				0

		210		4		-355.198        9		34.791		0		0				0

		211		4		-334.415        9		42.426		0		0				0

		212		3		-310.84         9		50.462		0		0				0

		213		3		-284.108        9		58.792		0		0				0

		214		56		-4325.78		40000		0		0				1

		215		56		-5468.69		40000		0		0				1

		216		56		-4148.06		40000		0		0				1

		217		56		-3113.72		40000		0		0				1

		218		55		-2099.88		40000		0		0				1

		219		55		-1060.6		40000		0		0				1

		220		55		0.000234203		40000		0		0				1

																				%%

																192				87.3

																28				12.7





140 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		60		0.000200597		40000		0		0				1

		2		60		277.37		40000		0		0				1

		3		60		546.592		40000		0		0				1

		4		60		800.813		40000		0		0				1

		5		60		1035.11		40000		0		0				1

		6		60		1246.83		40000		0		0				1

		7		60		1435.34		40000		0		0				1

		8		60		1601.51		40000		0		0				1

		9		61		1747.12		40000		0		0				1

		10		61		1874.4		40000		0		0				1

		11		61		1985.68		40000		0		0				1

		12		61		-1985.68		40000		0		0				1

		13		61		-1874.4		40000		0		0				1

		14		61		-1747.12		40000		0		0				1

		15		60		-1601.51		40000		0		0				1

		16		60		-1435.34		40000		0		0				1

		17		60		-1246.83		40000		0		0				1

		18		60		-1035.11		40000		0		0				1

		19		60		-800.813		40000		0		0				1

		20		60		-546.592		40000		0		0				1

		21		60		-277.37		40000		0		0				1

		22		60		0.000200597		40000		0		0				1

		23		59		0.000261126		40000		0		0				1

		24		59		379.834		40000		0		0				1

		25		59		747.012		40000		0		0				1

		26		60		1090.74		40000		0		0				1

		27		60		1403.4		40000		0		0				1

		28		60		1680.97		40000		0		0				1

		29		60		1922.74		40000		0		0				1

		30		60		2130.21		40000		0		0				1

		31		60		2306.85		40000		0		0				1

		32		61		2456.37		40000		0		0				1

		33		61		2582.65		40000		0		0				1

		34		61		-2582.65		40000		0		0				1

		35		61		-2456.37		40000		0		0				1

		36		60		-2306.85		40000		0		0				1

		37		60		-2130.21		40000		0		0				1

		38		60		-1922.74		40000		0		0				1

		39		60		-1680.97		40000		0		0				1

		40		60		-1403.4		40000		0		0				1

		41		60		-1090.74		40000		0		0				1

		42		59		-747.012		40000		0		0				1

		43		59		-379.834		40000		0		0				1

		44		59		0.000261126		40000		0		0				1

		45		59		0.000192024		40000		0		0				1

		46		59		461.374		40000		0		0				1

		47		59		906.231		40000		0		0				1

		48		59		1320.32		40000		0		0				1

		49		59		1693.2		40000		0		0				1

		50		59		2019.5		40000		0		0				1

		51		59		2298.01		40000		0		0				1

		52		60		2531.05		40000		0		0				1

		53		60		2723.11		40000		0		0				1

		54		60		2879.41		40000		0		0				1

		55		60		3005.79		40000		0		0				1

		56		60		-3005.79		40000		0		0				1

		57		60		-2879.41		40000		0		0				1

		58		60		-2723.11		40000		0		0				1

		59		60		-2531.05		40000		0		0				1

		60		59		-2298.01		40000		0		0				1

		61		59		-2019.5		40000		0		0				1

		62		59		-1693.2		40000		0		0				1

		63		59		-1320.32		40000		0		0				1

		64		59		-906.231		40000		0		0				1

		65		59		-461.374		40000		0		0				1

		66		59		0.000192024		40000		0		0				1

		67		58		0.000151182		40000		0		0				1

		68		58		533.493		40000		0		0				1

		69		58		1047.62		40000		0		0				1

		70		58		1525.36		40000		0		0				1

		71		59		1953.88		40000		0		0				1

		72		59		2325.65		40000		0		0				1

		73		59		2638.65		40000		0		0				1

		74		59		2895.42		40000		0		0				1

		75		59		3102.2		40000		0		0				1

		76		59		3266.28		40000		0		0				1

		77		60		3395.01		40000		0		0				1

		78		60		-3395.01		40000		0		0				1

		79		59		-3266.28		40000		0		0				1

		80		59		-3102.2		40000		0		0				1

		81		59		-2895.52		40000		0		0				1

		82		59		-2638.65		40000		0		0				1

		83		59		-2325.65		40000		0		0				1

		84		59		-1953.88		40000		0		0				1

		85		58		-1525.36		40000		0		0				1

		86		58		-1047.62		40000		0		0				1

		87		58		-533.493		40000		0		0				1

		88		58		0.000151182		40000		0		0				1

		89		58		0.000270491		40000		0		0				1

		90		58		599.534		40000		0		0				1

		91		58		1178.04		40000		0		0				1

		92		58		1716.55		40000		0		0				1

		93		58		2200.29		40000		0		0				1

		94		58		2619.11		40000		0		0				1

		95		58		2968.73		40000		0		0				1

		96		58		3251.15		40000		0		0				1

		97		58		3473.67		40000		0		0				1

		98		59		3646.35		40000		0		0				1

		99		59		3779.43		40000		0		0				1

		100		59		-3779.43		40000		0		0				1

		101		59		-3646.35		40000		0		0				1

		102		58		-3473.67		40000		0		0				1

		103		58		-3251.15		40000		0		0				1

		104		58		-2968.73		40000		0		0				1

		105		58		-2619.11		40000		0		0				1

		106		58		-2200.29		40000		0		0				1

		107		58		-1716.55		40000		0		0				1

		108		58		-1178.04		40000		0		0				1

		109		58		-599.534		40000		0		0				1

		110		58		0.000270491		40000		0		0				1

		111		57		0.000220558		40000		0		0				1

		112		57		716.803		40000		0		0				1

		113		57		1412.52		40000		0		0				1

		114		57		2068.55		40000		0		0				1

		115		57		2669.12		40000		0		0				1

		116		57		3198.6		40000		0		0				1

		117		57		3639.1		40000		0		0				1

		118		57		3977.17		40000		0		0				1

		119		57		4219.11		40000		0		0				1

		120		57		4391.05		40000		0		0				1

		121		57		4519.63		40000		0		0				1

		122		57		-4519.63		40000		0		0				1

		123		57		-4391.05		40000		0		0				1

		124		57		-4219.11		40000		0		0				1

		125		57		-3977.17		40000		0		0				1

		126		57		-3639.1		40000		0		0				1

		127		57		-3198.6		40000		0		0				1

		128		57		-2669.12		40000		0		0				1

		129		57		-2068.55		40000		0		0				1

		130		57		-1412.52		40000		0		0				1

		131		57		-716.802		40000		0		0				1

		132		57		0.000220558		40000		0		0				1

		133		57		0.000246585		40000		0		0				1

		134		57		816.411		40000		0		0				1

		135		57		1613.73		40000		0		0				1

		136		57		2378.36		40000		0		0				1

		137		57		3104.69		40000		0		0				1

		138		57		3787.7		40000		0		0				1

		139		56		4388.53		40000		0		0				1

		140		56		4792.59		40000		0		0				1

		141		55		4956.95		40000		0		0				1

		142		4		271.237         9		62.513		0		0				0

		143		4		287.77         9		57.699		0		0				0

		144		4		-287.77         9		57.699		0		0				0

		145		4		-271.237        9		62.513		0		0				0

		146		55		-4956.95		40000		0		0				1

		147		56		-4792.59		40000		0		0				1

		148		56		-4388.53		40000		0		0				1

		149		57		-3787.7		40000		0		0				1

		150		57		-3104.69		40000		0		0				1

		151		57		-2378.36		40000		0		0				1

		152		57		-1613.73		40000		0		0				1

		153		57		-816.411		40000		0		0				1

		154		57		0.000246585		40000		0		0				1

		155		56		0.000271261		40000		0		0				1

		156		56		900.447		40000		0		0				1

		157		56		1782.13		40000		0		0				1

		158		56		2636.2		40000		0		0				1

		159		56		3475.52		40000		0		0				1

		160		56		4354.95		40000		0		0				1

		161		56		5377.62		40000		0		0				1

		162		53		5984.48		40000		0		0				1

		163		4		271.201         9		62.523		0		0				0

		164		4		291.697         9		56.511		0		0				0

		165		4		309.726         9		50.826		0		0				0

		166		4		-309.726        9		50.826		0		0				0

		167		4		-291.697        9		56.511		0		0				0

		168		4		-271.201        9		62.523		0		0				0

		169		53		-5984.48		40000		0		0				1

		170		56		-5377.62		40000		0		0				1

		171		56		-4354.95		40000		0		0				1

		172		56		-3475.52		40000		0		0				1

		173		56		-2636.2		40000		0		0				1

		174		56		-1782.13		40000		0		0				1

		175		56		-900.447		40000		0		0				1

		176		56		0.000271261		40000		0		0				1

		177		55		0.000360929		40000		0		0				1

		178		55		972.402		40000		0		0				1

		179		56		1922.49		40000		0		0				1

		180		56		2839.15		40000		0		0				1

		181		56		3736.95		40000		0		0				1

		182		57		4700.1		40000		0		0				1

		183		58		6216.03		40000		0		0				1

		184		56		4365.7		40000		0		0				1

		185		3		288.085         9		57.605		0		0				0

		186		3		309.845         9		50.787		0		0				0

		187		3		328.974         9		44.339		0		0				0

		188		3		-328.974        9		44.339		0		0				0

		189		3		-309.845        9		50.787		0		0				0

		190		3		-288.085        9		57.605		0		0				0

		191		56		-4365.7		40000		0		0				1

		192		58		-6216.03		40000		0		0				1

		193		57		-4700.1		40000		0		0				1

		194		56		-3736.95		40000		0		0				1

		195		56		-2839.15		40000		0		0				1

		196		56		-1922.49		40000		0		0				1

		197		55		-972.402		40000		0		0				1

		198		55		0.000360929		40000		0		0				1

		199		55		0.000164245		40000		0		0				1

		200		55		1036.09		40000		0		0				1

		201		55		2042.5		40000		0		0				1

		202		55		2997.87		40000		0		0				1

		203		56		3889.34		40000		0		0				1

		204		57		4669.76		40000		0		0				1

		205		60		4742.64		40000		0		0				1

		206		3		277.153         9		60.826		0		0				0

		207		3		303.117         9		52.953		0		0				0

		208		3		325.992         9		45.373		0		0				0

		209		3		346.113         9		38.193		0		0				0

		210		3		-346.113        9		38.193		0		0				0

		211		3		-325.992        9		45.373		0		0				0

		212		3		-303.117        9		52.953		0		0				0

		213		3		-277.153        9		60.826		0		0				0

		214		60		-4742.64		40000		0		0				1

		215		57		-4669.76		40000		0		0				1

		216		56		-3889.34		40000		0		0				1

		217		55		-2997.87		40000		0		0				1

		218		55		-2042.5		40000		0		0				1

		219		55		-1036.09		40000		0		0				1

		220		55		0.000164245		40000		0		0				1

																				%%

																196				89.1

																24				10.9





150 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		59		0.00013328		40000		0		0				1

		2		60		261.636		40000		0		0				1

		3		60		515.087		40000		0		0				1

		4		60		753.377		40000		0		0				1

		5		60		971.564		40000		0		0				1

		6		60		1167		40000		0		0				1

		7		60		1339.1		40000		0		0				1

		8		60		1488.81		40000		0		0				1

		9		60		1617.9		40000		0		0				1

		10		60		1728.91		40000		0		0				1

		11		60		1824.2		40000		0		0				1

		12		60		-1824.2		40000		0		0				1

		13		60		-1728.91		40000		0		0				1

		14		60		-1617.9		40000		0		0				1

		15		60		-1488.81		40000		0		0				1

		16		60		-1339.1		40000		0		0				1

		17		60		-1167		40000		0		0				1

		18		60		-971.564		40000		0		0				1

		19		60		-753.377		40000		0		0				1

		20		60		-515.087		40000		0		0				1

		21		60		-261.636		40000		0		0				1

		22		59		0.00013328		40000		0		0				1

		23		59		0.000191864		40000		0		0				1

		24		59		368.628		40000		0		0				1

		25		59		724.82		40000		0		0				1

		26		59		1057.91		40000		0		0				1

		27		59		1360.25		40000		0		0				1

		28		59		1627.52		40000		0		0				1

		29		59		1859.02		40000		0		0				1

		30		59		2056.24		40000		0		0				1

		31		59		2222.32		40000		0		0				1

		32		60		2361.15		40000		0		0				1

		33		60		2476.75		40000		0		0				1

		34		60		-2476.75		40000		0		0				1

		35		60		-2361.15		40000		0		0				1

		36		59		-2222.32		40000		0		0				1

		37		59		-2056.24		40000		0		0				1

		38		59		-1859.02		40000		0		0				1

		39		59		-1627.52		40000		0		0				1

		40		59		-1360.25		40000		0		0				1

		41		59		-1057.91		40000		0		0				1

		42		59		-724.82		40000		0		0				1

		43		59		-368.627		40000		0		0				1

		44		59		0.000191864		40000		0		0				1

		45		58		0.000153648		40000		0		0				1

		46		58		454.879		40000		0		0				1

		47		58		894.391		40000		0		0				1

		48		58		1305.13		40000		0		0				1

		49		58		1677.06		40000		0		0				1

		50		58		2004.2		40000		0		0				1

		51		59		2284.61		40000		0		0				1

		52		59		2519.77		40000		0		0				1

		53		59		2713.67		40000		0		0				1

		54		59		2871.62		40000		0		0				1

		55		59		2999.11		40000		0		0				1

		56		59		-2999.11		40000		0		0				1

		57		59		-2871.62		40000		0		0				1

		58		59		-2713.67		40000		0		0				1

		59		59		-2519.77		40000		0		0				1

		60		59		-2284.61		40000		0		0				1

		61		58		-2004.2		40000		0		0				1

		62		58		-1677.06		40000		0		0				1

		63		58		-1305.13		40000		0		0				1

		64		58		-894.391		40000		0		0				1

		65		58		-454.879		40000		0		0				1

		66		58		0.000153648		40000		0		0				1

		67		58		0.000155978		40000		0		0				1

		68		58		529.994		40000		0		0				1

		69		58		1043.02		40000		0		0				1

		70		58		1524.08		40000		0		0				1

		71		58		1961.41		40000		0		0				1

		72		58		2346.99		40000		0		0				1

		73		58		2676.48		40000		0		0				1

		74		58		2950.04		40000		0		0				1

		75		58		3171.07		40000		0		0				1

		76		58		3346.08		40000		0		0				1

		77		58		3483.01		40000		0		0				1

		78		58		-3483.01		40000		0		0				1

		79		58		-3346.08		40000		0		0				1

		80		58		-3171.07		40000		0		0				1

		81		58		-2950.04		40000		0		0				1

		82		58		-2676.58		40000		0		0				1

		83		58		-2346.99		40000		0		0				1

		84		58		-1961.41		40000		0		0				1

		85		58		-1524.08		40000		0		0				1

		86		58		-1043.02		40000		0		0				1

		87		58		-529.994		40000		0		0				1

		88		58		0.000155978		40000		0		0				1

		89		57		0.000323342		40000		0		0				1

		90		57		596.788		40000		0		0				1

		91		57		1176.13		40000		0		0				1

		92		58		1722.71		40000		0		0				1

		93		58		2224.33		40000		0		0				1

		94		58		2671.55		40000		0		0				1

		95		57		3056.7		40000		0		0				1

		96		57		3374.41		40000		0		0				1

		97		57		3624.25		40000		0		0				1

		98		57		3812.89		40000		0		0				1

		99		57		3952.44		40000		0		0				1

		100		57		-3952.44		40000		0		0				1

		101		57		-3812.89		40000		0		0				1

		102		57		-3624.25		40000		0		0				1

		103		57		-3374.41		40000		0		0				1

		104		57		-3056.7		40000		0		0				1

		105		58		-2671.55		40000		0		0				1

		106		58		-2224.33		40000		0		0				1

		107		58		-1722.71		40000		0		0				1

		108		57		-1176.13		40000		0		0				1

		109		57		-596.788		40000		0		0				1

		110		57		0.000323342		40000		0		0				1

		111		57		0.000226875		40000		0		0				1

		112		57		710.397		40000		0		0				1

		113		57		1403.83		40000		0		0				1

		114		57		2067.35		40000		0		0				1

		115		57		2694.16		40000		0		0				1

		116		57		3281.83		40000		0		0				1

		117		57		3822.88		40000		0		0				1

		118		56		4285.51		40000		0		0				1

		119		56		4609.27		40000		0		0				1

		120		56		4769.8		40000		0		0				1

		121		56		4834.77		40000		0		0				1

		122		56		-4834.77		40000		0		0				1

		123		56		-4769.8		40000		0		0				1

		124		56		-4609.27		40000		0		0				1

		125		56		-4285.51		40000		0		0				1

		126		57		-3822.88		40000		0		0				1

		127		57		-3281.83		40000		0		0				1

		128		57		-2694.16		40000		0		0				1

		129		57		-2067.35		40000		0		0				1

		130		57		-1403.83		40000		0		0				1

		131		57		-710.397		40000		0		0				1

		132		57		0.000226875		40000		0		0				1

		133		56		0.000179146		40000		0		0				1

		134		56		802.822		40000		0		0				1

		135		56		1587.71		40000		0		0				1

		136		56		2343.56		40000		0		0				1

		137		57		3072.75		40000		0		0				1

		138		57		3800.04		40000		0		0				1

		139		57		4588.68		40000		0		0				1

		140		56		5527.28		40000		0		0				1

		141		53		6164.67		40000		0		0				1

		142		53		5437.37		40000		0		0				1

		143		4		279.45		960.16		0		0				0

		144		4		-279.45		960.16		0		0				0

		145		53		-5437.37		40000		0		0				1

		146		53		-6164.67		40000		0		0				1

		147		56		-5527.28		40000		0		0				1

		148		57		-4588.68		40000		0		0				1

		149		57		-3800.04		40000		0		0				1

		150		57		-3072.75		40000		0		0				1

		151		56		-2343.56		40000		0		0				1

		152		56		-1587.71		40000		0		0				1

		153		56		-802.822		40000		0		0				1

		154		56		0.000179146		40000		0		0				1

		155		56		0.000349641		40000		0		0				1

		156		56		880.509		40000		0		0				1

		157		56		1738.63		40000		0		0				1

		158		56		2559.15		40000		0		0				1

		159		56		3341.7		40000		0		0				1

		160		57		4109.6		40000		0		0				1

		161		58		4938.9		40000		0		0				1

		162		61		6095.73		40000		0		0				1

		163		3		264.26         9		64.452		0		0				0

		164		3		284.138         9		58.783		0		0				0

		165		3		301.593         9		53.437		0		0				0

		166		3		-301.593        9		53.437		0		0				0

		167		3		-284.138        9		58.783		0		0				0

		168		56		-4266.46		40000		0		0				1

		169		61		-6095.73		40000		0		0				1

		170		58		-4938.9		40000		0		0				1

		171		57		-4109.6		40000		0		0				1

		172		56		-3341.7		40000		0		0				1

		173		56		-2559.15		40000		0		0				1

		174		56		-1738.63		40000		0		0				1

		175		56		-880.508		40000		0		0				1

		176		56		0.000349641		40000		0		0				1

		177		55		0.000331029		40000		0		0				1

		178		55		948.96		40000		0		0				1

		179		55		1868.14		40000		0		0				1

		180		56		2733.03		40000		0		0				1

		181		56		3524.22		40000		0		0				1

		182		57		4209.67		40000		0		0				1

		183		58		4637.56		40000		0		0				1

		184		61		4368.78		40000		0		0				1

		185		3		281.225         9		59.642		0		0				0

		186		3		302.361         9		53.193		0		0				0

		187		3		320.935         9		47.101		0		0				0

		188		3		-320.935        9		47.101		0		0				0

		189		3		-302.361        9		53.193		0		0				0

		190		3		-281.225        9		59.642		0		0				0

		191		61		-4368.78		40000		0		0				1

		192		58		-4637.56		40000		0		0				1

		193		57		-4209.67		40000		0		0				1

		194		56		-3524.22		40000		0		0				1

		195		56		-2733.03		40000		0		0				1

		196		55		-1868.14		40000		0		0				1

		197		55		-948.96		40000		0		0				1

		198		55		0.000331029		40000		0		0				1

		199		55		3.40E-05		40000		0		0				1

		200		55		1012.59		40000		0		0				1

		201		55		1986.77		40000		0		0				1

		202		55		2886.61		40000		0		0				1

		203		56		3674.84		40000		0		0				1

		204		56		4289.27		40000		0		0				1

		205		57		4627.63		40000		0		0				1

		206		3		270.915         9		62.603		0		0				0

		207		3		296.157         9		55.139		0		0				0

		208		3		318.389		947.96		0		0				0

		209		3		337.94         9		41.168		0		0				0

		210		3		-337.94         9		41.168		0		0				0

		211		3		-318.389		947.96		0		0				0

		212		3		-296.157        9		55.139		0		0				0

		213		3		-270.915        9		62.603		0		0				0

		214		57		-4627.63		40000		0		0				1

		215		56		-4289.27		40000		0		0				1

		216		56		-3674.84		40000		0		0				1

		217		55		-2886.61		40000		0		0				1

		218		55		-1986.77		40000		0		0				1

		219		55		-1012.59		40000		0		0				1

		220		55		3.40E-05		40000		0		0				1

																				%%

																199				90.5

																21				9.5





160 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		59		0.000157384		40000		0		0				1

		2		59		255.385		40000		0		0				1

		3		59		502.665		40000		0		0				1

		4		59		734.916		40000		0		0				1

		5		59		947.151		40000		0		0				1

		6		59		1136.67		40000		0		0				1

		7		59		1302.87		40000		0		0				1

		8		59		1446.6		40000		0		0				1

		9		59		1569.72		40000		0		0				1

		10		59		1674.59		40000		0		0				1

		11		59		1763.68		40000		0		0				1

		12		59		-1763.68		40000		0		0				1

		13		59		-1674.59		40000		0		0				1

		14		59		-1569.72		40000		0		0				1

		15		59		-1446.6		40000		0		0				1

		16		59		-1302.87		40000		0		0				1

		17		59		-1136.67		40000		0		0				1

		18		59		-947.167		40000		0		0				1

		19		59		-734.916		40000		0		0				1

		20		59		-502.665		40000		0		0				1

		21		59		-255.384		40000		0		0				1

		22		59		0.000157384		40000		0		0				1

		23		58		0.000289027		40000		0		0				1

		24		58		366.325		40000		0		0				1

		25		58		720.955		40000		0		0				1

		26		58		1053.75		40000		0		0				1

		27		58		1357.23		40000		0		0				1

		28		58		1626.99		40000		0		0				1

		29		58		1861.59		40000		0		0				1

		30		58		2062.01		40000		0		0				1

		31		59		2230.85		40000		0		0				1

		32		59		2371.71		40000		0		0				1

		33		59		2488.48		40000		0		0				1

		34		59		-2488.48		40000		0		0				1

		35		59		-2371.71		40000		0		0				1

		36		59		-2230.85		40000		0		0				1

		37		58		-2062.01		40000		0		0				1

		38		58		-1861.59		40000		0		0				1

		39		58		-1626.99		40000		0		0				1

		40		58		-1357.23		40000		0		0				1

		41		58		-1053.75		40000		0		0				1

		42		58		-720.955		40000		0		0				1

		43		58		-366.325		40000		0		0				1

		44		58		0.000289027		40000		0		0				1

		45		58		0.000228366		40000		0		0				1

		46		58		454.183		40000		0		0				1

		47		58		894.734		40000		0		0				1

		48		58		1309.73		40000		0		0				1

		49		58		1690.03		40000		0		0				1

		50		58		2029.71		40000		0		0				1

		51		58		2325.83		40000		0		0				1

		52		58		2578.06		40000		0		0				1

		53		58		2788.37		40000		0		0				1

		54		58		2960.47		40000		0		0				1

		55		58		3099.26		40000		0		0				1

		56		58		-3099.26		40000		0		0				1

		57		58		-2960.47		40000		0		0				1

		58		58		-2788.37		40000		0		0				1

		59		58		-2578.06		40000		0		0				1

		60		58		-2325.83		40000		0		0				1

		61		58		-2029.71		40000		0		0				1

		62		58		-1690.03		40000		0		0				1

		63		58		-1309.73		40000		0		0				1

		64		58		-894.734		40000		0		0				1

		65		58		-454.183		40000		0		0				1

		66		58		0.000228366		40000		0		0				1

		67		57		0.000221544		40000		0		0				1

		68		57		528.274		40000		0		0				1

		69		57		1042.1		40000		0		0				1

		70		57		1529.04		40000		0		0				1

		71		57		1979.71		40000		0		0				1

		72		57		2387.75		40000		0		0				1

		73		57		2748.78		40000		0		0				1

		74		57		3059.6		40000		0		0				1

		75		57		3318.27		40000		0		0				1

		76		57		3525.31		40000		0		0				1

		77		57		3684.85		40000		0		0				1

		78		57		-3684.85		40000		0		0				1

		79		57		-3525.31		40000		0		0				1

		80		57		-3318.27		40000		0		0				1

		81		57		-3059.6		40000		0		0				1

		82		57		-2748.78		40000		0		0				1

		83		57		-2387.75		40000		0		0				1

		84		57		-1979.71		40000		0		0				1

		85		57		-1529.04		40000		0		0				1

		86		57		-1042.1		40000		0		0				1

		87		57		-528.274		40000		0		0				1

		88		57		0.000221544		40000		0		0				1

		89		57		0.000161204		40000		0		0				1

		90		57		592.224		40000		0		0				1

		91		57		1169.62		40000		0		0				1

		92		57		1720.05		40000		0		0				1

		93		57		2235.53		40000		0		0				1

		94		57		2711.81		40000		0		0				1

		95		57		3145.93		40000		0		0				1

		96		57		3532.45		40000		0		0				1

		97		57		3860.15		40000		0		0				1

		98		56		4115.52		40000		0		0				1

		99		56		4293		40000		0		0				1

		100		56		-4293		40000		0		0				1

		101		56		-4115.52		40000		0		0				1

		102		57		-3860.15		40000		0		0				1

		103		57		-3532.52		40000		0		0				1

		104		57		-3145.93		40000		0		0				1

		105		57		-2711.81		40000		0		0				1

		106		57		-2235.53		40000		0		0				1

		107		57		-1720.05		40000		0		0				1

		108		57		-1169.62		40000		0		0				1

		109		57		-592.223		40000		0		0				1

		110		57		0.000161204		40000		0		0				1

		111		56		0.000167766		40000		0		0				1

		112		56		697.95		40000		0		0				1

		113		57		1379.34		40000		0		0				1

		114		57		2032.18		40000		0		0				1

		115		57		2652.73		40000		0		0				1

		116		57		3248.26		40000		0		0				1

		117		57		3839.42		40000		0		0				1

		118		57		4463.34		40000		0		0				1

		119		56		5148.97		40000		0		0				1

		120		55		5771.9		40000		0		0				1

		121		53		5853.05		40000		0		0				1

		122		53		-5853.05		40000		0		0				1

		123		55		-5771.9		40000		0		0				1

		124		56		-5148.97		40000		0		0				1

		125		57		-4463.34		40000		0		0				1

		126		57		-3839.49		40000		0		0				1

		127		57		-3248.26		40000		0		0				1

		128		57		-2652.73		40000		0		0				1

		129		57		-2032.18		40000		0		0				1

		130		57		-1379.34		40000		0		0				1

		131		56		-697.95		40000		0		0				1

		132		56		0.000167766		40000		0		0				1

		133		56		0.000138922		40000		0		0				1

		134		56		783.981		40000		0		0				1

		135		56		1546.56		40000		0		0				1

		136		56		2271.81		40000		0		0				1

		137		57		2952.83		40000		0		0				1

		138		57		3594.68		40000		0		0				1

		139		57		4217.74		40000		0		0				1

		140		58		4868.01		40000		0		0				1

		141		61		5616.05		40000		0		0				1

		142		56		4043.12		40000		0		0				1

		143		3		272.375         9		62.191		0		0				0

		144		3		-272.375        9		62.191		0		0				0

		145		56		-4043.12		40000		0		0				1

		146		61		-5616.05		40000		0		0				1

		147		58		-4868.01		40000		0		0				1

		148		57		-4217.74		40000		0		0				1

		149		57		-3594.68		40000		0		0				1

		150		57		-2952.83		40000		0		0				1

		151		56		-2271.81		40000		0		0				1

		152		56		-1546.56		40000		0		0				1

		153		56		-783.981		40000		0		0				1

		154		56		0.000138922		40000		0		0				1

		155		55		0.000267562		40000		0		0				1

		156		55		858.872		40000		0		0				1

		157		55		1689.33		40000		0		0				1

		158		56		2467.15		40000		0		0				1

		159		56		3173.21		40000		0		0				1

		160		57		3787.79		40000		0		0				1

		161		58		4263.14		40000		0		0				1

		162		59		4454.11		40000		0		0				1

		163		61		4198.26		40000		0		0				1

		164		3		277.552         9		60.711		0		0				0

		165		3		294.506		955.65		0		0				0

		166		3		-294.506		955.65		0		0				0

		167		3		-277.552        9		60.711		0		0				0

		168		61		-4198.26		40000		0		0				1

		169		59		-4454.11		40000		0		0				1

		170		58		-4263.14		40000		0		0				1

		171		57		-3787.79		40000		0		0				1

		172		56		-3173.21		40000		0		0				1

		173		56		-2467.15		40000		0		0				1

		174		55		-1689.33		40000		0		0				1

		175		55		-858.872		40000		0		0				1

		176		55		0.000267562		40000		0		0				1

		177		55		0.000301265		40000		0		0				1

		178		55		928.36		40000		0		0				1

		179		55		1820.33		40000		0		0				1

		180		55		2642.35		40000		0		0				1

		181		56		3364.16		40000		0		0				1

		182		56		3950.14		40000		0		0				1

		183		57		4352.66		40000		0		0				1

		184		58		4544.87		40000		0		0				1

		185		3		275.058         9		61.428		0		0				0

		186		3		295.605		955.31		0		0				0

		187		3		313.656         9		49.537		0		0				0

		188		3		-313.656        9		49.537		0		0				0

		189		3		-295.605		955.31		0		0				0

		190		3		-275.058        9		61.428		0		0				0

		191		58		-4544.87		40000		0		0				1

		192		57		-4352.66		40000		0		0				1

		193		56		-3950.14		40000		0		0				1

		194		56		-3364.16		40000		0		0				1

		195		55		-2642.35		40000		0		0				1

		196		55		-1820.33		40000		0		0				1

		197		55		-928.36		40000		0		0				1

		198		55		0.000301265		40000		0		0				1

		199		54		0.000176945		40000		0		0				1

		200		54		995.569		40000		0		0				1

		201		54		1947.13		40000		0		0				1

		202		55		2813.42		40000		0		0				1

		203		55		3557.44		40000		0		0				1

		204		55		4147.21		40000		0		0				1

		205		56		4567.77		40000		0		0				1

		206		3		265.332         9		64.157		0		0				0

		207		3		289.883         9		57.062		0		0				0

		208		3		311.501         9		50.246		0		0				0

		209		3		330.511         9		43.802		0		0				0

		210		3		-330.511        9		43.802		0		0				0

		211		3		-311.501        9		50.246		0		0				0

		212		3		-289.883        9		57.062		0		0				0

		213		3		-265.332        9		64.157		0		0				0

		214		56		-4567.77		40000		0		0				1

		215		55		-4147.21		40000		0		0				1

		216		55		-3557.44		40000		0		0				1

		217		55		-2813.42		40000		0		0				1

		218		54		-1947.13		40000		0		0				1

		219		54		-995.569		40000		0		0				1

		220		54		0.000176945		40000		0		0				1

																				%%

																200				90.9

																20				9.1





170 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		58		0.000155021		40000		0		0				1

		2		58		255.049		40000		0		0				1

		3		58		502.237		40000		0		0				1

		4		58		734.796		40000		0		0				1

		5		58		947.801		40000		0		0				1

		6		58		1138.45		40000		0		0				1

		7		58		1305.91		40000		0		0				1

		8		58		1450.88		40000		0		0				1

		9		59		1575.04		40000		0		0				1

		10		59		1680.64		40000		0		0				1

		11		59		1770.08		40000		0		0				1

		12		59		-1770.08		40000		0		0				1

		13		59		-1680.64		40000		0		0				1

		14		59		-1575.04		40000		0		0				1

		15		58		-1450.88		40000		0		0				1

		16		58		-1305.91		40000		0		0				1

		17		58		-1138.45		40000		0		0				1

		18		58		-947.8		40000		0		0				1

		19		58		-734.795		40000		0		0				1

		20		58		-502.237		40000		0		0				1

		21		58		-255.049		40000		0		0				1

		22		58		0.000155021		40000		0		0				1

		23		58		0.000211907		40000		0		0				1

		24		58		366.865		40000		0		0				1

		25		58		723.033		40000		0		0				1

		26		58		1059.2		40000		0		0				1

		27		58		1368.41		40000		0		0				1

		28		58		1646.35		40000		0		0				1

		29		58		1891.23		40000		0		0				1

		30		58		2103.27		40000		0		0				1

		31		58		2284.19		40000		0		0				1

		32		58		2436.67		40000		0		0				1

		33		58		2563.96		40000		0		0				1

		34		58		-2563.96		40000		0		0				1

		35		58		-2436.67		40000		0		0				1

		36		58		-2284.19		40000		0		0				1

		37		58		-2103.27		40000		0		0				1

		38		58		-1891.23		40000		0		0				1

		39		58		-1646.35		40000		0		0				1

		40		58		-1368.41		40000		0		0				1

		41		58		-1059.2		40000		0		0				1

		42		58		-723.033		40000		0		0				1

		43		58		-366.865		40000		0		0				1

		44		58		0.000211907		40000		0		0				1

		45		57		0.000163626		40000		0		0				1

		46		57		452.394		40000		0		0				1

		47		57		892.684		40000		0		0				1

		48		57		1310.48		40000		0		0				1

		49		57		1698.16		40000		0		0				1

		50		58		2051.01		40000		0		0				1

		51		57		2366.72		40000		0		0				1

		52		57		2644.52		40000		0		0				1

		53		57		2884.61		40000		0		0				1

		54		57		3087.97		40000		0		0				1

		55		57		3256.46		40000		0		0				1

		56		57		-3256.46		40000		0		0				1

		57		57		-3087.97		40000		0		0				1

		58		57		-2884.61		40000		0		0				1

		59		57		-2644.52		40000		0		0				1

		60		57		-2366.72		40000		0		0				1

		61		58		-2051.01		40000		0		0				1

		62		57		-1698.16		40000		0		0				1

		63		57		-1310.48		40000		0		0				1

		64		57		-892.684		40000		0		0				1

		65		57		-452.394		40000		0		0				1

		66		57		0.000163626		40000		0		0				1

		67		57		0.000194944		40000		0		0				1

		68		57		522.487		40000		0		0				1

		69		57		1031.84		40000		0		0				1

		70		57		1517.24		40000		0		0				1

		71		57		1971.57		40000		0		0				1

		72		57		2391.55		40000		0		0				1

		73		57		2776.7		40000		0		0				1

		74		57		3127.44		40000		0		0				1

		75		57		3442.98		40000		0		0				1

		76		57		3720.21		40000		0		0				1

		77		57		3952.91		40000		0		0				1

		78		57		-3952.91		40000		0		0				1

		79		57		-3720.21		40000		0		0				1

		80		57		-3442.98		40000		0		0				1

		81		57		-3127.44		40000		0		0				1

		82		57		-2776.7		40000		0		0				1

		83		57		-2391.55		40000		0		0				1

		84		57		-1971.57		40000		0		0				1

		85		57		-1517.24		40000		0		0				1

		86		57		-1031.84		40000		0		0				1

		87		57		-522.487		40000		0		0				1

		88		57		0.000194944		40000		0		0				1

		89		57		0.000284271		40000		0		0				1

		90		57		582.058		40000		0		0				1

		91		57		1149.67		40000		0		0				1

		92		57		1691.4		40000		0		0				1

		93		57		2200.68		40000		0		0				1

		94		57		2677.15		40000		0		0				1

		95		57		3125.27		40000		0		0				1

		96		57		3553		40000		0		0				1

		97		57		3968.19		40000		0		0				1

		98		57		4371.51		40000		0		0				1

		99		56		4743.53		40000		0		0				1

		100		56		-4743.53		40000		0		0				1

		101		57		-4371.51		40000		0		0				1

		102		57		-3968.19		40000		0		0				1

		103		57		-3553		40000		0		0				1

		104		57		-3125.27		40000		0		0				1

		105		57		-2677.15		40000		0		0				1

		106		57		-2200.68		40000		0		0				1

		107		57		-1691.4		40000		0		0				1

		108		57		-1149.67		40000		0		0				1

		109		57		-582.058		40000		0		0				1

		110		57		0.000284271		40000		0		0				1

		111		56		0.000127313		40000		0		0				1

		112		56		680.911		40000		0		0				1

		113		56		1342.78		40000		0		0				1

		114		56		1970.46		40000		0		0				1

		115		57		2555.4		40000		0		0				1

		116		57		3096.99		40000		0		0				1

		117		57		3602.57		40000		0		0				1

		118		58		4090.36		40000		0		0				1

		119		58		4594.36		40000		0		0				1

		120		59		5197.84		40000		0		0				1

		121		61		6250.21		40000		0		0				1

		122		61		-6250.21		40000		0		0				1

		123		59		-5197.84		40000		0		0				1

		124		58		-4594.36		40000		0		0				1

		125		58		-4090.36		40000		0		0				1

		126		57		-3602.57		40000		0		0				1

		127		57		-3096.99		40000		0		0				1

		128		57		-2555.4		40000		0		0				1

		129		56		-1970.46		40000		0		0				1

		130		56		-1342.78		40000		0		0				1

		131		56		-680.911		40000		0		0				1

		132		56		0.000127313		40000		0		0				1

		133		55		0.000117482		40000		0		0				1

		134		56		764.577		40000		0		0				1

		135		56		1503.41		40000		0		0				1

		136		56		2194.53		40000		0		0				1

		137		56		2821.27		40000		0		0				1

		138		57		3371.65		40000		0		0				1

		139		57		3830.08		40000		0		0				1

		140		58		4163.16		40000		0		0				1

		141		60		4290.81		40000		0		0				1

		142		61		4107.18		40000		0		0				1

		143		3		266.232         9		63.909		0		0				0

		144		3		-266.232        9		63.909		0		0				0

		145		61		-4107.18		40000		0		0				1

		146		60		-4290.81		40000		0		0				1

		147		58		-4163.16		40000		0		0				1

		148		57		-3830.08		40000		0		0				1

		149		57		-3371.65		40000		0		0				1

		150		56		-2821.27		40000		0		0				1

		151		56		-2194.53		40000		0		0				1

		152		56		-1503.41		40000		0		0				1

		153		56		-764.577		40000		0		0				1

		154		55		0.000117482		40000		0		0				1

		155		55		0.000294111		40000		0		0				1

		156		55		841.201		40000		0		0				1

		157		55		1649.16		40000		0		0				1

		158		55		2394.41		40000		0		0				1

		159		56		3052.22		40000		0		0				1

		160		56		3600.89		40000		0		0				1

		161		57		4018.9		40000		0		0				1

		162		57		4285.6		40000		0		0				1

		163		58		4406.45		40000		0		0				1

		164		3		271.628         9		62.402		0		0				0

		165		3		288.109         9		57.598		0		0				0

		166		3		-288.109        9		57.598		0		0				0

		167		3		-271.628        9		62.402		0		0				0

		168		58		-4406.45		40000		0		0				1

		169		57		-4285.6		40000		0		0				1

		170		57		-4018.9		40000		0		0				1

		171		56		-3600.89		40000		0		0				1

		172		56		-3052.22		40000		0		0				1

		173		55		-2394.41		40000		0		0				1

		174		55		-1649.16		40000		0		0				1

		175		55		-841.201		40000		0		0				1

		176		55		0.000294111		40000		0		0				1

		177		54		0.000334223		40000		0		0				1

		178		54		914.88		40000		0		0				1

		179		54		1789.72		40000		0		0				1

		180		55		2587.99		40000		0		0				1

		181		55		3279.69		40000		0		0				1

		182		55		3843.03		40000		0		0				1

		183		56		4269.09		40000		0		0				1

		184		56		4566.07		40000		0		0				1

		185		3		269.527         9		62.993		0		0				0

		186		3		289.51         9		57.175		0		0				0

		187		3		307.059         9		51.691		0		0				0

		188		3		-307.059        9		51.691		0		0				0

		189		3		-289.51         9		57.175		0		0				0

		190		3		-269.527        9		62.993		0		0				0

		191		56		-4566.07		40000		0		0				1

		192		56		-4269.09		40000		0		0				1

		193		55		-3843.03		40000		0		0				1

		194		55		-3279.69		40000		0		0				1

		195		55		-2587.99		40000		0		0				1

		196		54		-1789.72		40000		0		0				1

		197		54		-914.88		40000		0		0				1

		198		54		0.000334223		40000		0		0				1

		199		54		0.000179083		40000		0		0				1

		200		54		987.791		40000		0		0				1

		201		54		1929.63		40000		0		0				1

		202		54		2783.02		40000		0		0				1

		203		54		3515.16		40000		0		0				1

		204		55		4108.49		40000		0		0				1

		205		55		4566.06		40000		0		0				1

		206		55		4909.21		40000		0		0				1

		207		3		284.244         9		58.752		0		0				0

		208		3		305.263         9		52.268		0		0				0

		209		3		323.748         9		46.143		0		0				0

		210		3		-323.748        9		46.143		0		0				0

		211		3		-305.263        9		52.268		0		0				0

		212		3		-284.244        9		58.752		0		0				0

		213		55		-4909.21		40000		0		0				1

		214		55		-4566.06		40000		0		0				1

		215		55		-4108.49		40000		0		0				1

		216		54		-3515.16		40000		0		0				1

		217		54		-2783.02		40000		0		0				1

		218		54		-1929.63		40000		0		0				1

		219		54		-987.79		40000		0		0				1

		220		54		0.000179083		40000		0		0				1

																				%%

																202				91.8

																18				8.2





180 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		58		0.000130069		40000		0		0				1

		2		58		256.212		40000		0		0				1

		3		58		504.935		40000		0		0				1

		4		58		739.688		40000		0		0				1

		5		58		955.693		40000		0		0				1

		6		58		1150.14		40000		0		0				1

		7		58		1322.05		40000		0		0				1

		8		58		1471.89		40000		0		0				1

		9		58		1601.08		40000		0		0				1

		10		58		1711.63		40000		0		0				1

		11		58		1805.75		40000		0		0				1

		12		58		-1805.75		40000		0		0				1

		13		58		-1711.63		40000		0		0				1

		14		58		-1601.08		40000		0		0				1

		15		58		-1471.89		40000		0		0				1

		16		58		-1322.05		40000		0		0				1

		17		58		-1150.14		40000		0		0				1

		18		58		-955.693		40000		0		0				1

		19		58		-739.688		40000		0		0				1

		20		58		-504.935		40000		0		0				1

		21		58		-256.212		40000		0		0				1

		22		58		0.000130069		40000		0		0				1

		23		57		0.000274742		40000		0		0				1

		24		57		365.067		40000		0		0				1

		25		57		720.26		40000		0		0				1

		26		57		1057.07		40000		0		0				1

		27		58		1369.24		40000		0		0				1

		28		58		1653.03		40000		0		0				1

		29		58		1906.95		40000		0		0				1

		30		58		2131.16		40000		0		0				1

		31		58		2326.91		40000		0		0				1

		32		58		2496.12		40000		0		0				1

		33		57		2641.08		40000		0		0				1

		34		57		-2641.08		40000		0		0				1

		35		58		-2496.12		40000		0		0				1

		36		58		-2326.91		40000		0		0				1

		37		58		-2131.16		40000		0		0				1

		38		58		-1906.95		40000		0		0				1

		39		58		-1653.03		40000		0		0				1

		40		58		-1369.24		40000		0		0				1

		41		57		-1057.07		40000		0		0				1

		42		57		-720.26		40000		0		0				1

		43		57		-365.067		40000		0		0				1

		44		57		0.000274742		40000		0		0				1

		45		57		0.000199026		40000		0		0				1

		46		57		445.889		40000		0		0				1

		47		57		880.242		40000		0		0				1

		48		57		1293.38		40000		0		0				1

		49		57		1678.64		40000		0		0				1

		50		57		2032.69		40000		0		0				1

		51		57		2354.89		40000		0		0				1

		52		57		2646.35		40000		0		0				1

		53		57		2908.86		40000		0		0				1

		54		57		3144.04		40000		0		0				1

		55		57		3352.94		40000		0		0				1

		56		57		-3352.94		40000		0		0				1

		57		57		-3144.04		40000		0		0				1

		58		57		-2908.86		40000		0		0				1

		59		57		-2646.35		40000		0		0				1

		60		57		-2354.89		40000		0		0				1

		61		57		-2032.69		40000		0		0				1

		62		57		-1678.64		40000		0		0				1

		63		57		-1293.38		40000		0		0				1

		64		57		-880.242		40000		0		0				1

		65		57		-445.889		40000		0		0				1

		66		57		0.000199026		40000		0		0				1

		67		57		0.000153554		40000		0		0				1

		68		57		511.496		40000		0		0				1

		69		57		1009.57		40000		0		0				1

		70		57		1483.21		40000		0		0				1

		71		57		1925.29		40000		0		0				1

		72		57		2333.55		40000		0		0				1

		73		57		2709.44		40000		0		0				1

		74		57		3057.41		40000		0		0				1

		75		57		3383.54		40000		0		0				1

		76		57		3694.05		40000		0		0				1

		77		57		3993.62		40000		0		0				1

		78		57		-3993.62		40000		0		0				1

		79		57		-3694.05		40000		0		0				1

		80		57		-3383.54		40000		0		0				1

		81		57		-3057.41		40000		0		0				1

		82		57		-2709.44		40000		0		0				1

		83		57		-2333.55		40000		0		0				1

		84		57		-1925.33		40000		0		0				1

		85		57		-1483.21		40000		0		0				1

		86		57		-1009.57		40000		0		0				1

		87		57		-511.496		40000		0		0				1

		88		57		0.000153554		40000		0		0				1

		89		56		0.000270452		40000		0		0				1

		90		56		567.653		40000		0		0				1

		91		56		1119.46		40000		0		0				1

		92		57		1642.43		40000		0		0				1

		93		57		2128.44		40000		0		0				1

		94		57		2574.9		40000		0		0				1

		95		57		2984.08		40000		0		0				1

		96		57		3363.59		40000		0		0				1

		97		58		3724.41		40000		0		0				1

		98		58		4082.61		40000		0		0				1

		99		58		4460.78		40000		0		0				1

		100		58		-4460.78		40000		0		0				1

		101		58		-4082.61		40000		0		0				1

		102		58		-3724.41		40000		0		0				1

		103		57		-3363.59		40000		0		0				1

		104		57		-2984.08		40000		0		0				1

		105		57		-2574.9		40000		0		0				1

		106		57		-2128.44		40000		0		0				1

		107		57		-1642.43		40000		0		0				1

		108		56		-1119.46		40000		0		0				1

		109		56		-567.653		40000		0		0				1

		110		56		0.000270452		40000		0		0				1

		111		56		0.000222409		40000		0		0				1

		112		56		663.977		40000		0		0				1

		113		56		1305.9		40000		0		0				1

		114		56		1906.89		40000		0		0				1

		115		56		2453.81		40000		0		0				1

		116		57		2938.76		40000		0		0				1

		117		57		3357.28		40000		0		0				1

		118		58		3705.7		40000		0		0				1

		119		58		3974.02		40000		0		0				1

		120		59		4138.47		40000		0		0				1

		121		61		4147.06		40000		0		0				1

		122		61		-4147.06		40000		0		0				1

		123		59		-4138.47		40000		0		0				1

		124		58		-3974.02		40000		0		0				1

		125		58		-3705.7		40000		0		0				1

		126		57		-3357.28		40000		0		0				1

		127		57		-2938.76		40000		0		0				1

		128		56		-2453.81		40000		0		0				1

		129		56		-1906.89		40000		0		0				1

		130		56		-1305.9		40000		0		0				1

		131		56		-663.977		40000		0		0				1

		132		56		0.000222409		40000		0		0				1

		133		55		0.000112233		40000		0		0				1

		134		55		749.559		40000		0		0				1

		135		55		1470.18		40000		0		0				1

		136		55		2136.48		40000		0		0				1

		137		56		2728.8		40000		0		0				1

		138		56		3233.13		40000		0		0				1

		139		57		3640.09		40000		0		0				1

		140		57		3941.54		40000		0		0				1

		141		58		4133.52		40000		0		0				1

		142		59		4221.35		40000		0		0				1

		143		59		4229.98		40000		0		0				1

		144		59		-4229.98		40000		0		0				1

		145		59		-4221.35		40000		0		0				1

		146		58		-4133.52		40000		0		0				1

		147		57		-3941.54		40000		0		0				1

		148		57		-3640.09		40000		0		0				1

		149		56		-3233.13		40000		0		0				1

		150		56		-2728.8		40000		0		0				1

		151		55		-2136.48		40000		0		0				1

		152		55		-1470.18		40000		0		0				1

		153		55		-749.559		40000		0		0				1

		154		55		0.000112233		40000		0		0				1

		155		55		0.000267754		40000		0		0				1

		156		55		830.793		40000		0		0				1

		157		55		1626.18		40000		0		0				1

		158		55		2354.6		40000		0		0				1

		159		55		2991.54		40000		0		0				1

		160		55		3521.16		40000		0		0				1

		161		56		3938.31		40000		0		0				1

		162		56		4246.41		40000		0		0				1

		163		57		4460.15		40000		0		0				1

		164		3		266.32         9		63.885		0		0				0

		165		3		282.344         9		59.313		0		0				0

		166		3		-282.344        9		59.313		0		0				0

		167		3		-266.32         9		63.885		0		0				0

		168		57		-4460.15		40000		0		0				1

		169		56		-4246.41		40000		0		0				1

		170		56		-3938.31		40000		0		0				1

		171		55		-3521.16		40000		0		0				1

		172		55		-2991.54		40000		0		0				1

		173		55		-2354.6		40000		0		0				1

		174		55		-1626.18		40000		0		0				1

		175		55		-830.793		40000		0		0				1

		176		55		0.000267754		40000		0		0				1

		177		54		0.000245862		40000		0		0				1

		178		54		910.262		40000		0		0				1

		179		54		1779.66		40000		0		0				1

		180		54		2571.3		40000		0		0				1

		181		54		3256.95		40000		0		0				1

		182		55		3821.88		40000		0		0				1

		183		55		4266.85		40000		0		0				1

		184		55		4605.49		40000		0		0				1

		185		3		264.605         9		64.357		0		0				0

		186		3		284.031         9		58.815		0		0				0

		187		3		301.088         9		53.596		0		0				0

		188		3		-301.088        9		53.596		0		0				0

		189		3		-284.031        9		58.815		0		0				0

		190		3		-264.605        9		64.357		0		0				0

		191		55		-4605.49		40000		0		0				1

		192		55		-4266.85		40000		0		0				1

		193		55		-3821.88		40000		0		0				1

		194		54		-3256.95		40000		0		0				1

		195		54		-2571.3		40000		0		0				1

		196		54		-1779.66		40000		0		0				1

		197		54		-910.262		40000		0		0				1

		198		54		0.000245862		40000		0		0				1

		199		54		0.000170433		40000		0		0				1

		200		54		989.091		40000		0		0				1

		201		54		1933.28		40000		0		0				1

		202		54		2791.14		40000		0		0				1

		203		54		3530.53		40000		0		0				1

		204		54		4136.27		40000		0		0				1

		205		54		4614.17		40000		0		0				1

		206		54		4984.63		40000		0		0				1

		207		3		279.214         9		60.229		0		0				0

		208		3		299.631         9		54.055		0		0				0

		209		3		317.595         9		48.227		0		0				0

		210		3		-317.595        9		48.227		0		0				0

		211		3		-299.631        9		54.055		0		0				0

		212		3		-279.214        9		60.229		0		0				0

		213		54		-4984.63		40000		0		0				1

		214		54		-4614.17		40000		0		0				1

		215		54		-4136.27		40000		0		0				1

		216		54		-3530.53		40000		0		0				1

		217		54		-2791.14		40000		0		0				1

		218		54		-1933.28		40000		0		0				1

		219		54		-989.091		40000		0		0				1

		220		54		0.000170433		40000		0		0				1

																				%%

																204				92.7

																16				7.3





190 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		58		0.000140451		40000		0		0				1

		2		58		255.278		40000		0		0				1

		3		58		503.433		40000		0		0				1

		4		58		738.305		40000		0		0				1

		5		58		955.359		40000		0		0				1

		6		58		1151.92		40000		0		0				1

		7		58		1327		40000		0		0				1

		8		58		1480.97		40000		0		0				1

		9		58		1615.08		40000		0		0				1

		10		58		1731.1		40000		0		0				1

		11		58		1831.01		40000		0		0				1

		12		58		-1831.01		40000		0		0				1

		13		58		-1731.1		40000		0		0				1

		14		58		-1615.08		40000		0		0				1

		15		58		-1480.97		40000		0		0				1

		16		58		-1327		40000		0		0				1

		17		58		-1151.92		40000		0		0				1

		18		58		-955.359		40000		0		0				1

		19		58		-738.305		40000		0		0				1

		20		58		-503.433		40000		0		0				1

		21		58		-255.278		40000		0		0				1

		22		58		0.000140451		40000		0		0				1

		23		57		0.000166492		40000		0		0				1

		24		57		358.908		40000		0		0				1

		25		57		708.223		40000		0		0				1

		26		57		1039.77		40000		0		0				1

		27		57		1347.72		40000		0		0				1

		28		57		1628.86		40000		0		0				1

		29		57		1882.26		40000		0		0				1

		30		57		2108.62		40000		0		0				1

		31		57		2309.65		40000		0		0				1

		32		57		2487.47		40000		0		0				1

		33		57		2644.29		40000		0		0				1

		34		57		-2644.29		40000		0		0				1

		35		57		-2487.47		40000		0		0				1

		36		57		-2309.65		40000		0		0				1

		37		57		-2108.62		40000		0		0				1

		38		57		-1882.26		40000		0		0				1

		39		57		-1628.86		40000		0		0				1

		40		57		-1347.72		40000		0		0				1

		41		57		-1039.77		40000		0		0				1

		42		57		-708.223		40000		0		0				1

		43		57		-358.908		40000		0		0				1

		44		57		0.000166492		40000		0		0				1

		45		57		0.000186823		40000		0		0				1

		46		57		435.232		40000		0		0				1

		47		57		858.547		40000		0		0				1

		48		57		1259.95		40000		0		0				1

		49		57		1632.67		40000		0		0				1

		50		57		1973.28		40000		0		0				1

		51		57		2281.99		40000		0		0				1

		52		57		2560.79		40000		0		0				1

		53		57		2813.12		40000		0		0				1

		54		58		3042.81		40000		0		0				1

		55		58		3253.53		40000		0		0				1

		56		58		-3253.53		40000		0		0				1

		57		58		-3042.81		40000		0		0				1

		58		57		-2813.12		40000		0		0				1

		59		57		-2560.79		40000		0		0				1

		60		57		-2281.99		40000		0		0				1

		61		57		-1973.28		40000		0		0				1

		62		57		-1632.67		40000		0		0				1

		63		57		-1259.95		40000		0		0				1

		64		57		-858.547		40000		0		0				1

		65		57		-435.232		40000		0		0				1

		66		57		0.000186823		40000		0		0				1

		67		56		0.000220712		40000		0		0				1

		68		56		498.014		40000		0		0				1

		69		56		981.417		40000		0		0				1

		70		57		1437.98		40000		0		0				1

		71		57		1859.46		40000		0		0				1

		72		57		2242.1		40000		0		0				1

		73		57		2585.88		40000		0		0				1

		74		57		2894.33		40000		0		0				1

		75		58		3172.41		40000		0		0				1

		76		58		3426.3		40000		0		0				1

		77		58		3662.59		40000		0		0				1

		78		58		-3662.59		40000		0		0				1

		79		58		-3426.3		40000		0		0				1

		80		58		-3172.41		40000		0		0				1

		81		57		-2894.33		40000		0		0				1

		82		57		-2585.88		40000		0		0				1

		83		57		-2242.1		40000		0		0				1

		84		57		-1859.46		40000		0		0				1

		85		57		-1437.98		40000		0		0				1

		86		56		-981.417		40000		0		0				1

		87		56		-498.014		40000		0		0				1

		88		56		0.000220712		40000		0		0				1

		89		56		0.000147549		40000		0		0				1

		90		56		553.118		40000		0		0				1

		91		56		1088.57		40000		0		0				1

		92		56		1591.46		40000		0		0				1

		93		56		2051.48		40000		0		0				1

		94		57		2463.36		40000		0		0				1

		95		57		2826.27		40000		0		0				1

		96		57		3142.09		40000		0		0				1

		97		58		3415.13		40000		0		0				1

		98		58		3649.06		40000		0		0				1

		99		59		3847.83		40000		0		0				1

		100		59		-3847.83		40000		0		0				1

		101		58		-3649.06		40000		0		0				1

		102		58		-3415.13		40000		0		0				1

		103		57		-3142.09		40000		0		0				1

		104		57		-2826.27		40000		0		0				1

		105		57		-2463.36		40000		0		0				1

		106		56		-2051.48		40000		0		0				1

		107		56		-1591.46		40000		0		0				1

		108		56		-1088.57		40000		0		0				1

		109		56		-553.117		40000		0		0				1

		110		56		0.000147549		40000		0		0				1

		111		55		0.000129428		40000		0		0				1

		112		55		651.2		40000		0		0				1

		113		55		1278.24		40000		0		0				1

		114		56		1860.26		40000		0		0				1

		115		56		2381.92		40000		0		0				1

		116		56		2833.85		40000		0		0				1

		117		56		3212.36		40000		0		0				1

		118		57		3516.61		40000		0		0				1

		119		57		3747.16		40000		0		0				1

		120		58		3906.19		40000		0		0				1

		121		59		3995.83		40000		0		0				1

		122		59		-3995.83		40000		0		0				1

		123		58		-3906.19		40000		0		0				1

		124		57		-3747.16		40000		0		0				1

		125		57		-3516.61		40000		0		0				1

		126		56		-3212.36		40000		0		0				1

		127		56		-2833.85		40000		0		0				1

		128		56		-2381.92		40000		0		0				1

		129		56		-1860.26		40000		0		0				1

		130		55		-1278.24		40000		0		0				1

		131		55		-651.2		40000		0		0				1

		132		55		0.000129428		40000		0		0				1

		133		55		0.000154788		40000		0		0				1

		134		55		741.567		40000		0		0				1

		135		55		1452.78		40000		0		0				1

		136		55		2107.02		40000		0		0				1

		137		55		2684.19		40000		0		0				1

		138		55		3173.08		40000		0		0				1

		139		56		3570.1		40000		0		0				1

		140		56		3878.05		40000		0		0				1

		141		56		4105.35		40000		0		0				1

		142		57		4263.26		40000		0		0				1

		143		57		4364.97		40000		0		0				1

		144		57		-4364.97		40000		0		0				1

		145		57		-4263.26		40000		0		0				1

		146		56		-4105.35		40000		0		0				1

		147		56		-3878.2		40000		0		0				1

		148		56		-3570.1		40000		0		0				1

		149		55		-3173.08		40000		0		0				1

		150		55		-2684.19		40000		0		0				1

		151		55		-2107.02		40000		0		0				1

		152		55		-1452.78		40000		0		0				1

		153		55		-741.567		40000		0		0				1

		154		55		0.000154788		40000		0		0				1

		155		54		0.000308634		40000		0		0				1

		156		54		828.551		40000		0		0				1

		157		54		1621.52		40000		0		0				1

		158		54		2347.25		40000		0		0				1

		159		54		2981.9		40000		0		0				1

		160		55		3512.81		40000		0		0				1

		161		55		3939.36		40000		0		0				1

		162		55		4270.56		40000		0		0				1

		163		55		4521.6		40000		0		0				1

		164		56		4708.37		40000		0		0				1

		165		3		277.17         9		60.821		0		0				0

		166		3		-277.17         9		60.821		0		0				0

		167		56		-4708.37		40000		0		0				1

		168		55		-4521.6		40000		0		0				1

		169		55		-4270.56		40000		0		0				1

		170		55		-3939.36		40000		0		0				1

		171		55		-3512.81		40000		0		0				1

		172		54		-2981.9		40000		0		0				1

		173		54		-2347.25		40000		0		0				1

		174		54		-1621.52		40000		0		0				1

		175		54		-828.551		40000		0		0				1

		176		54		0.000308634		40000		0		0				1

		177		54		0.000310413		40000		0		0				1

		178		54		913.841		40000		0		0				1

		179		54		1788.32		40000		0		0				1

		180		54		2587.53		40000		0		0				1

		181		54		3283.42		40000		0		0				1

		182		54		3861.13		40000		0		0				1

		183		54		4322.16		40000		0		0				1

		184		54		4680.95		40000		0		0				1

		185		54		4958.36		40000		0		0				1

		186		3		279.141		960.25		0		0				0

		187		3		295.701		955.28		0		0				0

		188		3		-295.701		955.28		0		0				0

		189		3		-279.141		960.25		0		0				0

		190		54		-4958.35		40000		0		0				1

		191		54		-4680.95		40000		0		0				1

		192		54		-4322.16		40000		0		0				1

		193		54		-3861.13		40000		0		0				1

		194		54		-3283.42		40000		0		0				1

		195		54		-2587.53		40000		0		0				1

		196		54		-1788.32		40000		0		0				1

		197		54		-913.841		40000		0		0				1

		198		54		0.000310413		40000		0		0				1

		199		53		0.000114748		40000		0		0				1

		200		53		997.59		40000		0		0				1

		201		53		1953.9		40000		0		0				1

		202		53		2829.8		40000		0		0				1

		203		53		3591.98		40000		0		0				1

		204		53		4219.97		40000		0		0				1

		205		53		4714		40000		0		0				1

		206		53		5094.63		40000		0		0				1

		207		3		274.79         9		61.504		0		0				0

		208		3		294.574         9		55.629		0		0				0

		209		3		312.001         9		50.082		0		0				0

		210		3		-312.001        9		50.082		0		0				0

		211		3		-294.574        9		55.629		0		0				0

		212		3		-274.79         9		61.504		0		0				0

		213		53		-5094.63		40000		0		0				1

		214		53		-4714		40000		0		0				1

		215		53		-4220.06		40000		0		0				1

		216		53		-3591.98		40000		0		0				1

		217		53		-2829.8		40000		0		0				1

		218		53		-1953.9		40000		0		0				1

		219		53		-997.59		40000		0		0				1

		220		53		0.000114748		40000		0		0				1

																				%%

																208				94.5

																12				5.5





200 Volt

		

				Beam #1 Data

				Point Data

		Ray #		Point #		X[mic]		Y[mic]		Z[mic]		I[A]

		-----------		-------------------		-----------------------		----------		-----------------		----------------

		1		57		0.000141625		40000		0		0				1

		2		57		250.96		40000		0		0				1

		3		57		495.009		40000		0		0				1

		4		57		726.201		40000		0		0				1

		5		57		940.198		40000		0		0				1

		6		57		1134.49		40000		0		0				1

		7		57		1308.24		40000		0		0				1

		8		57		1461.87		40000		0		0				1

		9		57		1596.63		40000		0		0				1

		10		57		1714.22		40000		0		0				1

		11		57		1816.51		40000		0		0				1

		12		57		-1816.51		40000		0		0				1

		13		57		-1714.22		40000		0		0				1

		14		57		-1596.63		40000		0		0				1

		15		57		-1461.87		40000		0		0				1

		16		57		-1308.24		40000		0		0				1

		17		57		-1134.49		40000		0		0				1

		18		57		-940.198		40000		0		0				1

		19		57		-726.201		40000		0		0				1

		20		57		-495.009		40000		0		0				1

		21		57		-250.96		40000		0		0				1

		22		57		0.000141625		40000		0		0				1

		23		57		0.000174583		40000		0		0				1

		24		57		349.685		40000		0		0				1

		25		57		689.585		40000		0		0				1

		26		57		1011.39		40000		0		0				1

		27		57		1309.26		40000		0		0				1

		28		57		1579.95		40000		0		0				1

		29		57		1823.01		40000		0		0				1

		30		57		2039.42		40000		0		0				1

		31		57		2231.37		40000		0		0				1

		32		57		2401.47		40000		0		0				1

		33		57		2552.43		40000		0		0				1

		34		57		-2552.42		40000		0		0				1

		35		57		-2401.47		40000		0		0				1

		36		57		-2231.37		40000		0		0				1

		37		57		-2039.42		40000		0		0				1

		38		57		-1823.01		40000		0		0				1

		39		57		-1579.95		40000		0		0				1

		40		57		-1309.26		40000		0		0				1

		41		57		-1011.39		40000		0		0				1

		42		57		-689.585		40000		0		0				1

		43		57		-349.685		40000		0		0				1

		44		57		0.000174583		40000		0		0				1

		45		56		0.000196738		40000		0		0				1

		46		56		423.522		40000		0		0				1

		47		56		834.353		40000		0		0				1

		48		57		1221.78		40000		0		0				1

		49		57		1578.36		40000		0		0				1

		50		57		1900.38		40000		0		0				1

		51		57		2187.35		40000		0		0				1

		52		57		2440.96		40000		0		0				1

		53		57		2664.92		40000		0		0				1

		54		57		2863.09		40000		0		0				1

		55		58		3039.53		40000		0		0				1

		56		58		-3039.53		40000		0		0				1

		57		57		-2863.09		40000		0		0				1

		58		57		-2664.92		40000		0		0				1

		59		57		-2440.96		40000		0		0				1

		60		57		-2187.35		40000		0		0				1

		61		57		-1900.38		40000		0		0				1

		62		57		-1578.36		40000		0		0				1

		63		57		-1221.78		40000		0		0				1

		64		56		-834.353		40000		0		0				1

		65		56		-423.522		40000		0		0				1

		66		56		0.000196738		40000		0		0				1

		67		56		0.000221042		40000		0		0				1

		68		56		485.964		40000		0		0				1

		69		56		956.054		40000		0		0				1

		70		56		1397.01		40000		0		0				1

		71		56		1799.45		40000		0		0				1

		72		56		2158.55		40000		0		0				1

		73		57		2473.45		40000		0		0				1

		74		57		2746.24		40000		0		0				1

		75		57		2980.41		40000		0		0				1

		76		58		3180.99		40000		0		0				1

		77		58		3352.36		40000		0		0				1

		78		58		-3352.36		40000		0		0				1

		79		58		-3180.99		40000		0		0				1

		80		57		-2980.41		40000		0		0				1

		81		57		-2746.24		40000		0		0				1

		82		57		-2473.45		40000		0		0				1

		83		56		-2158.55		40000		0		0				1

		84		56		-1799.45		40000		0		0				1

		85		56		-1397.01		40000		0		0				1

		86		56		-956.054		40000		0		0				1

		87		56		-485.964		40000		0		0				1

		88		56		0.000221042		40000		0		0				1

		89		56		0.000152494		40000		0		0				1

		90		56		542.121		40000		0		0				1

		91		56		1065.27		40000		0		0				1

		92		56		1553.27		40000		0		0				1

		93		56		1994.66		40000		0		0				1

		94		56		2383		40000		0		0				1

		95		56		2717.15		40000		0		0				1

		96		57		2998.75		40000		0		0				1

		97		57		3231.32		40000		0		0				1

		98		57		3419.72		40000		0		0				1

		99		58		3568.23		40000		0		0				1

		100		58		-3568.23		40000		0		0				1

		101		57		-3419.72		40000		0		0				1

		102		57		-3231.32		40000		0		0				1

		103		57		-2998.75		40000		0		0				1

		104		56		-2717.15		40000		0		0				1

		105		56		-2383		40000		0		0				1

		106		56		-1994.66		40000		0		0				1

		107		56		-1553.27		40000		0		0				1

		108		56		-1065.27		40000		0		0				1

		109		56		-542.121		40000		0		0				1

		110		56		0.000152494		40000		0		0				1

		111		55		0.000121772		40000		0		0				1

		112		55		644.79		40000		0		0				1

		113		55		1264.39		40000		0		0				1

		114		55		1837.26		40000		0		0				1

		115		55		2347.62		40000		0		0				1

		116		55		2786.81		40000		0		0				1

		117		56		3152.74		40000		0		0				1

		118		56		3448.33		40000		0		0				1

		119		56		3679.2		40000		0		0				1

		120		57		3853.43		40000		0		0				1

		121		57		3978.4		40000		0		0				1

		122		57		-3978.4		40000		0		0				1

		123		57		-3853.43		40000		0		0				1

		124		56		-3679.2		40000		0		0				1

		125		56		-3448.33		40000		0		0				1

		126		56		-3152.74		40000		0		0				1

		127		55		-2786.81		40000		0		0				1

		128		55		-2347.62		40000		0		0				1

		129		55		-1837.26		40000		0		0				1

		130		55		-1264.39		40000		0		0				1

		131		55		-644.79		40000		0		0				1

		132		55		0.000121772		40000		0		0				1

		133		54		0.000162819		40000		0		0				1

		134		54		740.848		40000		0		0				1

		135		54		1451.42		40000		0		0				1

		136		54		2105.11		40000		0		0				1

		137		55		2682.19		40000		0		0				1

		138		55		3171.98		40000		0		0				1

		139		55		3573.24		40000		0		0				1

		140		55		3892.18		40000		0		0				1

		141		55		4138.89		40000		0		0				1

		142		56		4326.08		40000		0		0				1

		143		56		4464.86		40000		0		0				1

		144		56		-4464.86		40000		0		0				1

		145		56		-4326.08		40000		0		0				1

		146		55		-4138.89		40000		0		0				1

		147		55		-3892.18		40000		0		0				1

		148		55		-3573.24		40000		0		0				1

		149		55		-3171.98		40000		0		0				1

		150		55		-2682.19		40000		0		0				1

		151		54		-2105.11		40000		0		0				1

		152		54		-1451.42		40000		0		0				1

		153		54		-740.848		40000		0		0				1

		154		54		0.000162819		40000		0		0				1

		155		54		0.000241249		40000		0		0				1

		156		54		833.387		40000		0		0				1

		157		54		1632.9		40000		0		0				1

		158		54		2367.64		40000		0		0				1

		159		54		3013.78		40000		0		0				1

		160		54		3557.75		40000		0		0				1

		161		54		3998.33		40000		0		0				1

		162		54		4345.08		40000		0		0				1

		163		54		4613.66		40000		0		0				1

		164		54		4820.54		40000		0		0				1

		165		3		272.575         9		62.135		0		0				0

		166		3		-272.575        9		62.135		0		0				0

		167		54		-4820.54		40000		0		0				1

		168		54		-4613.66		40000		0		0				1

		169		54		-4345.17		40000		0		0				1

		170		54		-3998.33		40000		0		0				1

		171		54		-3557.75		40000		0		0				1

		172		54		-3013.78		40000		0		0				1

		173		54		-2367.64		40000		0		0				1

		174		54		-1632.9		40000		0		0				1

		175		54		-833.404		40000		0		0				1

		176		54		0.000241249		40000		0		0				1

		177		54		0.000316003		40000		0		0				1

		178		54		923.195		40000		0		0				1

		179		54		1810.58		40000		0		0				1

		180		53		2628.59		40000		0		0				1

		181		53		3348.57		40000		0		0				1

		182		53		3950.99		40000		0		0				1

		183		53		4431.24		40000		0		0				1

		184		53		4802.18		40000		0		0				1

		185		53		5087.04		40000		0		0				1

		186		3		274.849         9		61.487		0		0				0

		187		3		290.882         9		56.759		0		0				0

		188		3		-290.882        9		56.759		0		0				0

		189		3		-274.849        9		61.487		0		0				0

		190		53		-5087.04		40000		0		0				1

		191		53		-4802.18		40000		0		0				1

		192		53		-4431.43		40000		0		0				1

		193		53		-3950.99		40000		0		0				1

		194		53		-3348.57		40000		0		0				1

		195		53		-2628.59		40000		0		0				1

		196		54		-1810.58		40000		0		0				1

		197		54		-923.195		40000		0		0				1

		198		54		0.000316003		40000		0		0				1

		199		53		0.000104877		40000		0		0				1

		200		53		1009.97		40000		0		0				1

		201		53		1984.56		40000		0		0				1

		202		53		2889.22		40000		0		0				1

		203		53		3690.53		40000		0		0				1

		204		52		4357.41		40000		0		0				1

		205		52		4874.49		40000		0		0				1

		206		52		5257.68		40000		0		0				1

		207		3		271.018         9		62.574		0		0				0

		208		3		290.094         9		56.998		0		0				0

		209		3		306.936		951.73		0		0				0

		210		3		-306.936		951.73		0		0				0

		211		3		-290.094        9		56.998		0		0				0

		212		3		-271.018        9		62.574		0		0				0

		213		52		-5257.68		40000		0		0				1

		214		52		-4874.49		40000		0		0				1

		215		52		-4357.41		40000		0		0				1

		216		53		-3690.53		40000		0		0				1

		217		53		-2889.22		40000		0		0				1

		218		53		-1984.56		40000		0		0				1

		219		53		-1009.97		40000		0		0				1

		220		53		0.000104877		40000		0		0				1

																				%%

																208				94.5

																12				5.5






